## Dashboard
BIOMEDICAL ENGINEERING DASHBOARD - ENGINEER'S PERFORMANCE
This shape represents a slicer. Slicers are supported in Excel 2010 or later. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2003 or earlier, the slicer cannot be used.
This shape represents a slicer. Slicers are supported in Excel 2010 or later. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2003 or earlier, the slicer cannot be used.
NO. OF CALLS ATTENDED
TOTAL HOURS SPENT
533
AVG. RESPONSE TIME
4 Min
AVG. RESOLUTION TIME
32 Min
BIOMEDICAL ENGINEERS
MONTH & YEAR
OVERALL DEPT WISE (%)
HAFIZ V MOHAMMED RAIYYAN
12792 Hours : 00 Minutes

## Background Analysis
NO. OF CALLS ATTENDED |  |  |  | NO. OF CALLS PENDING |  |  |  | AVG. RESPONSE TIME |  |  |  | AVG. RESOLUTION TIME
Row Labels | Count of START MONTH & YEAR |  |  | Row Labels | Count of STATUS |  |  | Row Labels | Average of RESPONSE TIME (IN MINUTES) |  |  | Row Labels | Average of TIME DOWN
HAFIZ | 533 |  |  | HAFIZ | 533 |  |  | HAFIZ | 0.0031739 |  |  | HAFIZ | 0.022277
JERLIN | 428 |  |  | JERLIN | 428 |  |  | JERLIN | 0.0063263 |  |  | JERLIN | 0.018106
RAMKIE | 276 |  |  | RAMKIE | 276 |  |  | RAMKIE | 0.0078804 |  |  | Apr 2024 | 0.024231
SANTHOSH | 287 |  |  | SANTHOSH | 287 |  |  | SANTHOSH | 0.0038811 |  |  | Aug 2023 | 0.016554
SOORAJ | 363 |  |  | SOORAJ | 363 |  |  | SOORAJ | 0.0031987 |  |  | Aug 2024 | 0.024028
THANISH | 207 |  |  | THANISH | 207 |  |  | THANISH | 0.0039016 |  |  | Dec 2023 | 0.013889
Grand Total | 2094 |  |  | Grand Total | 2094 |  |  | Grand Total | 0.0046117 |  |  | Feb 2024 | 0.015951
[12 empty] | Jan 2024 | 0.013244
[12 empty] | Jul 2023 | 0.012566
[12 empty] | Jul 2024 | 0.019508
[12 empty] | Jun 2023 | 0.027778
[12 empty] | Jun 2024 | 0.019949
AVG. RESOLUTION TIME | [7 empty] | AVG. RESPONSE TIME |  |  |  | Mar 2024 | 0.017157
Row Labels | Average of TIME DOWN | [6 empty] | Row Labels | Average of RESPONSE TIME (IN MINUTES) |  |  | May 2024 | 0.031927
HAFIZ | 0.022277 | [6 empty] | HAFIZ | 0.0031739 |  |  | Nov 2023 | 0.012032
JERLIN | 0.018106 | [6 empty] | JERLIN | 0.0063263 |  |  | Oct 2023 | 0.01558
RAMKIE | 0.022985 | [6 empty] | RAMKIE | 0.0078804 |  |  | Sep 2023 | 0.017178
SANTHOSH | 0.024381 | [6 empty] | SANTHOSH | 0.0038811 |  |  | RAMKIE | 0.022985
SOORAJ | 0.040658 | [6 empty] | SOORAJ | 0.0031987 |  |  | SANTHOSH | 0.024381
THANISH | 0.038721 | [6 empty] | THANISH | 0.0039016 |  |  | SOORAJ | 0.040658
Grand Total | 0.026618 | [6 empty] | Grand Total | 0.0046117 |  |  | THANISH | 0.038721
[12 empty] | Grand Total | 0.026618
CRITICAL VS NON-CRITICAL VS SEMI-CRITICAL
Count of STATUS | Column Labels
ENGINEER | CRITICAL | NON-CRITICAL | SEMI-CRITICAL | Grand Total
HAFIZ | 131 | 104 | 298 | 533
JERLIN | 80 | 91 | 257 | 428
RAMKIE | 71 | 75 | 130 | 276
SANTHOSH | 77 | 62 | 148 | 287
SOORAJ | 124 | 55 | 184 | 363
THANISH | 53 | 35 | 119 | 207
Grand Total | 536 | 422 | 1136 | 2094
Count of DEPARTMENT | Column Labels
Row Labels | HAFIZ | JERLIN | RAMKIE | SANTHOSH | SOORAJ | THANISH | Grand Total
BLOODBANK | 0.2388% | 0% | 0.382% | 0.1433% | 0.04776% | 0% | 0.8118%
CATHLAB | 0.7641% | 0.6686% | 0.8596% | 0.382% | 0.5253% | 0.2865% | 3.486%
CSSD | 0.4776% | 0.191% | 0.1433% | 0.2865% | 0.8118% | 0.191% | 2.101%
DIALYSIS | 0.382% | 0.3343% | 0.6208% | 0.382% | 0.6686% | 0.04776% | 2.436%
EMERGENCY | 1.098% | 0.9074% | 0.7163% | 0.5253% | 0.4298% | 0.2388% | 3.916%
ENDOSCOPY | 0.4298% | 0.3343% | 0.09551% | 0.09551% | 0% | 0.09551% | 1.051%
ICU | 8.787% | 6.065% | 2.961% | 4.298% | 3.82% | 2.579% | 28.51%
LABORATORY | 0.6686% | 0.191% | 1.289% | 0.191% | 0.2388% | 0.04776% | 2.627%
OPD | 2.053% | 1.91% | 1.194% | 2.245% | 0.4776% | 1.91% | 9.79%
OT | 5.253% | 5.158% | 2.579% | 2.77% | 6.781% | 2.197% | 24.74%
PHYSIOTHERAPY | 0.09551% | 0.191% | 0.6208% | 0.2388% | 0% | 0.191% | 1.337%
RADIOLOGY | 0.6686% | 1.098% | 0.8596% | 0.4776% | 1.146% | 1.242% | 5.492%
WARD | 4.537% | 3.391% | 0.8596% | 1.671% | 2.388% | 0.8596% | 13.71%
Grand Total | 25.45% | 20.44% | 13.18% | 13.71% | 17.34% | 9.885% | 100%
[7 empty] | Count of DEPARTMENT | Column Labels
[7 empty] | Row Labels | BLOODBANK | CATHLAB | CSSD | DIALYSIS | EMERGENCY | ENDOSCOPY | ICU | LABORATORY | OPD | OT | PHYSIOTHERAPY | RADIOLOGY | WARD | Grand Total
[7 empty] | HAFIZ | 0.2388% | 0.7641% | 0.4776% | 0.382% | 1.098% | 0.4298% | 8.787% | 0.6686% | 2.053% | 5.253% | 0.09551% | 0.6686% | 4.537% | 25.45%
[7 empty] | JERLIN | 0% | 0.6686% | 0.191% | 0.3343% | 0.9074% | 0.3343% | 6.065% | 0.191% | 1.91% | 5.158% | 0.191% | 1.098% | 3.391% | 20.44%
[7 empty] | RAMKIE | 0.382% | 0.8596% | 0.1433% | 0.6208% | 0.7163% | 0.09551% | 2.961% | 1.289% | 1.194% | 2.579% | 0.6208% | 0.8596% | 0.8596% | 13.18%
[7 empty] | SANTHOSH | 0.1433% | 0.382% | 0.2865% | 0.382% | 0.5253% | 0.09551% | 4.298% | 0.191% | 2.245% | 2.77% | 0.2388% | 0.4776% | 1.671% | 13.71%
[7 empty] | SOORAJ | 0.04776% | 0.5253% | 0.8118% | 0.6686% | 0.4298% | 0% | 3.82% | 0.2388% | 0.4776% | 6.781% | 0% | 1.146% | 2.388% | 17.34%
[7 empty] | THANISH | 0% | 0.2865% | 0.191% | 0.04776% | 0.2388% | 0.09551% | 2.579% | 0.04776% | 1.91% | 2.197% | 0.191% | 1.242% | 0.8596% | 9.885%
[7 empty] | Grand Total | 0.8118% | 3.486% | 2.101% | 2.436% | 3.916% | 1.051% | 28.51% | 2.627% | 9.79% | 24.74% | 1.337% | 5.492% | 13.71% | 100%
CONTRIBUTIONS
Row Labels | Count of STATUS |  |  |  | Row Labels | Count of TIME DOWN
HAFIZ | 25.45% |  |  |  | HAFIZ | 1901-06-16
Apr 2024 | 2.053% |  |  |  | JERLIN | 1901-03-03
Aug 2023 | 1.91% |  |  |  | RAMKIE | 1900-10-02
Aug 2024 | 1.003% |  |  |  | SANTHOSH | 1900-10-13
Dec 2023 | 2.245% |  |  |  | SOORAJ | 1900-12-28
Feb 2024 | 1.576% |  |  |  | THANISH | 1900-07-25
Jan 2024 | 2.006% |  |  |  | Grand Total | 1905-09-24
Jul 2023 | 1.003%
Jul 2024 | 1.289%
Jun 2023 | 0.5253%
Jun 2024 | 1.242%
Mar 2024 | 2.436%
May 2024 | 2.483%
Nov 2023 | 2.006%
Oct 2023 | 1.958%
Sep 2023 | 1.719%
JERLIN | 20.44%
RAMKIE | 13.18%
SANTHOSH | 13.71%
SOORAJ | 17.34%
THANISH | 9.885%
Grand Total | 100%

## Other Analysis
TYPE VS AVG DOWNTIME
TYPE OF EQUIPMENT | AVERAGE TIME DOWN | Count of STATUS | Column Labels
CRITICAL | 0.023914 | TYPE OF EQUIPMENT | CLOSED CALLS | PENDING CALLS | Grand Total
NON-CRITICAL | 0.029079 | CRITICAL | 517 | 19 | 536
SEMI-CRITICAL | 0.026979 | NON-CRITICAL | 407 | 15 | 422
Grand Total | 0.026618 | SEMI-CRITICAL | 1099 | 37 | 1136
|  | Grand Total | 2023 | 71 | 2094
PEAK CALL TIME ANALYSIS
Row Labels | Count of CALL DATE
12 AM | 41
1 AM | 44
2 AM | 29
3 AM | 21
4 AM | 11
5 AM | 17
6 AM | 39
7 AM | 96
8 AM | 107
9 AM | 150
10 AM | 219
11 AM | 180
12 PM | 134
1 PM | 124
2 PM | 154
3 PM | 165
4 PM | 122
5 PM | 105
6 PM | 53
7 PM | 41
8 PM | 45
9 PM | 62
10 PM | 80
11 PM | 55
Grand Total | 2094

## Dataset
SL.NO | CALL DATE | START MONTH & YEAR | CALL RECEIVED TIME | CALL ATTENDED TIME | RESPONSE TIME (IN MINUTES) | EQUIPMENT NAME | MANUFACTURER | MODEL | ENGINEER | DEPARTMENT | PROBLEM REPORTED | ACTION TAKEN | SPARE USED | PENDING REASON | END DATE | END MONTH & YEAR | END TIME | STARTING DATE & TIME | ENDING DATE & TIME | TIME DOWN | STATUS | TYPE
1 | 2023-06-01 | Jun 2023 | 9:45 | 9:50 | 0.0034722 | DIGITAL RADIOGRAPHIC UNIT | CARESTREAM | DRF ASCEND ANOLOG HF | SOORAJ | RADIOLOGY | COLLIMATION LIGHT NOT WORKING | FUSE GOT DEFECTIVE. CHANGED FUSE AND FOUND TO BE WORKING FINE. | FUSE | NO | 2023-06-01 | June 2023 | 10:00 | 2023-06-01 09:45 | 2023-06-01 10:00 | 0:15 | COMPLETED | SEMI-CRITICAL
2 | 2023-06-01 | Jun 2023 | 13:25 | 13:30 | 0.0034722 | CATHLAB - MONOPLANE | PHILIPS | AZURION 7M12 | SOORAJ | CATHLAB | COLLIMATION SENSOR IS DEFECTIVE MESSAGE ON SCREEN | RESTARTED THE MACHINE AND ISSUE RECTIFIED. | NO | NO | 2023-06-01 | June 2023 | 13:35 | 2023-06-01 13:25 | 2023-06-01 13:35 | 0:10 | COMPLETED | CRITICAL | NO. OF CALLS ATTENDED
3 | 2023-06-01 | Jun 2023 | 17:20 | 17:25 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | EMERGENCY | COT FUNCTIONS NOT WORKING | CONTROL BOX FOUND DEFECTIVE. NEW CONTROL BOX CHANGED AND ISSUE RECTIFIED. | CONTROL BOX | NO | 2023-06-01 | June 2023 | 17:50 | 2023-06-01 17:20 | 2023-06-01 17:50 | 0:30 | COMPLETED | NON-CRITICAL | 2094
4 | 2023-06-01 | Jun 2023 | 17:50 | 17:50 | 0 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | OPD | BATTERY INDICATION RED | FOUND ISSUE WITH CONTROL BOX, CHANGED CONTROL BOX and issue rectified. | CONTROL BOX | NO | 2023-06-01 | June 2023 | 18:20 | 2023-06-01 17:50 | 2023-06-01 18:20 | 0:30 | COMPLETED | NON-CRITICAL | NO. OF CALLS PENDING
5 | 2023-06-01 | Jun 2023 | 18:20 | 18:20 | 0 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | LABORATORY | BRAKED PEDAL MISALIGNED | ALLIGNMENT CHANGED AND RECTIFIED. | NO | NO | 2023-06-01 | June 2023 | 18:50 | 2023-06-01 18:20 | 2023-06-01 18:50 | 0:30 | COMPLETED | NON-CRITICAL | 71
6 | 2023-06-01 | Jun 2023 | 10:20 | 10:25 | 0.0034722 | DEFIBRILLATOR MONITOR | MINDRAY MEDICAL INDIA | BENEHEART D2 | SANTHOSH | OT | INTERNOL PADS HEAD BROKEN | NEED TO REPLACE NEW PADS | INTERNAL PADS | NO | 2023-06-01 | June 2023 | 10:39 | 2023-06-01 10:20 | 2023-06-01 10:39 | 0:19 | COMPLETED | CRITICAL | AVG. RESPONSE TIME (IN MIN)
7 | 2023-06-02 | Jun 2023 | 8:00 | 8:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | SANTHOSH | OT | CIRCUIT LEAK TEST FAILED | CLEANED BELLOWS,CHECKED WITH ANOTHER CIRCUIT | NO | NO | 2023-06-02 | June 2023 | 8:15 | 2023-06-02 08:00 | 2023-06-02 08:15 | 0:15 | COMPLETED | CRITICAL | 0.0046117
8 | 2023-06-03 | Jun 2023 | 13:10 | 13:15 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SOORAJ | ICU | CARTRIDGE EMPTY | REPLACED CARTRIDGE AND ISSUE RECTIFIED. | CARTRIDGE | NO | 2023-06-03 | June 2023 | 13:25 | 2023-06-03 13:10 | 2023-06-03 13:25 | 0:15 | COMPLETED | SEMI-CRITICAL | AVG. RESOLUTION TIME
9 | 2023-06-03 | Jun 2023 | 13:30 | 13:35 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SOORAJ | CATHLAB | NO BATTERY BACKUP ISSUE | POWER SOCKET NOT WORKING. CHANGED POWER SOCKET AND ISSUE RECTIFIED, | POWER SOCKET | NO | 2023-06-03 | June 2023 | 13:45 | 2023-06-03 13:30 | 2023-06-03 13:45 | 0:15 | COMPLETED | CRITICAL | 0.026618
10 | 2023-06-03 | Jun 2023 | 13:20 | 13:25 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SANTHOSH | ICU | NOT WORKING | RESET DONE | NO | NO | 2023-06-03 | June 2023 | 13:25 | 2023-06-03 13:20 | 2023-06-03 13:25 | 0:05 | COMPLETED | NON-CRITICAL
11 | 2023-06-04 | Jun 2023 | 10:00 | 10:05 | 0.0034722 | ECMO MACHINE | GETINGE | ROTOFLOW | SANTHOSH | ICU | NO BATTERY BACKUP | REPLACED NEW BATTERIES | BATTERY | NO | 2023-06-04 | June 2023 | 10:20 | 2023-06-04 10:00 | 2023-06-04 10:20 | 0:20 | COMPLETED | CRITICAL
12 | 2023-06-05 | Jun 2023 | 13:05 | 13:10 | 0.0034722 | ORTHOPANTOMOGRAPHY | SKANRAY DENTAL | MYRAY HYPERION X5 | SOORAJ | OPD | SYSTEM NOT CONNECTING TO MACHINE | CACHE DATA CLEARED AND ISSUE RECTIFIED. | NO | NO | 2023-06-05 | June 2023 | 13:15 | 2023-06-05 13:05 | 2023-06-05 13:15 | 0:10 | COMPLETED | SEMI-CRITICAL
13 | 2023-06-05 | Jun 2023 | 6:10 | 6:20 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S7 | RAMKIE | RADIOLOGY | UNOBLE TO SEND PACS | CHECKED AND RECONNECTED THE ETHERNET CABLE AND SPOOLER CLEANED, PROBLEM RECTIFIED | NO | NO | 2023-06-05 | June 2023 | 6:30 | 2023-06-05 06:10 | 2023-06-05 06:30 | 0:20 | COMPLETED | SEMI-CRITICAL
14 | 2023-06-05 | Jun 2023 | 7:30 | 7:35 | 0.0034722 | BINOCULAR MICROSCOPE | MAGNUS | CH20iLED | RAMKIE | LABORATORY | PROBLEM IN SLIDE MOVEMENT | TIGHTENED THE NUT IN ADJUSTMENT KNOB AND RECTIFIED | NO | NO | 2023-06-05 | June 2023 | 7:45 | 2023-06-05 07:30 | 2023-06-05 07:45 | 0:15 | COMPLETED | SEMI-CRITICAL
15 | 2023-06-06 | Jun 2023 | 7:48 | 7:50 | 0.0013889 | HEART LUNG MACHINE | GETINGE | HL20-4PUMPS | SOORAJ | OT | SHOWING AIR BUBBLE SENSOR ALARM | FOUND ALARM STOPPED WORKING TEMPORAIRLY DUE TO HIGH HUMIDITY. REMOVED AND REFIXED THE SENSOR AFTER 30 MINS. ISSUE RECTIFIED. | NO | NO | 2023-06-06 | June 2023 | 8:30 | 2023-06-06 07:48 | 2023-06-06 08:30 | 0:42 | COMPLETED | CRITICAL
16 | 2023-06-06 | Jun 2023 | 9:23 | 9:25 | 0.0013889 | ELECTRICAL SUCTION | YUWELL | 7A-23B | SOORAJ | OT | POWER NOT SUFFICIENT | CHECKED AND FOUND FILTER ISSUE. CLEANED FILTER AND ISSUE RECTIFIED. | NO | NO | 2023-06-06 | June 2023 | 9:30 | 2023-06-06 09:23 | 2023-06-06 09:30 | 0:07 | COMPLETED | SEMI-CRITICAL
17 | 2023-06-06 | Jun 2023 | 9:50 | 9:55 | 0.0034722 | OT TABLE | MAQUET | MEERA | SOORAJ | OT | REMOTE SHOWING ERROR | CHECEKD AND FOUND PORT NOT WORKING. CONNECTED TO ALTERNATE PORT AND WORKUNG FNE | NO | NO | 2023-06-06 | June 2023 | 10:10 | 2023-06-06 09:50 | 2023-06-06 10:10 | 0:20 | COMPLETED | NON-CRITICAL
18 | 2023-06-06 | Jun 2023 | 11:20 | 11:25 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | RAMKIE | EMERGENCY | NOT WORKING | FOUND POWER CABLE GOT DISCONNECTED, SO FIXED AND RECTIFIED | NO | NO | 2023-06-06 | June 2023 | 11:35 | 2023-06-06 11:20 | 2023-06-06 11:35 | 0:15 | COMPLETED | NON-CRITICAL
19 | 2023-06-07 | Jun 2023 | 8:24 | 8:25 | 0.00069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | SOORAJ | OT | MIXTURE GAS BOARD ISSUE. | MIXTURE GAS BOARD REPLACED AND ISSUE RECTIFIED. | MIXTURE GAS BOARD | NO | 2023-06-07 | June 2023 | 9:00 | 2023-06-07 08:24 | 2023-06-07 09:00 | 0:36 | COMPLETED | CRITICAL
20 | 2023-06-07 | Jun 2023 | 9:00 | 9:05 | 0.0034722 | HEART LUNG MACHINE | GETINGE | HL20-4PUMPS | SOORAJ | OT | ER 20 ERROR ON DISPLAY | FOUND ISSUE WITH SENSOR BAORD DUE TO HUMIDITY. SENSOR BOARD TAKEN OUT AND REFIXED AFTER 30 MINS. ISSU ERECTIFIED | NO | NO | 2023-06-07 | June 2023 | 9:30 | 2023-06-07 09:00 | 2023-06-07 09:30 | 0:30 | COMPLETED | CRITICAL
21 | 2023-06-07 | Jun 2023 | 13:40 | 13:45 | 0.0034722 | CAMERA SYSTEM | KARL STORZ | TC200 | RAMKIE | OPD | LAP UNIT RECOEDING GOT STRUCK AND SHOWS LOADING | CABLES RECONNECTED AND RESTARTED, NOW IT IS WORKING GOOD | NO | NO | 2023-06-07 | June 2023 | 13:55 | 2023-06-07 13:40 | 2023-06-07 13:55 | 0:15 | COMPLETED | SEMI-CRITICAL
22 | 2023-06-07 | Jun 2023 | 15:00 | 15:10 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | WARD | BED NOT WORKING | CHECKED THE BED CONDITION WITH THE SERVICE PERSON. REPLACED TWO HEIGHT ACTUATOR AND A CONTROL BOX. | NO | NO | 2023-06-07 | June 2023 | 15:30 | 2023-06-07 15:00 | 2023-06-07 15:30 | 0:30 | COMPLETED | NON-CRITICAL
23 | 2023-06-08 | Jun 2023 | 9:40 | 9:45 | 0.0034722 | HEART LUNG MACHINE | MAQUET | HL20-4PUMPS | SOORAJ | ICU | BURNT SMELL ON HLT MACHINE | SUSPECT SMELL DUE TO NEGLIGIBLE SHORT CIRCUIT. | NO | NO | 2023-06-08 | June 2023 | 10:15 | 2023-06-08 09:40 | 2023-06-08 10:15 | 0:35 | COMPLETED | CRITICAL
24 | 2023-06-08 | Jun 2023 | 14:30 | 14:35 | 0.0034722 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | RAMKIE | WARD | LOCKER GOT STRUCK | SERVICED AND RECTIFIED | NO | NO | 2023-06-08 | June 2023 | 14:45 | 2023-06-08 14:30 | 2023-06-08 14:45 | 0:15 | COMPLETED | NON-CRITICAL
25 | 2023-06-08 | Jun 2023 | 16:00 | 16:10 | 0.0069444 | OT TABLE | VM SURGICALS | ENTERPRISE 9000 | HAFIZ | EMERGENCY | REMOTE NOT WORKING PROPERLY AT EMERGENCY OT. | ACCOMPANIED WITH THE SERVICE PERSON FOR OT BED REMOTE. TIGHTENED THE SCREW OF THE BED. MADE SOME SMALL ADJUSTMENT. SOLVED THE PROBLEM. | NO | NO | 2023-06-08 | June 2023 | 16:30 | 2023-06-08 16:00 | 2023-06-08 16:30 | 0:30 | COMPLETED | NON-CRITICAL
26 | 2023-06-11 | Jun 2023 | 11:30 | 11:40 | 0.0069444 | PATIENT MONITOR | MINDRAY | ePM 12c | HAFIZ | ICU | ISSUE WITH THE CENTRAL MONITORING SYSTEM | REPLACED FEW LAN CABLE AND RECTIFIED THE PROBLEM WITH THE SERVICE PERSON. | LAN CABLE | NO | 2023-06-11 | June 2023 | 12:30 | 2023-06-11 11:30 | 2023-06-11 12:30 | 1:00 | COMPLETED | SEMI-CRITICAL
27 | 2023-06-12 | Jun 2023 | 10:50 | 10:55 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | RAMKIE | ICU | 02 TEST FAILED | CHECKED AND FOUND THE PROBLEM 02 SENSOR, NEED TO REPLACE NEW ONE | O2 SENSOR | NO | 2023-06-12 | June 2023 | 11:05 | 2023-06-12 10:50 | 2023-06-12 11:05 | 0:15 | COMPLETED | CRITICAL
28 | 2023-06-12 | Jun 2023 | 10:30 | 10:40 | 0.0069444 | VIDEO PROCESSING SYSTEM, ENDOSCOPE | OLYMPUS | CV-190 | HAFIZ | OPD | PROBLEM WHILE CAPTURING THE FREEZED IMAGE FROM THE WINDOW. | CONTACTED THE COMPANY PERSON AND UPDATED THE SOFTWARE. PROBLEM SOLVED. | NO | NO | 2023-06-12 | June 2023 | 11:00 | 2023-06-12 10:30 | 2023-06-12 11:00 | 0:30 | COMPLETED | SEMI-CRITICAL
29 | 2023-06-12 | Jun 2023 | 19:30 | 19:35 | 0.0034722 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | SANTHOSH | ICU | AIR IN LINE | RECONNECTED AND CHANGED MODE | NO | NO | 2023-06-12 | June 2023 | 19:40 | 2023-06-12 19:30 | 2023-06-12 19:40 | 0:10 | COMPLETED | NON-CRITICAL
30 | 2023-06-13 | Jun 2023 | 0:08 | 0:13 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SOORAJ | WARD | COT FUNCTIONS LCKED | RESET DONE AND ISSUE RECTIFIED. | NO | NO | 2023-06-13 | June 2023 | 0:15 | 2023-06-13 00:08 | 2023-06-13 00:15 | 0:07 | COMPLETED | NON-CRITICAL
31 | 2023-06-13 | Jun 2023 | 0:45 | 0:50 | 0.0034722 | VENTILATOR, UNIVERSAL | MAQUET | SERVO-I | SOORAJ | ICU | SET O2 CONCENTRATION NOT ATTAING | CHANGED TO ANOTHER PORT AND FOUND SET O2 WAS ATTAINING. ISSUE RECTIFED. | NO | NO | 2023-06-13 | June 2023 | 0:55 | 2023-06-13 00:45 | 2023-06-13 00:55 | 0:10 | COMPLETED | CRITICAL
32 | 2023-06-13 | Jun 2023 | 14:50 | 14:56 | 0.0041667 | HOLTER,RECORDER | WIPRO GE HEALTH CARE | SEER 1000 | SANTHOSH | OPD | RED INDICATION | RESET DONE | NO | NO | 2023-06-13 | June 2023 | 15:02 | 2023-06-13 14:50 | 2023-06-13 15:02 | 0:12 | COMPLETED | SEMI-CRITICAL
33 | 2023-06-13 | Jun 2023 | 18:00 | 18:05 | 0.0034722 | OT TABLE | MAQUET | ALPHA CLASSIC PRO | SANTHOSH | OT | REMOTE NOT WORKING | REFIXED REMOTE AND CHECKED | NO | NO | 2023-06-13 | June 2023 | 18:15 | 2023-06-13 18:00 | 2023-06-13 18:15 | 0:15 | COMPLETED | NON-CRITICAL
34 | 2023-06-14 | Jun 2023 | 23:30 | 23:35 | 0.0034722 | STEAM STERILIZER | GETINGE | GSS67H102E | SOORAJ | CSSD | START FUNCTION DISABLED | PRESSURE VALVE PARTIALLY CLOSED. OPENED VALVE AND ISSUE RECTIFIED. | NO | NO | 2023-06-14 | June 2023 | 23:45 | 2023-06-14 23:30 | 2023-06-14 23:45 | 0:15 | COMPLETED | SEMI-CRITICAL
35 | 2023-06-14 | Jun 2023 | 18:50 | 19:00 | 0.0069444 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | RAMKIE | OT | GOT STRUCK | RESTARTED AND RECTIFIED | NO | NO | 2023-06-14 | June 2023 | 19:05 | 2023-06-14 18:50 | 2023-06-14 19:05 | 0:15 | COMPLETED | SEMI-CRITICAL
36 | 2023-06-14 | Jun 2023 | 17:00 | 17:10 | 0.0069444 | ARTHROSCOPY | SMITH & NEPHEW | 72204354 | HAFIZ | OT | REPORTED THAT VIDEO GET DISTORTED DURING PROCEDURE. | INSPECTED LENS CAMERA SYSTEM & FOUND CAMERA CONNECTOR PINS HAVING DIRTY AND DUSTS. CLEARED THE HEAD CABLE CONNECTOR AND CHECKED. PROBLEM RECTIFIED. | NO | NO | 2023-06-14 | June 2023 | 17:45 | 2023-06-14 17:00 | 2023-06-14 17:45 | 0:45 | COMPLETED | SEMI-CRITICAL
37 | 2023-06-14 | Jun 2023 | 14:10 | 14:15 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | SANTHOSH | ICU | O2 CELL FAILURE | NEW O2 CELL REPLACED | O2 CELL | NO | 2023-06-14 | June 2023 | 14:25 | 2023-06-14 14:10 | 2023-06-14 14:25 | 0:15 | COMPLETED | CRITICAL
38 | 2023-06-14 | Jun 2023 | 17:00 | 17:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | SANTHOSH | ICU | DIDNT TURN ON | PROBLEM WITH BOARD | CONTROLLER BOARD REPLACED | NO | 2023-06-14 | June 2023 | 17:15 | 2023-06-14 17:00 | 2023-06-14 17:15 | 0:15 | COMPLETED | SEMI-CRITICAL
39 | 2023-06-15 | Jun 2023 | 0:15 | 0:20 | 0.0034722 | FLUID WARMER SYSTEM | MATRIX | MN05 | SOORAJ | OT | warmer not warming up | WATER IN TANK CHANGED AND ISSUE RECTIFIED. | NO | NO | 2023-06-15 | June 2023 | 0:35 | 2023-06-15 00:15 | 2023-06-15 00:35 | 0:20 | COMPLETED | SEMI-CRITICAL
40 | 2023-06-15 | Jun 2023 | 0:40 | 0:45 | 0.0034722 | VENTILATOR, UNIVERSAL | MAQUET | SERVO-I | SOORAJ | ICU | CHANGE TRIGGER VALUE ALARM | CHANGED THE SETTINGS TO ADULT | NO | NO | 2023-06-15 | June 2023 | 0:50 | 2023-06-15 00:40 | 2023-06-15 00:50 | 0:10 | COMPLETED | CRITICAL
41 | 2023-06-15 | Jun 2023 | 21:50 | 21:55 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SOORAJ | WARD | COT NOT SWITCHING ON | FOUND POWER CORD DISCONNECTED. RECONNECTED AND ISSUE RECTIFIED. | NO | NO | 2023-06-15 | June 2023 | 22:00 | 2023-06-15 21:50 | 2023-06-15 22:00 | 0:10 | COMPLETED | NON-CRITICAL
42 | 2023-06-15 | Jun 2023 | 19:00 | 19:05 | 0.0034722 | DIALYSIS MACHINE | FRESENIUS | 4008 S | RAMKIE | OPD | BLD ALARM | BLD CALIBRATED AND CALIBRATED | NO | NO | 2023-06-15 | June 2023 | 19:10 | 2023-06-15 19:00 | 2023-06-15 19:10 | 0:10 | COMPLETED | CRITICAL
43 | 2023-06-15 | Jun 2023 | 16:10 | 16:15 | 0.0034722 | VENTILATOR, ADULT | MAQUET | SERVO-I | SANTHOSH | ICU | PLUG TOP LOOSENED | CUT DOWN THE TEARED WIRE AND FIXED TIGHTENED | NO | NO | 2023-06-15 | June 2023 | 16:15 | 2023-06-15 16:10 | 2023-06-15 16:15 | 0:05 | COMPLETED | CRITICAL
44 | 2023-06-16 | Jun 2023 | 18:10 | 18:15 | 0.0034722 | VENTILATOR, TRANSPORT | RESMED | ASTRAL 100 | SANTHOSH | ICU | NV MASK REBREATHING ALARM | CIRCUIT CALIBRATION DONE | NO | NO | 2023-06-16 | June 2023 | 18:25 | 2023-06-16 18:10 | 2023-06-16 18:25 | 0:15 | COMPLETED | CRITICAL
45 | 2023-06-17 | Jun 2023 | 23:23 | 23:25 | 0.0013889 | PATIENT MONITOR | PHILIPS | MX 550 | SOORAJ | OT | MONITOR MAKING NOISE AND DISPLAY NOT SHOWING | BATTERY GONE DEFECTIVE. REMOVED BATTERY AND CONNECTED ON POWER SUPPLY. FOUND WORKING PROPERLY. AS OF NOW MONITOR IS WORKING ON POWER SUPPLY. | NO | NO | 2023-06-17 | June 2023 | 23:40 | 2023-06-17 23:23 | 2023-06-17 23:40 | 0:17 | COMPLETED | SEMI-CRITICAL
47 | 2023-06-17 | Jun 2023 | 15:10 | 15:15 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | SANTHOSH | ICU | DOESNT PUSH IMAGE TO PACS | REFRESHED SPOOLER | NO | NO | 2023-06-17 | June 2023 | 15:25 | 2023-06-17 15:10 | 2023-06-17 15:25 | 0:15 | COMPLETED | SEMI-CRITICAL
48 | 2023-06-18 | Jun 2023 | 10:30 | 10:35 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | RAMKIE | RADIOLOGY | ECHO SCAN QUALITY IS NOT GOOD | CHECKED AND REGISTERED CALL AND COMPANY PERSON SUGGESTED TO CHANGE THE PROBE C1-6-D | NO | C1-6-D PROBE | 2023-06-18 | June 2023 | 10:45 | 2023-06-18 10:30 | 2023-06-18 10:45 | 0:15 | COMPLETED | SEMI-CRITICAL
49 | 2023-06-20 | Jun 2023 | 14:20 | 14:30 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | RAMKIE | WARD | NOT WORKING | RESET DONE AND PROBLEM RECTIFIED | NO | NO | 2023-06-20 | June 2023 | 14:45 | 2023-06-20 14:20 | 2023-06-20 14:45 | 0:25 | COMPLETED | NON-CRITICAL
50 | 2023-06-21 | Jun 2023 | 14:40 | 14:45 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | SOORAJ | ICU | DISPLAY NOT SWITHCING ON | IV2 FLEX BOARD CHANGED AND ISSUE RECTIFIED. | IV2 FLEX BOARD | NO | 2023-06-21 | June 2023 | 15:00 | 2023-06-21 14:40 | 2023-06-21 15:00 | 0:20 | COMPLETED | SEMI-CRITICAL
51 | 2023-06-21 | Jun 2023 | 15:00 | 15:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SOORAJ | EMERGENCY | PRINT OUT NOT WORKING. FOUND THE RECOMMENDED PAPER NOT BEING USED. INFORMED END USER | RECOMMENDED PAPER NOT USED BY USER. AFTER USING RECOMMENDED PAPER, PRINTING WORKS. | NO | NO | 2023-06-21 | June 2023 | 15:20 | 2023-06-21 15:00 | 2023-06-21 15:20 | 0:20 | COMPLETED | CRITICAL
52 | 2023-06-21 | Jun 2023 | 15:45 | 15:50 | 0.0034722 | IVUS MACHINE | PHILIPS | CORE SERIES S5 | SOORAJ | CATHLAB | SERVICE REQUIRED MESSAGE SHOWING. | SOFTWARE UPDATE DONE AND ISSUE RECTIFIED. | NO | NO | 2023-06-21 | June 2023 | 16:00 | 2023-06-21 15:45 | 2023-06-21 16:00 | 0:15 | COMPLETED | SEMI-CRITICAL
53 | 2023-06-21 | Jun 2023 | 17:20 | 17:30 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | EMERGENCY | REPORTED THAT PRINTER WAS NOT WORKING IN THE DEFIBRILLATOR. | ACCOMPANIED WITH THE SERVICE PERSON AND HE SUGGESTED TO USE WHITE PAPER INSTEAD OF RED AND GREEN PAPER FOR PRINTING. IT IS RECTIFIED. | NO | NO | 2023-06-21 | June 2023 | 17:40 | 2023-06-21 17:20 | 2023-06-21 17:40 | 0:20 | COMPLETED | CRITICAL
54 | 2023-06-21 | Jun 2023 | 11:50 | 11:55 | 0.0034722 | ORTHOPANTOMOGRAPHY | SKANRAY DENTAL | MYRAY HYPERION X5 | SANTHOSH | OPD | PATIENT IMAGE CANT RETRIEVAL | CONFIGURED IP CHECKED CONNECTIONS | NO | NO | 2023-06-21 | June 2023 | 12:05 | 2023-06-21 11:50 | 2023-06-21 12:05 | 0:15 | COMPLETED | SEMI-CRITICAL
55 | 2023-06-21 | Jun 2023 | 13:10 | 13:15 | 0.0034722 | INTRA ORAL XRAY | SKANRAY DENTAL | INTRA SKANDC | SANTHOSH | OPD | REMOTE NOT WORKING | NEW REMOTE REPLACED | REMOTE DISPLAY | NO | 2023-06-21 | June 2023 | 13:20 | 2023-06-21 13:10 | 2023-06-21 13:20 | 0:10 | COMPLETED | SEMI-CRITICAL
56 | 2023-06-21 | Jun 2023 | 13:30 | 13:35 | 0.0034722 | INTRA ORAL XRAY | SKANRAY DENTAL | INTRA SKANDC | SANTHOSH | OPD | SENSOR DISCONNECTED | CHANGED ANOTHER PORT | NO | NO | 2023-06-21 | June 2023 | 13:40 | 2023-06-21 13:30 | 2023-06-21 13:40 | 0:10 | COMPLETED | SEMI-CRITICAL
57 | 2023-06-23 | Jun 2023 | 12:50 | 13:00 | 0.0069444 | PRESSURE INJECTOR | LF OPTI VONTAGE | ILLUMENO NEO | HAFIZ | CATHLAB | REPORTED AN ISSUE IN THE PLUNGER OF THE INJECTOR MACHINE | THE PROBLEM WAS WITH BOARD. IT GOT DAMAGED. INFORMED THE SERVICE PERSON. | NO | NO | 2023-06-23 | June 2023 | 13:30 | 2023-06-23 12:50 | 2023-06-23 13:30 | 0:40 | COMPLETED | SEMI-CRITICAL
58 | 2023-06-23 | Jun 2023 | 13:10 | 13:15 | 0.0034722 | PRESSURE INJECTOR | LF OPTI VONTAGE | ILLUMENO NEO | SANTHOSH | CATHLAB | ERROR 0004 | NEED TO REPLACE BOARD | MAIN BOARD | NO | 2023-06-23 | June 2023 | 13:25 | 2023-06-23 13:10 | 2023-06-23 13:25 | 0:15 | COMPLETED | SEMI-CRITICAL
59 | 2023-06-24 | Jun 2023 | 12:00 | 12:10 | 0.0069444 | PRESSURE INJECTOR | LF OPTI VONTAGE | ILLUMENO NEO | HAFIZ | CATHLAB | REPORTED AN ISSUE IN THE PLUNGER OF THE INJECTOR MACHINE | ACCOMPANIED WITH THE SERVICE PERSON. SWAPPED THE CONTROL BOX FROM FIXED TO PORTABLE INJECTOR MACHINE. NOW, THE PORTABLE INJECTOR IS WORKING FINE, AND THE FIXED INJECTOR IS DOWN. | CONTROL BOX | NO | 2023-06-24 | June 2023 | 15:00 | 2023-06-24 12:00 | 2023-06-24 15:00 | 3:00 | COMPLETED | SEMI-CRITICAL
60 | 2023-06-25 | Jun 2023 | 15:30 | 15:35 | 0.0034722 | PRESSURE INJECTOR | LF OPTI VONTAGE | ILLUMENO NEO | JERLIN | CATHLAB | NOT WORKING | FOUND ISSUE WITH MAIN BOARD & INTERFACING. MAIN BOARD & INTERFACING BOARD REPLACED AND ISSUE RECTIFIED. | MAIN BOARD AND INTERFACING BOARD | NO | 2023-06-25 | June 2023 | 16:30 | 2023-06-25 15:30 | 2023-06-25 16:30 | 1:00 | COMPLETED | SEMI-CRITICAL
61 | 2023-06-25 | Jun 2023 | 10:00 | 10:05 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 9000 | SANTHOSH | WARD | NOT WORKING | RESET DONE | NO | NO | 2023-06-25 | June 2023 | 10:20 | 2023-06-25 10:00 | 2023-06-25 10:20 | 0:20 | COMPLETED | NON-CRITICAL
62 | 2023-06-26 | Jun 2023 | 11:55 | 12:00 | 0.0034722 | EXAMINATION COUCH | MIDMARK JANAK | MX400A | SOORAJ | DIALYSIS | CHAIR NOT WORKING | TRIED RESETTING AND STILL ISSUE PERSIST. SUSPECTING CONTROL BOX. | NO | CONTRL BOX | 2023-06-26 | June 2023 | 12:15 | 2023-06-26 11:55 | 2023-06-26 12:15 | 0:20 | PENDING | NON-CRITICAL
63 | 2023-06-26 | Jun 2023 | 1:20 | 1:25 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | RAMKIE | WARD | ECG MALFUNCTION | REFIXED ECG CABLE AND OPERATIONAL CHECK PASSED, READY TO USE | NO | NO | 2023-06-26 | June 2023 | 1:35 | 2023-06-26 01:20 | 2023-06-26 01:35 | 0:15 | COMPLETED | CRITICAL
64 | 2023-06-26 | Jun 2023 | 9:10 | 9:20 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | ICU | REPORTED AN ISSUE IN THE DEFIBRILLATOR (I.E., ECG LEAD TEST FAILED). | ECG CABLE REMOVED AND INSERTED PROPERLY. RESTARTED THE DEVICE AND PASSED THE TEST. | NO | NO | 2023-06-26 | June 2023 | 9:25 | 2023-06-26 09:10 | 2023-06-26 09:25 | 0:15 | COMPLETED | CRITICAL
65 | 2023-06-27 | Jun 2023 | 18:40 | 18:45 | 0.0034722 | VENTILATOR, ADULT | MAQUET | SERVO-I | HAFIZ | ICU | INTERNAL LEAKGE ERROR IN SERVO-I VENTILATOR | REPLACED A NEW CASSETTE FROM OUR DEPARTMENT. STILL THE ERROR PERSISTS. INFORMED THE COMPANY PERSON. | EXPIRATORY CASSETTE | NO | 2023-06-27 | June 2023 | 19:00 | 2023-06-27 18:40 | 2023-06-27 19:00 | 0:20 | COMPLETED | CRITICAL
66 | 2023-06-28 | Jun 2023 | 14:30 | 14:45 | 0.010417 | PRESSURE INJECTOR | LF OPTI VONTAGE | ILLUMENO NEO | JERLIN | CATHLAB | NOT WORKING | FOUND ISSUE WITH INTERFACING CABLE'S CONNECTOR. CONNECTOR REPLACED AND ISSUE RECTIFIED. | CABLE CONNECTOR | NO | 2023-06-28 | June 2023 | 16:00 | 2023-06-28 14:30 | 2023-06-28 16:00 | 1:30 | COMPLETED | SEMI-CRITICAL
67 | 2023-06-28 | Jun 2023 | 11:30 | 11:45 | 0.010417 | VENTILATOR, ADULT | MAQUET | SERVO-I | JERLIN | ICU | INTERNAL LEAKAGE TEST FAILED | CHECKED AND RECTIFIED THE ISSUE | NO | NO | 2023-06-28 | June 2023 | 11:45 | 2023-06-28 11:30 | 2023-06-28 11:45 | 0:15 | COMPLETED | CRITICAL
68 | 2023-06-28 | Jun 2023 | 21:30 | 21:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | ICU | NOT WORKING | RESET DONE AND PROBLEM RECTIFIED | NO | NO | 2023-06-28 | June 2023 | 21:45 | 2023-06-28 21:30 | 2023-06-28 21:45 | 0:15 | COMPLETED | NON-CRITICAL
69 | 2023-06-28 | Jun 2023 | 2:30 | 2:35 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | RAMKIE | EMERGENCY | ECG MALFUNCTION | REFIXED ECG CABLE AND OPERATIONAL CHECK PASSED, READY TO USE | NO | NO | 2023-06-28 | June 2023 | 2:45 | 2023-06-28 02:30 | 2023-06-28 02:45 | 0:15 | COMPLETED | CRITICAL
70 | 2023-06-28 | Jun 2023 | 13:25 | 13:30 | 0.0034722 | BABY WEIGHING MACHINE | SECA | 354 | HAFIZ | ICU | BABY WEIGHING SCALE NOT WORKING. | CHECKED AND FIXED THE ISSUE BY RESTARTING THE DEVICE. | NO | NO | 2023-06-28 | June 2023 | 13:45 | 2023-06-28 13:25 | 2023-06-28 13:45 | 0:20 | COMPLETED | NON-CRITICAL
71 | 2023-06-29 | Jun 2023 | 11:45 | 12:00 | 0.010417 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | JERLIN | ICU | TO CHECK O2 PRESSURE | CHECKED AND FOUND CORRECT | NO | NO | 2023-06-29 | June 2023 | 12:10 | 2023-06-29 11:45 | 2023-06-29 12:10 | 0:25 | COMPLETED | CRITICAL
72 | 2023-06-29 | Jun 2023 | 14:00 | 14:10 | 0.0069444 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | JERLIN | OPD | COMMUNICATION PROBLEM | CHECKED AND RECTIFIED THE ISSUE | NO | NO | 2023-06-29 | June 2023 | 14:25 | 2023-06-29 14:00 | 2023-06-29 14:25 | 0:25 | COMPLETED | SEMI-CRITICAL
73 | 2023-06-29 | Jun 2023 | 4:30 | 4:45 | 0.010417 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | RAMKIE | WARD | MALFUNTION | CHECKED AND NO ISSUES FOUND | NO | NO | 2023-06-29 | June 2023 | 14:55 | 2023-06-29 04:30 | 2023-06-29 14:55 | 10:25 | COMPLETED | NON-CRITICAL
74 | 2023-06-30 | Jun 2023 | 11:00 | 11:10 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 450 | JERLIN | ICU | NIBP NOT WORKING | FOUND PROBLEM WITH THE CONNECTOR . REPLACED THE NIBP HOSE | NIBP HOSE | NO | 2023-06-30 | June 2023 | 11:25 | 2023-06-30 11:00 | 2023-06-30 11:25 | 0:25 | COMPLETED | SEMI-CRITICAL
75 | 2023-06-30 | Jun 2023 | 22:30 | 22:35 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | RAMKIE | ICU | NOT WORKING | RESET DONE AND PROBLEM RECTIFIED | NO | NO | 2023-06-30 | June 2023 | 22:45 | 2023-06-30 22:30 | 2023-06-30 22:45 | 0:15 | COMPLETED | SEMI-CRITICAL
76 | 2023-06-30 | Jun 2023 | 23:25 | 23:30 | 0.0034722 | NITRIC OXIDE DELIVERY SYSTEM | WIPRO GE HEALTH CARE | SLE | RAMKIE | ICU | FLOW FLUCTUATION | PRESSURE ADJUSTED AND IT WORKING | NO | NO | 2023-06-30 | June 2023 | 23:35 | 2023-06-30 23:25 | 2023-06-30 23:35 | 0:10 | COMPLETED | CRITICAL
77 | 2023-06-30 | Jun 2023 | 21:45 | 21:55 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | ICU | NOT WORKING | RESET DONE AND PROBLEM RECTIFIED | NO | NO | 2023-06-30 | June 2023 | 22:10 | 2023-06-30 21:45 | 2023-06-30 22:10 | 0:25 | COMPLETED | NON-CRITICAL
78 | 2023-07-01 | Jul 2023 | 8:00 | 8:05 | 0.0034722 | OT TABLE | MAQUET | MEERA | SOORAJ | OT | TABLE NOT WORKING | CHECKED AND FOUND POWER CHORD DEFECTIVE. REPLACED ANOTHER POWER CHORD AND ISSUE RECTIFIED. | POWERCORD | NO | 2023-07-01 | July 2023 | 8:10 | 2023-07-01 08:00 | 2023-07-01 08:10 | 0:10 | COMPLETED | NON-CRITICAL
79 | 2023-07-01 | Jul 2023 | 6:10 | 6:25 | 0.010417 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | RAMKIE | CATHLAB | NOT WORKING | CHECKED AND RESET DONE, NOW IT WORKING | NO | NO | 2023-07-01 | July 2023 | 6:20 | 2023-07-01 06:10 | 2023-07-01 06:20 | 0:10 | COMPLETED | SEMI-CRITICAL
80 | 2023-07-01 | Jul 2023 | 10:00 | 10:15 | 0.010417 | MAMMOGRAPH UNIT | HOLOGIC | SELENIA DIMENSIONS | JERLIN | RADIOLOGY | NOT SWITCHING ON | REFIXED THE POWER CORD AND RECTIFIED THE ISSUE. | NO | NO | 2023-07-01 | July 2023 | 10:25 | 2023-07-01 10:00 | 2023-07-01 10:25 | 0:25 | COMPLETED | SEMI-CRITICAL
81 | 2023-07-01 | Jul 2023 | 13:15 | 13:25 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | WARD | DISCONNECTED POWER CORD | REFIXED THE POWER CORD AND RECTIFIED THE ISSUE. | NO | NO | 2023-07-01 | July 2023 | 13:40 | 2023-07-01 13:15 | 2023-07-01 13:40 | 0:25 | COMPLETED | NON-CRITICAL
82 | 2023-07-01 | Jul 2023 | 19:00 | 19:05 | 0.0034722 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | JERLIN | ICU | NOT SWITCHING ON | FIXED WITH APOWER CORD AND SWITCHED ON. RECTIFIED THE ISSUE. | NO | NO | 2023-07-01 | July 2023 | 19:10 | 2023-07-01 19:00 | 2023-07-01 19:10 | 0:10 | COMPLETED | NON-CRITICAL
83 | 2023-07-01 | Jul 2023 | 17:30 | 17:35 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | HAFIZ | ICU | PRE-USE CHECK TEST FAILED. | RESTARTED THE VENTILATOR, RERUN THE TEST. NOW, WORKING GOOD. | NO | NO | 2023-07-01 | July 2023 | 17:45 | 2023-07-01 17:30 | 2023-07-01 17:45 | 0:15 | COMPLETED | CRITICAL
84 | 2023-07-02 | Jul 2023 | 11:55 | 12:00 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | SOORAJ | ICU | SLAVE MONITOR NOT WORKING. | FOUND CABLE VGA CABLE DISCONNECTED. RECONNECTED AND ISSUE RECTIFIED. | NO | NO | 2023-07-02 | July 2023 | 12:05 | 2023-07-02 11:55 | 2023-07-02 12:05 | 0:10 | COMPLETED | SEMI-CRITICAL
85 | 2023-07-02 | Jul 2023 | 12:55 | 13:00 | 0.0034722 | BLOOD CULTURE DETECTION SYSTEM | BACT ALERT 3D | BIOMERIUX | SOORAJ | LABORATORY | NOT SWITCHING ON | POWER CABLE DISCONNECTED. RECONNECTED AND ISSUE RECTIFIED. | NO | NO | 2023-07-02 | July 2023 | 13:05 | 2023-07-02 12:55 | 2023-07-02 13:05 | 0:10 | COMPLETED | SEMI-CRITICAL
86 | 2023-07-02 | Jul 2023 | 5:20 | 5:25 | 0.0034722 | VENTILATOR, ADULT | MAQUET | SERVO-I | RAMKIE | ICU | PRE USE CHECK TEST FAILED | EXPIRATORY CASSETTE DIAPRAHM CLEAND AND PROBLEM RECTIFIED | NO | NO | 2023-07-02 | July 2023 | 5:35 | 2023-07-02 05:20 | 2023-07-02 05:35 | 0:15 | COMPLETED | CRITICAL
87 | 2023-07-03 | Jul 2023 | 9:18 | 9:20 | 0.0013889 | PATIENT MONITOR | PHILIPS | MX 550 | SOORAJ | OT | ETCO2 NOT SHWOING | WATER TRAP DEFECTIVE. CHANGED WATER TRAP AND ISSUE RECTIFIED. | WATER TRAP | NO | 2023-07-03 | July 2023 | 9:30 | 2023-07-03 09:18 | 2023-07-03 09:30 | 0:12 | COMPLETED | SEMI-CRITICAL
88 | 2023-07-03 | Jul 2023 | 14:10 | 14:20 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 450 | JERLIN | EMERGENCY | NIBP NOT FUNCTIONING PROPERLY | FOUND LEAKAGE IN BP CUFF. REPLACED WITH NEW PEDIATRIC BP CUFF AND RECTIFIED THE ISSUE. | PEDIATRIC BP CUFF | NO | 2023-07-03 | July 2023 | 14:25 | 2023-07-03 14:10 | 2023-07-03 14:25 | 0:15 | COMPLETED | SEMI-CRITICAL
89 | 2023-07-03 | Jul 2023 | 19:30 | 19:35 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SANTHOSH | ICU | AUDIO TEST FAILED WHILE OPERATIONAL CHECK | RERUN OPERATIONAL CHECK,NOW ITS WORKING | NO | NO | 2023-07-03 | July 2023 | 19:40 | 2023-07-03 19:30 | 2023-07-03 19:40 | 0:10 | COMPLETED | CRITICAL
90 | 2023-07-03 | Jul 2023 | 10:20 | 10:22 | 0.0013889 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | ICU | OCCLUSION ERROR | RESTARTED THE INFUSION PUMP, CHANGED THE OCCLUSION LEVEL. NOW, WORKING GOOD. | NO | NO | 2023-07-03 | July 2023 | 10:28 | 2023-07-03 10:20 | 2023-07-03 10:28 | 0:08 | COMPLETED | NON-CRITICAL
91 | 2023-07-03 | Jul 2023 | 10:30 | 10:35 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | HAFIZ | ICU | PRE-USE CHECK FAILED. EXPIRATORY CASSETTE ERROR. | REMOVED THE EXPIRATORY CASSETTE, CLEANED AND REFIXED INTO THE VENTILATOR. RERUN THE PRE-USE CHECK AND PASSED. NOW, WORKING FINE. | NO | NO | 2023-07-03 | July 2023 | 10:45 | 2023-07-03 10:30 | 2023-07-03 10:45 | 0:15 | COMPLETED | CRITICAL
92 | 2023-07-04 | Jul 2023 | 10:30 | 10:35 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID E95 | SOORAJ | OPD | 4V PROBE NOT DETECTING | FOUND ISSUE WITH 4V PROBE AND NEEDS REPLACEMENT. | NO | TEE PROBE | 2023-07-04 | July 2023 | 10:45 | 2023-07-04 10:30 | 2023-07-04 10:45 | 0:15 | PENDING | SEMI-CRITICAL
93 | 2023-07-04 | Jul 2023 | 15:20 | 15:25 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | RAMKIE | RADIOLOGY | ABDOMAN PROBE NOT WORKING | CHANGED PROBE TO ANOTHER PORT AND NOW IT WORKING | NO | NO | 2023-07-04 | July 2023 | 15:35 | 2023-07-04 15:20 | 2023-07-04 15:35 | 0:15 | COMPLETED | SEMI-CRITICAL
94 | 2023-07-04 | Jul 2023 | 13:10 | 13:15 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | HAFIZ | ICU | ABG MACHINE NOT WORKING. CLOT SAMPLE RUNNED. | RUN THE HEPARIN TEST 4-5 TIMES, SAME RESULTS. A NEW CATRIDGE IS REPLACED AND CVP TEST DONE. | CARTRIDGE | NO | 2023-07-04 | July 2023 | 13:25 | 2023-07-04 13:10 | 2023-07-04 13:25 | 0:15 | COMPLETED | SEMI-CRITICAL
95 | 2023-07-05 | Jul 2023 | 9:15 | 9:20 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S7 | SOORAJ | RADIOLOGY | NOT SHOWING SCAN PAGE | DISCONNECTED ALL PROBES AND RESTARTED THE SYSTEM. ISSUE RECTIFIED. | NO | NO | 2023-07-05 | July 2023 | 9:27 | 2023-07-05 09:15 | 2023-07-05 09:27 | 0:12 | COMPLETED | SEMI-CRITICAL
96 | 2023-07-05 | Jul 2023 | 12:40 | 12:45 | 0.0034722 | OT TABLE | MAQUET | MEERA | SOORAJ | OT | BATTERY INDICATION RED | BATTERY DEPLETED. KEPT ON CHARGE AND ISSUE WAS RECTIFIED. | NO | NO | 2023-07-05 | July 2023 | 13:00 | 2023-07-05 12:40 | 2023-07-05 13:00 | 0:20 | COMPLETED | NON-CRITICAL
97 | 2023-07-05 | Jul 2023 | 16:30 | 16:45 | 0.010417 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | RAMKIE | OT | OCCLUSION ERROR | USED FOR FEEDING, AFTER DILUTION IT WORKING GOOD | NO | NO | 2023-07-05 | July 2023 | 16:55 | 2023-07-05 16:30 | 2023-07-05 16:55 | 0:25 | COMPLETED | NON-CRITICAL
98 | 2023-07-05 | Jul 2023 | 8:25 | 8:30 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SANTHOSH | ICU | NOT WORKING | CHECKED CONTROLLER BOX ITS DEFECT | NO | CONTROLLER BOX | 2023-07-05 | July 2023 | 8:30 | 2023-07-05 08:25 | 2023-07-05 08:30 | 0:05 | PENDING | NON-CRITICAL
99 | 2023-07-05 | Jul 2023 | 16:50 | 16:52 | 0.0013889 | EXAMINATION COUCH | MAXLIFE HEALTH | VBC-1001 | HAFIZ | DIALYSIS | NOT WORKING | RESTARTED THE MACHINE. FIXED THE POWERCHORD PROPERLY. WORKING GOOD. | NO | NO | 2023-07-05 | July 2023 | 17:00 | 2023-07-05 16:50 | 2023-07-05 17:00 | 0:10 | COMPLETED | NON-CRITICAL
100 | 2023-07-06 | Jul 2023 | 18:50 | 18:55 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | RAMKIE | OT | CIRCUIT LEAK TEST FAILED | CHECKED AND RECONNECTED THE SODA LIME CHAMBER AND RECTIED | NO | NO | 2023-07-06 | July 2023 | 19:10 | 2023-07-06 18:50 | 2023-07-06 19:10 | 0:20 | COMPLETED | CRITICAL
101 | 2023-07-06 | Jul 2023 | 0:15 | 0:20 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | SANTHOSH | ICU | SUDDENLY NOT WORKING | RESET POWER CARD AND CHECKED CHARGING | NO | NO | 2023-07-06 | July 2023 | 0:25 | 2023-07-06 00:15 | 2023-07-06 00:25 | 0:10 | COMPLETED | SEMI-CRITICAL
102 | 2023-07-06 | Jul 2023 | 4:40 | 4:45 | 0.0034722 | BIPAP MACHINE | PHILIPS | A40 | SANTHOSH | WARD | LOW CIRCUIT LEAK | CHECKED AND CHANGED SETTINGS | NO | NO | 2023-07-06 | July 2023 | 4:45 | 2023-07-06 04:40 | 2023-07-06 04:45 | 0:05 | COMPLETED | SEMI-CRITICAL
103 | 2023-07-06 | Jul 2023 | 13:15 | 13:20 | 0.0034722 | CCO MONITOR | EDWARDS | HEMOSPHERE HEM1 | HAFIZ | ICU | NOT WORKING | FAULT OXIMETRY CABLE IS REPLACED WITH THE NEW OXIMETRY CABLE. | OXIMETRY CABLE | NO | 2023-07-06 | July 2023 | 13:35 | 2023-07-06 13:15 | 2023-07-06 13:35 | 0:20 | COMPLETED | CRITICAL
104 | 2023-07-06 | Jul 2023 | 14:40 | 14:45 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | WARD | ISSUE IN THE DEFIBRILLATOR MONITOR | REMOVED THE ECG CABLE AND POWER CHORD. REFIXED AND RUNNED THE OPERATIONAL CHECK. NOW, WORKING GOOD. | NO | NO | 2023-07-06 | July 2023 | 14:55 | 2023-07-06 14:40 | 2023-07-06 14:55 | 0:15 | COMPLETED | CRITICAL
105 | 2023-07-07 | Jul 2023 | 16:00 | 16:10 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | ICU | NOT WORKING | POWER CORD WAS NOT PROPERLY CONNECTED TO THE CONTROL BOX. FIXED IT AND FOUND IT'S WORKING WELL. | NO | NO | 2023-07-07 | July 2023 | 16:15 | 2023-07-07 16:00 | 2023-07-07 16:15 | 0:15 | COMPLETED | NON-CRITICAL
106 | 2023-07-07 | Jul 2023 | 1:25 | 1:28 | 0.0020833 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | SANTHOSH | OT | CIRCUIT LEAK TEST FAILED | ISSUE RECTIFIED | NO | NO | 2023-07-07 | July 2023 | 1:35 | 2023-07-07 01:25 | 2023-07-07 01:35 | 0:10 | COMPLETED | CRITICAL
107 | 2023-07-08 | Jul 2023 | 9:25 | 9:30 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | SOORAJ | OT | LOW O2% ALARM | TEMPORAILY REPLACED AN O2 SENSOR FROM ANOTHER MACHINE AND ISSUE RECTIFIED. | O2 SENSOR | NO | 2023-07-08 | July 2023 | 9:40 | 2023-07-08 09:25 | 2023-07-08 09:40 | 0:15 | COMPLETED | CRITICAL
108 | 2023-07-08 | Jul 2023 | 10:45 | 10:50 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SOORAJ | OT | SERVICE CHECK INDIACATION | PRE OP CHECK DONE AND ISSUE RECTIFIED | NO | NO | 2023-07-08 | July 2023 | 10:55 | 2023-07-08 10:45 | 2023-07-08 10:55 | 0:10 | COMPLETED | CRITICAL
109 | 2023-07-08 | Jul 2023 | 19:20 | 19:45 | 0.017361 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | OT | NOT WORKING | AFTER RESET IT WORKING | NO | NO | 2023-07-08 | July 2023 | 19:55 | 2023-07-08 19:20 | 2023-07-08 19:55 | 0:35 | COMPLETED | NON-CRITICAL
110 | 2023-07-09 | Jul 2023 | 16:45 | 16:48 | 0.0020833 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | CATRIDGE GOT EMPTY | REPLACED NEW CATRIDGE | CARTRIDGE | NO | 2023-07-09 | July 2023 | 18:00 | 2023-07-09 16:45 | 2023-07-09 18:00 | 1:15 | COMPLETED | SEMI-CRITICAL
111 | 2023-07-09 | Jul 2023 | 6:30 | 6:30 | 0 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | SANTHOSH | ICU | PLUGTOP BROKEN | REPLACED NEW PLUG TOP | PLUG TOP | NO | 2023-07-09 | July 2023 | 6:45 | 2023-07-09 06:30 | 2023-07-09 06:45 | 0:15 | COMPLETED | SEMI-CRITICAL
112 | 2023-07-09 | Jul 2023 | 16:28 | 16:30 | 0.0013889 | USG MACHINE | MINDRAY MEDICAL INDIA | MX7 | HAFIZ | WARD | NOT WORKING | CHECKED THE CONNECTIONS AND FOUND THAT POWERCHORD NOT FIXED PROPERLY.FIXED THE POWERCHORD AND RESTARTED THE MACHINE. NOW, WORKING GOOD. | NO | NO | 2023-07-09 | July 2023 | 16:35 | 2023-07-09 16:28 | 2023-07-09 16:35 | 0:07 | COMPLETED | SEMI-CRITICAL
113 | 2023-07-09 | Jul 2023 | 16:37 | 16:40 | 0.0020833 | DIALYSIS MACHINE | FRESENIUS | 4008 S | HAFIZ | DIALYSIS | FLOW ERROR | CHECKED THE PRESSURE IN RO PLANT. IT IS IN NORMAL RANGE AND WORKING GOOD. | NO | NO | 2023-07-09 | July 2023 | 16:45 | 2023-07-09 16:37 | 2023-07-09 16:45 | 0:08 | COMPLETED | CRITICAL
114 | 2023-07-09 | Jul 2023 | 17:10 | 17:12 | 0.0013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | ICU | BUTTONS NOT WORKING | CHECKED THE BED AND FOUND THAT IT WAS LOCKED. UNLOCKED THE BUTTON AND NOW WORKING FINE. | NO | NO | 2023-07-09 | July 2023 | 17:15 | 2023-07-09 17:10 | 2023-07-09 17:15 | 0:05 | COMPLETED | NON-CRITICAL
115 | 2023-07-10 | Jul 2023 | 12:20 | 12:25 | 0.0034722 | OT TABLE | MAQUET | MEERA | RAMKIE | OT | SLIDING MOVEMENT PROBLEM | AFTER CONTINUES MOVEMENT IT GOT SOLVED, NEED TO DO ADJUSTMENT USING DEVICE | NO | NO | 2023-07-10 | July 2023 | 12:35 | 2023-07-10 12:20 | 2023-07-10 12:35 | 0:15 | COMPLETED | NON-CRITICAL
116 | 2023-07-10 | Jul 2023 | 17:00 | 17:15 | 0.010417 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | WARD | BROKEN PLUG TOP | DEFECTIVE PLUG TOP REPLACED WITH A NEW ONE. | PLUG TOP | NO | 2023-07-10 | July 2023 | 17:25 | 2023-07-10 17:00 | 2023-07-10 17:25 | 0:25 | COMPLETED | NON-CRITICAL
117 | 2023-07-11 | Jul 2023 | 23:00 | 23:05 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SOORAJ | ICU | PO2 FIXING | RAN HEPARIN FIVE TIMES AND ISSUE RECTIFIED. | NO | NO | 2023-07-11 | July 2023 | 23:10 | 2023-07-11 23:00 | 2023-07-11 23:10 | 0:10 | COMPLETED | SEMI-CRITICAL
118 | 2023-07-11 | Jul 2023 | 13:40 | 13:45 | 0.0034722 | VENTILATOR, ADULT | MAQUET | SERVO-I | RAMKIE | ICU | PRE USE CHECK TEST FAILED | REPLACED THE EXP CASSETTE | EXPIRATORY CASSETTE | NO | 2023-07-11 | July 2023 | 13:55 | 2023-07-11 13:40 | 2023-07-11 13:55 | 0:15 | COMPLETED | CRITICAL
119 | 2023-07-11 | Jul 2023 | 15:30 | 15:35 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | JERLIN | RADIOLOGY | PROBES NOT DETECTED | CLEANED AND REFIXED THE PROBES. ISSUE RECTIFIED. | NO | NO | 2023-07-11 | July 2023 | 15:40 | 2023-07-11 15:30 | 2023-07-11 15:40 | 0:10 | COMPLETED | SEMI-CRITICAL
120 | 2023-07-11 | Jul 2023 | 18:00 | 18:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | JERLIN | ICU | ECG WAVE ISSUE | REFIXED CABLE ,NOW WORKING | NO | NO | 2023-07-11 | July 2023 | 18:15 | 2023-07-11 18:00 | 2023-07-11 18:15 | 0:15 | COMPLETED | SEMI-CRITICAL
121 | 2023-07-11 | Jul 2023 | 10:55 | 11:00 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | HAFIZ | ICU | EXCESSIVE LEAKAGE ERROR | REMOVED THE EXPIRATORY CASSETTE, CLEANED AND REFIXED INTO THE VENTILATOR. RERUN THE PRE-USE CHECK AND PASSED. NOW, WORKING FINE. | NO | NO | 2023-07-11 | July 2023 | 11:10 | 2023-07-11 10:55 | 2023-07-11 11:10 | 0:15 | COMPLETED | CRITICAL
122 | 2023-07-12 | Jul 2023 | 4:30 | 4:40 | 0.0069444 | IABP MACHINE | MAQUET | CS300 | SOORAJ | ICU | KEYS NOT FUNCTIONING | RESTARTED THE SYSTEM AND ISSUE RECTIFIED. | NO | NO | 2023-07-12 | July 2023 | 4:45 | 2023-07-12 04:30 | 2023-07-12 04:45 | 0:15 | COMPLETED | CRITICAL
123 | 2023-07-12 | Jul 2023 | 15:30 | 15:35 | 0.0034722 | DEFIBRILLATOR MONITOR | EFFICIA DFM 100 | CN32624356 | RAMKIE | EMERGENCY | NOT READY TO USE ECG CABLE MALFUNCTION | CHECKED AND RECONNECTED AND OPERATIONAL CHECK PASSED | NO | NO | 2023-07-12 | July 2023 | 15:45 | 2023-07-12 15:30 | 2023-07-12 15:45 | 0:15 | COMPLETED | CRITICAL
124 | 2023-07-12 | Jul 2023 | 18:00 | 18:15 | 0.010417 | VIDEO CALPOSCOPE | BORZE | HD-INTELLIO | JERLIN | OPD | NOT ABLE TO CAPTURE IMAGES | IMAGE CAPTURING ISSUE RECTIFIED | NO | NO | 2023-07-12 | July 2023 | 18:30 | 2023-07-12 18:00 | 2023-07-12 18:30 | 0:30 | COMPLETED | SEMI-CRITICAL
125 | 2023-07-12 | Jul 2023 | 14:00 | 14:05 | 0.0034722 | CAMERA SYSTEM,ICG | KARL STORZ | TC200 | HAFIZ | OT | NOT WORKING | ACCOMPANIED WITH THE COMPANY PERSON. NEED TO REPLACE THE SCB CABLE. | NO | SCB CABLE | 2023-07-12 | July 2023 | 14:15 | 2023-07-12 14:00 | 2023-07-12 14:15 | 0:15 | PENDING | SEMI-CRITICAL
126 | 2023-07-14 | Jul 2023 | 16:45 | 16:50 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | RAMKIE | OT | O2 CELL ISSUE | REPLACED O2 SENSOR | JUMBO O2 SENSOR | NO | 2023-07-14 | July 2023 | 17:00 | 2023-07-14 16:45 | 2023-07-14 17:00 | 0:15 | COMPLETED | CRITICAL
127 | 2023-07-14 | Jul 2023 | 16:30 | 16:40 | 0.0069444 | BMI MACHINE | AKAS MEDICAL | QVS-100 | JERLIN | OPD | AIR PRESSURE ERROR | FIXED THE CONNECTOR IN THE TUBING AND NOW IT'S WORKING WELL. | NO | NO | 2023-07-14 | July 2023 | 16:50 | 2023-07-14 16:30 | 2023-07-14 16:50 | 0:20 | COMPLETED | NON-CRITICAL
128 | 2023-07-14 | Jul 2023 | 17:45 | 17:50 | 0.0034722 | RADIANT HEAT WARMER | WIPRO GE HEALTH CARE | LULLABY | JERLIN | ICU | TEMPERATURE PROBE NOT WORKING | CHECKED AND REFIXED THE TEMPERATURE PROBE AND FOUND IT'S WORKING GOOD. | NO | NO | 2023-07-14 | July 2023 | 17:55 | 2023-07-14 17:45 | 2023-07-14 17:55 | 0:10 | COMPLETED | SEMI-CRITICAL
129 | 2023-07-17 | Jul 2023 | 15:20 | 15:25 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | SOORAJ | ICU | ECG NOT SHOWING | ECG BOARD ISSUE IN MMX MODULE. REPLACED BOARD AND ISSUE RECTIFIED. | MAIN BOARD | NO | 2023-07-17 | July 2023 | 15:40 | 2023-07-17 15:20 | 2023-07-17 15:40 | 0:20 | COMPLETED | SEMI-CRITICAL
130 | 2023-07-17 | Jul 2023 | 10:20 | 10:30 | 0.0069444 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | RAMKIE | OT | NOT WORKING | NO ISSUES FOUND WORKING GOOD | NO | NO | 2023-07-17 | July 2023 | 10:40 | 2023-07-17 10:20 | 2023-07-17 10:40 | 0:20 | COMPLETED | NON-CRITICAL
131 | 2023-07-17 | Jul 2023 | 14:10 | 14:15 | 0.0034722 | VENTILATOR, PEDIATRIC | MINDRAY MEDICAL INDIA | SV 300 | SANTHOSH | ICU | PRE-USE CHECK FAILED | CHECKED O2 SENSOR, ITS DEFECT | NO | O2 SENSOR | 2023-07-17 | July 2023 | 14:25 | 2023-07-17 14:10 | 2023-07-17 14:25 | 0:15 | PENDING | CRITICAL
132 | 2023-07-17 | Jul 2023 | 14:15 | 14:20 | 0.0034722 | VENTILATOR, UNIVERSAL | MINDRAY MEDICAL INDIA | SV 300 | SANTHOSH | ICU | PRE-USE CHECK FAILED | CHECKED O2 SENSOR, ITS DEFECT | NO | O2 SENSOR | 2023-07-17 | July 2023 | 14:30 | 2023-07-17 14:15 | 2023-07-17 14:30 | 0:15 | PENDING | CRITICAL
133 | 2023-07-18 | Jul 2023 | 8:20 | 8:25 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | RAMKIE | OT | CIRCUIT LEAK TEST FAILED | SODA LIME CHAMBER REDUCED QUANTITY AND BELLOWS CLEANED | NO | NO | 2023-07-18 | July 2023 | 8:35 | 2023-07-18 08:20 | 2023-07-18 08:35 | 0:15 | COMPLETED | CRITICAL
134 | 2023-07-18 | Jul 2023 | 13:10 | 13:15 | 0.0034722 | VIDEO CALPOSCOPE | BORZE | HD-INTELLIO | JERLIN | OPD | NOT ABLE TO PRINT REPORT | CHECKED AND PRINTER ISSUE RECTIFIED | NO | NO | 2023-07-18 | July 2023 | 13:25 | 2023-07-18 13:10 | 2023-07-18 13:25 | 0:15 | COMPLETED | SEMI-CRITICAL
135 | 2023-07-18 | Jul 2023 | 13:00 | 13:05 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SANTHOSH | ICU | SIDE RAIL DIDNT LOCK | ISSUE RECTIFIED | NO | NO | 2023-07-18 | July 2023 | 13:10 | 2023-07-18 13:00 | 2023-07-18 13:10 | 0:10 | COMPLETED | NON-CRITICAL
136 | 2023-07-18 | Jul 2023 | 1:00 | 1:07 | 0.0048611 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | HAFIZ | ICU | FLOW TRANSDUCER TEST FAILED. | RERUN THE PRE-USE CHECK WITH THE PATIENT CIRCUIT. PASSED THE TEST AND NOW WORKING FINE. | NO | NO | 2023-07-18 | July 2023 | 1:20 | 2023-07-18 01:00 | 2023-07-18 01:20 | 0:20 | COMPLETED | CRITICAL
137 | 2023-07-18 | Jul 2023 | 7:55 | 7:57 | 0.0013889 | VENTILATOR, ADULT | MAQUET | SERVO-I | HAFIZ | ICU | PRESSURE TRANSDUCER TEST FAILED | REMOVED THE EXPIRATORY CASSETTE, CLEANED AND REFIXED INTO THE VENTILATOR. RERUN THE PRE-USE CHECK AND PASSED. NOW, WORKING FINE. | NO | NO | 2023-07-18 | July 2023 | 8:07 | 2023-07-18 07:55 | 2023-07-18 08:07 | 0:12 | COMPLETED | CRITICAL
138 | 2023-07-19 | Jul 2023 | 17:00 | 17:10 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 450 | JERLIN | EMERGENCY | NOT SWITCHING ON | REFIXED A POWER CORD AND OBSERVED THAT IT'S WORKING GOOD. | NO | NO | 2023-07-19 | July 2023 | 17:15 | 2023-07-19 17:00 | 2023-07-19 17:15 | 0:15 | COMPLETED | SEMI-CRITICAL
139 | 2023-07-19 | Jul 2023 | 14:10 | 14:15 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | WARD | SIDE RAIL NOT WORKING | FIXED THE SIDE RAIL | NO | NO | 2023-07-19 | July 2023 | 14:25 | 2023-07-19 14:10 | 2023-07-19 14:25 | 0:15 | COMPLETED | NON-CRITICAL
140 | 2023-07-19 | Jul 2023 | 14:30 | 14:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | ICU | SIDE RAIL NOT WORKING | FIXED THE SIDE RAIL | NO | NO | 2023-07-19 | July 2023 | 14:45 | 2023-07-19 14:30 | 2023-07-19 14:45 | 0:15 | COMPLETED | NON-CRITICAL
141 | 2023-07-19 | Jul 2023 | 16:40 | 16:50 | 0.0069444 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | JERLIN | EMERGENCY | NIBP NOT FUNCTIONING PROPERLY | REPLACED WITH A NEW BP CUFF AND RECTIFIED THE ISSUE. | BP CUFF | NO | 2023-07-19 | July 2023 | 17:00 | 2023-07-19 16:40 | 2023-07-19 17:00 | 0:20 | COMPLETED | SEMI-CRITICAL
142 | 2023-07-19 | Jul 2023 | 15:30 | 15:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | ICU | SIDE RAIL NOT WORKING | FIXED THE SIDE RAIL | NO | NO | 2023-07-19 | July 2023 | 15:45 | 2023-07-19 15:30 | 2023-07-19 15:45 | 0:15 | COMPLETED | NON-CRITICAL
143 | 2023-07-19 | Jul 2023 | 15:10 | 15:15 | 0.0034722 | INTRA ORAL XRAY | SKANRAY DENTAL | INTRA SKANDC | SANTHOSH | OPD | NOT TURNING ON | BURNT SMELL OCCURS NEED TO REPLACE THE BOARD | NO | MAIN BOARD | 2023-07-19 | July 2023 | 15:20 | 2023-07-19 15:10 | 2023-07-19 15:20 | 0:10 | PENDING | SEMI-CRITICAL
144 | 2023-07-19 | Jul 2023 | 5:35 | 5:40 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | HAFIZ | WARD | PATIENT WARMER NOT WORKING | REFIXED THE CABLE AND RESTARTED THE MACHINE. NOW WORKING GOOD. | NO | NO | 2023-07-19 | July 2023 | 5:46 | 2023-07-19 05:35 | 2023-07-19 05:46 | 0:11 | COMPLETED | SEMI-CRITICAL
145 | 2023-07-20 | Jul 2023 | 14:30 | 14:40 | 0.0069444 | PRESSURE INJECTOR | LF OPTI VONTAGE | ILLUMENO NEO | JERLIN | CATHLAB | INJECTOR CASE NOT FIXED | FIXED THE CASE PROPERLY AND RECTIFIED THE ISSUE. | NO | NO | 2023-07-20 | July 2023 | 14:50 | 2023-07-20 14:30 | 2023-07-20 14:50 | 0:20 | COMPLETED | SEMI-CRITICAL
146 | 2023-07-20 | Jul 2023 | 21:50 | 21:55 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | HAFIZ | ICU | NOT GETTING PROPER VALUES | RERUN THE HEPARIN 4-5 TIMES. NOW WORKING GOOD. | NO | NO | 2023-07-20 | July 2023 | 22:18 | 2023-07-20 21:50 | 2023-07-20 22:18 | 0:28 | COMPLETED | SEMI-CRITICAL
147 | 2023-07-21 | Jul 2023 | 17:20 | 17:30 | 0.0069444 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | JERLIN | ICU | CASSETTE FAILED | PRE USE CHECK OUT PERFORMED AND FOUND IN WORKING CONDITION. | NO | NO | 2023-07-21 | July 2023 | 17:45 | 2023-07-21 17:20 | 2023-07-21 17:45 | 0:25 | COMPLETED | CRITICAL
148 | 2023-07-22 | Jul 2023 | 8:30 | 8:40 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | ENDOSCOPY | CROSS MARK IN DEFIBRILATOR | FIXED ECG CABLE AND PERFORMED OP CHECK. ISSUE SOLVED. | NO | NO | 2023-07-22 | July 2023 | 8:50 | 2023-07-22 08:30 | 2023-07-22 08:50 | 0:20 | COMPLETED | CRITICAL
149 | 2023-07-22 | Jul 2023 | 9:00 | 9:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | WARD | PRINTER NOT WORKING | FOUND THAT PROPER PAPER WAS NOT FIXED IN THE PRINTER. CROSS CHECKED WITH APPROPRIATE PAPER AND FOUND IT'S WORKING GOOD. | PAPER ROLL | NO | 2023-07-22 | July 2023 | 9:20 | 2023-07-22 09:00 | 2023-07-22 09:20 | 0:20 | COMPLETED | CRITICAL
150 | 2023-07-22 | Jul 2023 | 17:20 | 17:25 | 0.0034722 | ACT MACHINE | HELENO LABORATORIES | MINI II | SANTHOSH | ICU | SAMPLE DOESNT ROTATE | REFIXED THE CABLES ,ITS RECTIFIES | NO | NO | 2023-07-22 | July 2023 | 17:45 | 2023-07-22 17:20 | 2023-07-22 17:45 | 0:25 | COMPLETED | SEMI-CRITICAL
151 | 2023-07-23 | Jul 2023 | 12:20 | 12:45 | 0.017361 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | OT | NOT WORKING | CONNECTED POWER AND RECTIFIED | NO | NO | 2023-07-23 | July 2023 | 12:55 | 2023-07-23 12:20 | 2023-07-23 12:55 | 0:35 | COMPLETED | NON-CRITICAL
152 | 2023-07-23 | Jul 2023 | 12:35 | 12:40 | 0.0034722 | VENTILATOR, ADULT | MAQUET | SERVO-I | SANTHOSH | ICU | LEAK TEST FAILED | CHECKED WITH ANOTHER CASSETE ,ITS RECTIFIED NOW | EXPIRATORY CASSETTE | NO | 2023-07-23 | July 2023 | 12:40 | 2023-07-23 12:35 | 2023-07-23 12:40 | 0:05 | COMPLETED | CRITICAL
153 | 2023-07-23 | Jul 2023 | 15:10 | 15:15 | 0.0034722 | PULSE OXIMETER,PORTABLE | COVIDIEN | NELLCOR | SANTHOSH | WARD | SPO2 NOT WORKING | REFIXED CABLE ,NOW WORKING | NO | NO | 2023-07-23 | July 2023 | 15:20 | 2023-07-23 15:10 | 2023-07-23 15:20 | 0:10 | COMPLETED | SEMI-CRITICAL
154 | 2023-07-24 | Jul 2023 | 16:20 | 16:30 | 0.0069444 | DIALYSIS MACHINE | FRESENIUS | 4008 S | RAMKIE | DIALYSIS | V188 | CLEANED RINSE FILTER | RINSE FILTER | NO | 2023-07-24 | July 2023 | 16:40 | 2023-07-24 16:20 | 2023-07-24 16:40 | 0:20 | COMPLETED | CRITICAL
155 | 2023-07-25 | Jul 2023 | 7:45 | 7:50 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | SOORAJ | ICU | BATTERY MALFUNCTION MESSAGE | BATTERY GONE DEFECTIVE. DEFECTIVE BATTERY TAKEN TO DEPARTMENT. AS OF NOW MONITOR WORKING ON AC SUPPLY. | NO | BATTERY | 2023-07-25 | July 2023 | 8:25 | 2023-07-25 07:45 | 2023-07-25 08:25 | 0:40 | PENDING | SEMI-CRITICAL
156 | 2023-07-25 | Jul 2023 | 13:15 | 13:20 | 0.0034722 | HEATER COOLER | GETINGE | HCU40 | SOORAJ | OT | MAKING SOUND WHILE RUNNING. | DEAIRING DONE AND ISSUE RECTIFIED. | NO | NO | 2023-07-25 | July 2023 | 13:35 | 2023-07-25 13:15 | 2023-07-25 13:35 | 0:20 | COMPLETED | CRITICAL
157 | 2023-07-25 | Jul 2023 | 17:30 | 17:40 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | RAMKIE | ICU | NOT SWITCHING ON | POWER LOOSE CONNECTION | NO | NO | 2023-07-25 | July 2023 | 17:50 | 2023-07-25 17:30 | 2023-07-25 17:50 | 0:20 | COMPLETED | SEMI-CRITICAL
158 | 2023-07-25 | Jul 2023 | 18:17 | 18:20 | 0.0020833 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | HAFIZ | ICU | IV STAND CLAMP BROKEN | REPLACED WITH THE FAULTY SYRINGE CLAMP. NOW, WORKING GOOD. | NO | NO | 2023-07-25 | July 2023 | 18:27 | 2023-07-25 18:17 | 2023-07-25 18:27 | 0:10 | COMPLETED | NON-CRITICAL
159 | 2023-07-26 | Jul 2023 | 10:20 | 10:30 | 0.0069444 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX | RAMKIE | OT | ERROR 121 | NEED TO DO CALIBRATION | NO | NO | 2023-07-26 | July 2023 | 10:35 | 2023-07-26 10:20 | 2023-07-26 10:35 | 0:15 | COMPLETED | SEMI-CRITICAL
160 | 2023-07-27 | Jul 2023 | 14:40 | 14:43 | 0.0020833 | VENTILATOR, ADULT | MAQUET | SERVO-I | HAFIZ | ICU | INTERNAL LEAKAGE TEST FAILED AND PRESSURE TRANSDUCER TEST FAILED | REMOVED THE EXPIRATORY CASSETTE, CLEANED AND REFIXED INTO THE VENTILATOR. RERUN THE PRE-USE CHECK AND PASSED. NOW, WORKING FINE. | NO | NO | 2023-07-27 | July 2023 | 14:53 | 2023-07-27 14:40 | 2023-07-27 14:53 | 0:13 | COMPLETED | CRITICAL
161 | 2023-07-27 | Jul 2023 | 14:55 | 14:57 | 0.0013889 | VENTILATOR, ADULT | MAQUET | SERVO-I | HAFIZ | ICU | INTERNAL LEAKAGE TEST FAILED | REMOVED THE EXPIRATORY CASSETTE, CLEANED AND REFIXED INTO THE VENTILATOR. RERUN THE PRE-USE CHECK AND PASSED. NOW, WORKING FINE. | NO | NO | 2023-07-27 | July 2023 | 15:10 | 2023-07-27 14:55 | 2023-07-27 15:10 | 0:15 | COMPLETED | CRITICAL
162 | 2023-07-28 | Jul 2023 | 7:30 | 7:35 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | SOORAJ | OT | ETCO2 NOT SHOWING | ADJUSTED THE WATER TRAP POSITION ON GAS MODULE AND ISSUE RECTIFIED. | NO | NO | 2023-07-28 | July 2023 | 7:45 | 2023-07-28 07:30 | 2023-07-28 07:45 | 0:15 | COMPLETED | SEMI-CRITICAL
163 | 2023-07-28 | Jul 2023 | 7:48 | 7:50 | 0.0013889 | FLUID WARMER SYSTEM | MATRIX | MN05 | SOORAJ | OT | NO CIRCULATION IN BLANKET. | BLANKET LEAK. CHANGED BLANKET AND ISSUE RECTIFIED. | HEMOTHERM BLANKET | NO | 2023-07-28 | July 2023 | 7:55 | 2023-07-28 07:48 | 2023-07-28 07:55 | 0:07 | COMPLETED | SEMI-CRITICAL
164 | 2023-07-28 | Jul 2023 | 9:00 | 9:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SOORAJ | OT | NO BATTERY BACKUP | BATTERY DRAINED COMPLETELY, KEPT ON CHARGING AND ISSUE RECTIFIED. | NO | NO | 2023-07-28 | July 2023 | 11:05 | 2023-07-28 09:00 | 2023-07-28 11:05 | 2:05 | COMPLETED | CRITICAL
165 | 2023-07-28 | Jul 2023 | 10:50 | 10:55 | 0.0034722 | HEATER COOLER | GETINGE | HCU40 | SOORAJ | OT | LEAK ON HEMOTHERM BLANKET | BLANKET HAD MULTIPLE HOLES. NEW BLANKET GIVEN AS REPLACEMENT. ISSUE RECTIFIED. | HEMOTHERM BLANKET | NO | 2023-07-28 | July 2023 | 11:00 | 2023-07-28 10:50 | 2023-07-28 11:00 | 0:10 | COMPLETED | SEMI-CRITICAL
166 | 2023-07-28 | Jul 2023 | 16:30 | 16:40 | 0.0069444 | PLEURA VIDEOSCOPE,FLEXIBLE | OLYMPUS | LTF-160(E) | RAMKIE | ENDOSCOPY | LEAK IN WORKING CHANNEL | COMPANY PERSON CHECKED AND TAKEN TO INHOUSE SERVICE | NO | NEED SERVICE | 2023-07-28 | July 2023 | 17:00 | 2023-07-28 16:30 | 2023-07-28 17:00 | 0:30 | PENDING | SEMI-CRITICAL
167 | 2023-07-28 | Jul 2023 | 17:35 | 17:40 | 0.0034722 | AUTOMATED ELISA READER | COBAS E411 | ROCHE | HAFIZ | BLOODBANK | THE REAGENT ROTOR DIDN'T REACH THE HOME POSITION. | REMOVED THE REAGENT ROTOR AND FOUND SMALL DUST PARTICLE INSIDE. REMOVED IT AND THEN PLACED THE ROTOR AT THE PROPER POSITION. NOW WORKINH GOOD. | NO | NO | 2023-07-28 | July 2023 | 17:50 | 2023-07-28 17:35 | 2023-07-28 17:50 | 0:15 | COMPLETED | SEMI-CRITICAL
168 | 2023-07-29 | Jul 2023 | 8:10 | 8:15 | 0.0034722 | IABP MACHINE | GETINGE | CS300 | RAMKIE | OT | HELIUM CYLINDER LEAK | REPLACED THE WASHER AND THE CYLINDER IS GOOD | WASHER | NO | 2023-07-29 | July 2023 | 8:20 | 2023-07-29 08:10 | 2023-07-29 08:20 | 0:10 | COMPLETED | CRITICAL
169 | 2023-07-29 | Jul 2023 | 14:10 | 14:14 | 0.0027778 | ACT MACHINE | HELENO LABORATORIES | ACTALYKE | HAFIZ | CATHLAB | ACT MACHINE NOT WORKING | REMOVED THE POWERCHORD. REFIXED AND RUNNED WITH ANOTHER BLOOD SAMPLE. NOW, WORKING GOOD. | NO | NO | 2023-07-29 | July 2023 | 14:20 | 2023-07-29 14:10 | 2023-07-29 14:20 | 0:10 | COMPLETED | SEMI-CRITICAL
170 | 2023-07-31 | Jul 2023 | 10:50 | 10:55 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | SOORAJ | ICU | MONITOR BROKEN | MONITOR MISHANDLED AND PHYSICAL DAMAGE ONSERVED ON DISPLAY. OBSERVED BOOTING OK, DISPLAY OK, TOUCH NOT WORKING. | NO | DISPLAY | 2023-07-31 | July 2023 | 11:30 | 2023-07-31 10:50 | 2023-07-31 11:30 | 0:40 | PENDING | SEMI-CRITICAL
171 | 2023-07-31 | Jul 2023 | 19:35 | 19:40 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | SUDDENLY SHUT DOWN | HARDWARE FAILURE ,FIXED NOW | NO | NO | 2023-07-31 | July 2023 | 19:50 | 2023-07-31 19:35 | 2023-07-31 19:50 | 0:15 | COMPLETED | SEMI-CRITICAL
172 | 2023-08-01 | Aug 2023 | 23:30 | 23:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | WARD | COT NOT WORKING | FOUND THAT THE REMOTE WAS LOCKED. LOCK RELEASED AND ISSUE RECTIFIED. | NO | NO | 2023-08-01 | August 2023 | 23:40 | 2023-08-01 23:30 | 2023-08-01 23:40 | 0:10 | COMPLETED | NON-CRITICAL
173 | 2023-08-01 | Aug 2023 | 1:00 | 1:10 | 0.0069444 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | JERLIN | ICU | STOPS WORKING WITHIN A FEW MINUTES AFTER SWITCHING ON. | FIXED THE TEMPERATURE PROBE PROPERLY AND RECTIFIED THE ISSUE. | NO | NO | 2023-08-01 | August 2023 | 1:20 | 2023-08-01 01:00 | 2023-08-01 01:20 | 0:20 | COMPLETED | SEMI-CRITICAL
174 | 2023-08-01 | Aug 2023 | 22:00 | 22:15 | 0.010417 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | JERLIN | EMERGENCY | V3 ERROR | REPLACED THE DEFECTIVE CABLE WITH A NEW ONE AND RECTIFIED THE ISSUE | 10-LEADS ECG CABLE | NO | 2023-08-01 | August 2023 | 22:30 | 2023-08-01 22:00 | 2023-08-01 22:30 | 0:30 | COMPLETED | SEMI-CRITICAL
175 | 2023-08-01 | Aug 2023 | 18:30 | 18:35 | 0.0034722 | CENTRIFUGE | REMI | R-8C PLUS | SANTHOSH | LABORATORY | CENTRIFUGE MAKING NOISE | CHECKED WHEIGHT BALANCE,ALTERNED WHEIGHTS | NO | NO | 2023-08-01 | August 2023 | 19:00 | 2023-08-01 18:30 | 2023-08-01 19:00 | 0:30 | COMPLETED | SEMI-CRITICAL
176 | 2023-08-01 | Aug 2023 | 11:40 | 11:50 | 0.0069444 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | NOT WORKING (CO-Ox HARDWARE FAILURE) | ACCOMAPANIED WITH THE COMPANY PERSON. CHANGED THE EPROM OF THE CURRENTLY USING CATRIDGE. CVP TEST DONE.NOW, WORKING FINE. | EPROM | NO | 2023-08-01 | August 2023 | 12:00 | 2023-08-01 11:40 | 2023-08-01 12:00 | 0:20 | COMPLETED | SEMI-CRITICAL
177 | 2023-08-01 | Aug 2023 | 13:50 | 13:55 | 0.0034722 | CENTRIFUGE TABLE TOP | REMI | R-8C PLUS | HAFIZ | LABORATORY | UNIT NOISE PROBLEM | PLACED THE SAMPLE IN EVEN ORDER. TIGTHENED THE SCREON THE LEFT SIDE. NOW, IT IS RECTIFIED | NO | NO | 2023-08-01 | August 2023 | 14:00 | 2023-08-01 13:50 | 2023-08-01 14:00 | 0:10 | COMPLETED | SEMI-CRITICAL
178 | 2023-08-01 | Aug 2023 | 16:00 | 16:05 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | WARD | ECG LEAD TEST FAILED. | FOUND FAULT IN ECG CABLE. GIVEN ANOTHER ECG CABLE FROM THE DEPARTMENT. PASSED THE OPERATIONAL CHECK. NOW, WORKING GOOD. | ECG CABLE | NO | 2023-08-01 | August 2023 | 16:15 | 2023-08-01 16:00 | 2023-08-01 16:15 | 0:15 | COMPLETED | CRITICAL
179 | 2023-08-01 | Aug 2023 | 10:10 | 10:20 | 0.0069444 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | RAMKIE | EMERGENCY | V4 LEAD NOT AVAILABLE | CHECKED AND FOUND THAT CABLE DAMAGE, CONFIRMED THE PROBLME WITH ECG CABLE | NO | ECG CABLE | 2023-08-01 | August 2023 | 10:30 | 2023-08-01 10:10 | 2023-08-01 10:30 | 0:20 | PENDING | SEMI-CRITICAL
180 | 2023-08-02 | Aug 2023 | 5:30 | 5:40 | 0.0069444 | PATIENT MONITOR | PHILIPS | CM12 | JERLIN | EMERGENCY | NEED TO CHANGE THE NIBP SETTINGS | CHANGE THE SETTINGS ACCORDING TO THE PATIENT'S CONDITION. | NO | NO | 2023-08-02 | August 2023 | 5:45 | 2023-08-02 05:30 | 2023-08-02 05:45 | 0:15 | COMPLETED | SEMI-CRITICAL
181 | 2023-08-02 | Aug 2023 | 8:50 | 8:52 | 0.0013889 | DYONICS POWER CONTROL UNIT,ARTHROSCOPY | SMITH & NEPHEW | 72200873 | HAFIZ | OT | PROBLEM IN DISPLAY DURING CASE | REMOVED THE CABLE AND REFIXED PROPERLY. NOW WORKING GOOD. | NO | NO | 2023-08-02 | August 2023 | 8:55 | 2023-08-02 08:50 | 2023-08-02 08:55 | 0:05 | COMPLETED | SEMI-CRITICAL
182 | 2023-08-02 | Aug 2023 | 9:50 | 10:00 | 0.0069444 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | RAMKIE | ICU | CATRIDGE REPLACEMENT | REPLACED NEW CATRIDGE | CARTRIDGE | NO | 2023-08-02 | August 2023 | 10:10 | 2023-08-02 09:50 | 2023-08-02 10:10 | 0:20 | COMPLETED | SEMI-CRITICAL
183 | 2023-08-03 | Aug 2023 | 22:30 | 22:40 | 0.0069444 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | ICU | IV POLE CLAMP BROKEN | CHANGED THE BACK PANEL OF THE INFUSION PUMP AND RECTIFIED THE ISSUE. | NO | NO | 2023-08-03 | August 2023 | 22:55 | 2023-08-03 22:30 | 2023-08-03 22:55 | 0:25 | COMPLETED | NON-CRITICAL
184 | 2023-08-03 | Aug 2023 | 6:20 | 6:30 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | JERLIN | OT | CIRCUIT LEAK TEST FAILED | REFIXED THE RESPIRATORY UNIT AND CANISTER PROPERLY AND RE-RAN THE TEST . TEST PASSED AND ISSUE RECTIFIED | NO | NO | 2023-08-03 | August 2023 | 6:45 | 2023-08-03 06:20 | 2023-08-03 06:45 | 0:25 | COMPLETED | CRITICAL
185 | 2023-08-03 | Aug 2023 | 23:00 | 23:10 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | PRINTER NOT WORKING | FIXED THE PRINTER PAPER PROPERLY AND RECTIFIED THE ISSUE | NO | NO | 2023-08-03 | August 2023 | 23:15 | 2023-08-03 23:00 | 2023-08-03 23:15 | 0:15 | COMPLETED | SEMI-CRITICAL
186 | 2023-08-03 | Aug 2023 | 17:30 | 17:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SANTHOSH | ICU | NOT GETTING ON | OBSERVED POWER DISCONNECTED | NO | NO | 2023-08-03 | August 2023 | 17:40 | 2023-08-03 17:30 | 2023-08-03 17:40 | 0:10 | COMPLETED | NON-CRITICAL
187 | 2023-08-03 | Aug 2023 | 19:30 | 19:35 | 0.0034722 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | ePM 12c | SANTHOSH | ICU | SPO2 NOT WORKING | REFIXED THE CABLE | NO | NO | 2023-08-03 | August 2023 | 19:45 | 2023-08-03 19:30 | 2023-08-03 19:45 | 0:15 | COMPLETED | SEMI-CRITICAL
188 | 2023-08-03 | Aug 2023 | 8:22 | 8:25 | 0.0020833 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | RERUN THE TEST AND PASSED. NOW, WORKING FINE. | NO | NO | 2023-08-03 | August 2023 | 8:30 | 2023-08-03 08:22 | 2023-08-03 08:30 | 0:08 | COMPLETED | CRITICAL
189 | 2023-08-03 | Aug 2023 | 16:10 | 16:15 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | HAFIZ | OT | NO BATTERY BACKUP ISSUE | ACCOMPANIED WITH THE COMPANY PERSON.CHECKED THE MACHINE MANUALLY AND WITH THE HELP OF COMPANY APPLICATION. BATTERY REPLACED. | BATTERY | NO | 2023-08-03 | August 2023 | 16:30 | 2023-08-03 16:10 | 2023-08-03 16:30 | 0:20 | COMPLETED | CRITICAL
190 | 2023-08-03 | Aug 2023 | 15:10 | 15:15 | 0.0034722 | NITRIC OXIDE DELIVERY SYSTEM | WIPRO GE HEALTH CARE | SLE | SOORAJ | ICU | CHANGE SENSORS MESSAGE | FOUND NO & NO2 SENSOR TO BE DEFECTIVE. SENSORS TO BE REPLACED,CURRENTLY WORKING | NO | NO | 2023-08-03 | August 2023 | 15:30 | 2023-08-03 15:10 | 2023-08-03 15:30 | 0:20 | COMPLETED | CRITICAL
191 | 2023-08-03 | Aug 2023 | 15:25 | 15:25 | 0 | NITRIC OXIDE DELIVERY SYSTEM | WIPRO GE HEALTH CARE | SLE | SOORAJ | ICU | CALIBRATION REQUIRED MESSAGE | SENSOR CALIBRATION DONE AND ISSUE RECTIFIED | NO | NO | 2023-08-03 | August 2023 | 15:55 | 2023-08-03 15:25 | 2023-08-03 15:55 | 0:30 | COMPLETED | CRITICAL
192 | 2023-08-04 | Aug 2023 | 12:55 | 13:00 | 0.0034722 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | ePM 12c | HAFIZ | ICU | SpO2 CABLE NOT WORKING | REMOVED THE CABLE, CLEANED AND REFIXED PROPERLY. NOW, WORKING GOOD. | NO | NO | 2023-08-04 | August 2023 | 13:10 | 2023-08-04 12:55 | 2023-08-04 13:10 | 0:15 | COMPLETED | SEMI-CRITICAL
193 | 2023-08-04 | Aug 2023 | 16:30 | 16:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SOORAJ | WARD | COT NOT WORKING | FOUND BATTERY DRAINED. CONENCTED ON MAINS AND ISSUE RECTIFIED. | NO | NO | 2023-08-04 | August 2023 | 16:40 | 2023-08-04 16:30 | 2023-08-04 16:40 | 0:10 | COMPLETED | NON-CRITICAL
194 | 2023-08-05 | Aug 2023 | 23:30 | 23:40 | 0.0069444 | FETAL HEART MONITOR | WIPRO GE HEALTH CARE | COROMETRICS 259 | JERLIN | WARD | TOCO PROBE NOT DETECTED | FOUND THAT THE PROBE WAS DETACHED FROM THE CONNECTRO. REPLACED WITH A NEW PROBE. | TOCO PROBE | NO | 2023-08-06 | August 2023 | 0:10 | 2023-08-05 23:30 | 2023-08-06 00:10 | 0:40 | COMPLETED | SEMI-CRITICAL
195 | 2023-08-05 | Aug 2023 | 23:30 | 23:40 | 0.0069444 | FETAL HEART MONITOR | WIPRO GE HEALTH CARE | COROMETRICS 259 | JERLIN | WARD | TIME NOT DISPLAYING PROPERLY | CHANGED THE SETTINGS AND TIME PROPERLY FIXED. | NO | NO | 2023-08-05 | August 2023 | 23:50 | 2023-08-05 23:30 | 2023-08-05 23:50 | 0:20 | COMPLETED | SEMI-CRITICAL
196 | 2023-08-05 | Aug 2023 | 9:15 | 9:20 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | HAFIZ | OPD | ARTIFACTS IN ECG RECORDINGS | REMOVED THE CABLE, CLEANED AND REFIXED PROPERLY. TESTED WITH A PATIENT AND NO ARTIFACT FOUND IN THE RECORDINGS.WORKING FINE. | NO | NO | 2023-08-05 | August 2023 | 9:35 | 2023-08-05 09:15 | 2023-08-05 09:35 | 0:20 | COMPLETED | SEMI-CRITICAL
197 | 2023-08-05 | Aug 2023 | 14:30 | 14:33 | 0.0020833 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX | SOORAJ | OT | ERROR MESSAGE WHILE USING DIATHERMY. | ISSUE RECTIFIED BY ADJUSTING RHEOSTAT | NO | NO | 2023-08-05 | August 2023 | 14:45 | 2023-08-05 14:30 | 2023-08-05 14:45 | 0:15 | COMPLETED | SEMI-CRITICAL
198 | 2023-08-06 | Aug 2023 | 10:05 | 10:10 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | HAFIZ | WARD | RESPIRATORY RATE NOT DISPLAYING | REMOVED THE CABLE, CLEANED AND REFIXED PROPERLY. NOW, WORKING GOOD. | NO | NO | 2023-08-06 | August 2023 | 10:15 | 2023-08-06 10:05 | 2023-08-06 10:15 | 0:10 | COMPLETED | SEMI-CRITICAL
199 | 2023-08-06 | Aug 2023 | 10:16 | 10:20 | 0.0027778 | BABY WEIGHING SCALE | SECA | 354 | HAFIZ | ICU | NOT WORKING | REMOVED THE BATTERY, REFIXED PROPERLY. NOW, WORKING GOOD. | NO | NO | 2023-08-06 | August 2023 | 10:26 | 2023-08-06 10:16 | 2023-08-06 10:26 | 0:10 | COMPLETED | NON-CRITICAL
200 | 2023-08-06 | Aug 2023 | 15:15 | 15:20 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SOORAJ | WARD | COT NOT CHARGING | FOUND POWER CHORD DISCONNECTED. CONNECTED AND ISSUE RECTIFIED. | NO | NO | 2023-08-06 | August 2023 | 15:25 | 2023-08-06 15:15 | 2023-08-06 15:25 | 0:10 | COMPLETED | NON-CRITICAL
201 | 2023-08-07 | Aug 2023 | 11:00 | 11:10 | 0.0069444 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | JERLIN | OPD | V4 ERROR | REPLACED THE DEFECTIVE CABLE WITH A NEW ONE AND RECTIFIED THE ISSUE | CABLE CONNECTOR | NO | 2023-08-07 | August 2023 | 11:25 | 2023-08-07 11:00 | 2023-08-07 11:25 | 0:25 | COMPLETED | SEMI-CRITICAL
202 | 2023-08-07 | Aug 2023 | 9:10 | 9:17 | 0.0048611 | VENTILATOR, ADULT | MAQUET | SERVO-I | SANTHOSH | ICU | O2 SENSOR FAILURE | NEED TO REPLACE O2 SENSOR | NO | O2 SENSOR | 2023-08-07 | August 2023 | 9:35 | 2023-08-07 09:10 | 2023-08-07 09:35 | 0:25 | PENDING | CRITICAL
203 | 2023-08-07 | Aug 2023 | 12:30 | 12:37 | 0.0048611 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SANTHOSH | OT | CALIBRATION FAILED | O2 SENSOR CLEANED AND REFIXED | NO | NO | 2023-08-07 | August 2023 | 12:40 | 2023-08-07 12:30 | 2023-08-07 12:40 | 0:10 | COMPLETED | CRITICAL
204 | 2023-08-07 | Aug 2023 | 17:15 | 17:20 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | SANTHOSH | OPD | ECG WAVE GETTING NOISE | CABLES CHECKED AND REFIXED | NO | NO | 2023-08-07 | August 2023 | 17:30 | 2023-08-07 17:15 | 2023-08-07 17:30 | 0:15 | COMPLETED | SEMI-CRITICAL
205 | 2023-08-07 | Aug 2023 | 23:06 | 23:10 | 0.0027778 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | WARD | SpO2 CABLE NOT WORKING | CHECKED AND REPLACED WITH ANOTHER CABLE. NOW, WORKING GOOD. | SPO2 PROBE | NO | 2023-08-07 | August 2023 | 23:16 | 2023-08-07 23:06 | 2023-08-07 23:16 | 0:10 | COMPLETED | SEMI-CRITICAL
206 | 2023-08-07 | Aug 2023 | 12:30 | 12:35 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S7 | SOORAJ | RADIOLOGY | HANGING ISSUE | SPOOLER WAS FULL. CLEARED SPOOLER AND ISSUE RECTIFIED. | NO | NO | 2023-08-07 | August 2023 | 12:45 | 2023-08-07 12:30 | 2023-08-07 12:45 | 0:15 | COMPLETED | SEMI-CRITICAL
207 | 2023-08-08 | Aug 2023 | 13:00 | 13:05 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | SOORAJ | RADIOLOGY | PROBE NOT SHOWING. | DISCONNECTED ALL PROBES AND RESTARTED THE SYSTEM. ISSUE RECTIFIED. | NO | NO | 2023-08-08 | August 2023 | 13:15 | 2023-08-08 13:00 | 2023-08-08 13:15 | 0:15 | COMPLETED | SEMI-CRITICAL
208 | 2023-08-08 | Aug 2023 | 17:00 | 17:10 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC V2 | SOORAJ | EMERGENCY | PROBE CONNECTED BUT NOT SHOWING ON SCREEN | FOUND PROBE TO BE LOOSELY CONNECTED. FIXED PROPERLY AND ISSUE RECTIFIED. | NO | NO | 2023-08-08 | August 2023 | 17:15 | 2023-08-08 17:00 | 2023-08-08 17:15 | 0:15 | COMPLETED | SEMI-CRITICAL
209 | 2023-08-09 | Aug 2023 | 18:00 | 18:10 | 0.0069444 | ECMO MACHINE | GETINGE | ROTAFLOW | JERLIN | ICU | BATTERY FAILED | REPLACED THE OLD BATTERY WITH NEW ONE. | BATTERY | NO | 2023-08-09 | August 2023 | 18:30 | 2023-08-09 18:00 | 2023-08-09 18:30 | 0:30 | COMPLETED | CRITICAL
210 | 2023-08-10 | Aug 2023 | 17:00 | 17:10 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | CARTRIDGE EXPIRED | CHANGED THE CARTRIDGE WITH NEW ONE AND CVP DONE | CARTRIDGE | NO | 2023-08-10 | August 2023 | 17:40 | 2023-08-10 17:00 | 2023-08-10 17:40 | 0:40 | COMPLETED | SEMI-CRITICAL
211 | 2023-08-10 | Aug 2023 | 8:10 | 8:15 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | SANTHOSH | ICU | HIGH O2 CONCENTRATION | O2 SENSOR REMAINING CAPACITY 0% | O2 SENSOR | NO | 2023-08-10 | August 2023 | 8:25 | 2023-08-10 08:10 | 2023-08-10 08:25 | 0:15 | COMPLETED | CRITICAL
212 | 2023-08-10 | Aug 2023 | 23:40 | 23:45 | 0.0034722 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | HAFIZ | WARD | NIBP NOT WORKING | CHECKED THE MACHINE AND FOUND THAT NIBP PARAMETER WAS TURNED OFF. CHANGED THE SETTINGS AND REFIXED THE NIBP PROBE.CHECKED AND WORKING. | NO | NO | 2023-08-10 | August 2023 | 23:50 | 2023-08-10 23:40 | 2023-08-10 23:50 | 0:10 | COMPLETED | SEMI-CRITICAL
213 | 2023-08-10 | Aug 2023 | 12:50 | 12:55 | 0.0034722 | PRESSURE INJECTOR | LF OPTI VONTAGE | ILLUMENO NEO | SOORAJ | CATHLAB | ERROR 56 MESSAGE | ISSUE RECTIFIED BY RESTARTING THE SYSTEM. | NO | NO | 2023-08-10 | August 2023 | 13:05 | 2023-08-10 12:50 | 2023-08-10 13:05 | 0:15 | COMPLETED | SEMI-CRITICAL
214 | 2023-08-10 | Aug 2023 | 14:40 | 14:45 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SOORAJ | ICU | CARTRIDGE EXPIRED | NEW CARTRIDGE CHANGED & CVP PASSED. | CARTRIDGE | NO | 2023-08-10 | August 2023 | 15:30 | 2023-08-10 14:40 | 2023-08-10 15:30 | 0:50 | COMPLETED | SEMI-CRITICAL
215 | 2023-08-11 | Aug 2023 | 19:00 | 19:15 | 0.010417 | DIALYSIS MACHINE | FRESENIUS | 4008 S | JERLIN | DIALYSIS | LOW TEMPERATURE ERROR | CHANGED THE FLOW MOTOR AND RECTIFIED THE ISSUE | FLOW MOTOR | NO | 2023-08-11 | August 2023 | 19:30 | 2023-08-11 19:00 | 2023-08-11 19:30 | 0:30 | COMPLETED | CRITICAL
216 | 2023-08-11 | Aug 2023 | 11:20 | 11:25 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 9000 | SANTHOSH | ICU | SHOWS 410 ERROR | POWER SUPPLY BOARD REPLACED | NO | NO | 2023-08-11 | August 2023 | 11:50 | 2023-08-11 11:20 | 2023-08-11 11:50 | 0:30 | COMPLETED | NON-CRITICAL
217 | 2023-08-11 | Aug 2023 | 11:15 | 11:20 | 0.0034722 | ANKLE MOTION DEVICE | ANKLE MOTION PRIVATE LIMITED | AM-03 | SANTHOSH | PHYSIOTHERAPY | POWER ADAPTOR WIRE PHYSICAL DAMAGE | CABLE SOLDERED AND RECTIFIED | NO | NO | 2023-08-11 | August 2023 | 11:30 | 2023-08-11 11:15 | 2023-08-11 11:30 | 0:15 | COMPLETED | SEMI-CRITICAL
218 | 2023-08-11 | Aug 2023 | 19:00 | 19:05 | 0.0034722 | PRESSURE INJECTOR,DUAL | LF OPTI VONTAGE | 846700A | HAFIZ | RADIOLOGY | NOT WORKING | RESTARTED THE MACHINE. FOLLOWED THE INSTRUCTIONS IN THE SELF TEST. NOW, WORKING GOOD. | NO | NO | 2023-08-11 | August 2023 | 19:15 | 2023-08-11 19:00 | 2023-08-11 19:15 | 0:15 | COMPLETED | SEMI-CRITICAL
219 | 2023-08-11 | Aug 2023 | 21:50 | 21:55 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | ICU | BRAKE NOT WORKING | FOUND THAT A SCREW WAS MISSING AT ONE SIDE. THE OTHER SIDE'S SCREW WAS LOOSE. TIGHTENED IT. NOW, PARTIALLY WORKING | NO | SCREW NEEDED | 2023-08-11 | August 2023 | 22:06 | 2023-08-11 21:50 | 2023-08-11 22:06 | 0:16 | PENDING | NON-CRITICAL
220 | 2023-08-11 | Aug 2023 | 23:30 | 23:34 | 0.0027778 | HUMIDIFIER AIR | FISHER AND PAYKEL | MR850 | HAFIZ | ICU | ALARM INDICATION | FOUND THAT WATER IS BELOW THE MARKED LEVEL AND LOW LEVEL WATER INDICATION ALARM IS BLINKING. FILLED THE WATER LEVEL TO THE NORMAL RANGE. | NO | NO | 2023-08-11 | August 2023 | 23:45 | 2023-08-11 23:30 | 2023-08-11 23:45 | 0:15 | COMPLETED | SEMI-CRITICAL
221 | 2023-08-11 | Aug 2023 | 14:20 | 14:30 | 0.0069444 | CENTRIFUGE TABLE TOP | REMI | NEYA 4 | RAMKIE | LABORATORY | ABNORMAL NOISE | CHECKED AND COMP PERSON CONFIMED THE PROBLEM WITH MOTOR SHAFT | NO | MOTOR | 2023-08-11 | August 2023 | 14:40 | 2023-08-11 14:20 | 2023-08-11 14:40 | 0:20 | PENDING | SEMI-CRITICAL
222 | 2023-08-11 | Aug 2023 | 14:55 | 15:05 | 0.0069444 | CCO MONITOR | EDWARDS | HEMOSPHERE HEM1 | RAMKIE | ICU | NOT DETECTING THE OXIMETRY CABLE | NEED TO REPLACE THE NEW CABLE | NO | OXIMETRY CABLE | 2023-08-11 | August 2023 | 15:20 | 2023-08-11 14:55 | 2023-08-11 15:20 | 0:25 | PENDING | CRITICAL
223 | 2023-08-12 | Aug 2023 | 11:00 | 11:05 | 0.0034722 | BABY WEIGHING SCALE WITH INFANTOMETER | SECA | 334 | JERLIN | OPD | NOT WORKING | CHECKED THE MACHINE AND FOUND IT'S WORKING GOOD. | NO | NO | 2023-08-12 | August 2023 | 11:10 | 2023-08-12 11:00 | 2023-08-12 11:10 | 0:10 | COMPLETED | NON-CRITICAL
224 | 2023-08-12 | Aug 2023 | 10:15 | 10:20 | 0.0034722 | BABY WEIGHING SCALE WITH INFANTOMETER | SECA | 334 | SANTHOSH | OPD | NOT WORKING | FOUND BATTERY DRY,REPLACED NEW BATTERIES | NO | NO | 2023-08-12 | August 2023 | 10:30 | 2023-08-12 10:15 | 2023-08-12 10:30 | 0:15 | COMPLETED | NON-CRITICAL
225 | 2023-08-12 | Aug 2023 | 13:10 | 13:15 | 0.0034722 | PLASMA STERILIZER | JOHNSON&JOHNSON | STERRAD 100NX | SOORAJ | CSSD | H202 ADJUSTMENT MESSAGE | MACHINE RESTARTED AND ISSUE RECTIFIED | NO | NO | 2023-08-12 | August 2023 | 13:30 | 2023-08-12 13:10 | 2023-08-12 13:30 | 0:20 | COMPLETED | SEMI-CRITICAL
226 | 2023-08-12 | Aug 2023 | 13:35 | 13:35 | 0 | TABLE TOP STEAM STERILIZER | GETINGE | HS 22K7 | SOORAJ | CSSD | SERVICE MESSAGE | WATER CHANGED AND BOWIE DICK TEST DONE AND PASSED. ISSUE RECTIFIED. | NO | NO | 2023-08-12 | August 2023 | 13:45 | 2023-08-12 13:35 | 2023-08-12 13:45 | 0:10 | COMPLETED | SEMI-CRITICAL
227 | 2023-08-13 | Aug 2023 | 16:00 | 16:10 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | ICU | BRAKE LEVER BROKEN | FIXED IT AND ISSUE RECTIFIED | NO | NO | 2023-08-13 | August 2023 | 16:20 | 2023-08-13 16:00 | 2023-08-13 16:20 | 0:20 | COMPLETED | NON-CRITICAL
228 | 2023-08-13 | Aug 2023 | 16:00 | 16:10 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 450 | JERLIN | ICU | SPO2 NOT WORKING | OUTER SHEATH WAS WORNOUT , FIXED IT AND FOUND IN GOOD WORKING CONDITION. | NO | NO | 2023-08-13 | August 2023 | 16:15 | 2023-08-13 16:00 | 2023-08-13 16:15 | 0:15 | COMPLETED | SEMI-CRITICAL
229 | 2023-08-14 | Aug 2023 | 12:00 | 12:15 | 0.010417 | RADIOGRAPHIC UNIT | CARESTREAM | DRF ASCEND ANOLOG HF | JERLIN | RADIOLOGY | BEAM IR FAULT | REFIXED THE CONNECTORS IN GENERATOR AND ISSUE RECTIFIED. | NO | NO | 2023-08-14 | August 2023 | 12:30 | 2023-08-14 12:00 | 2023-08-14 12:30 | 0:30 | COMPLETED | SEMI-CRITICAL
230 | 2023-08-14 | Aug 2023 | 7:45 | 7:50 | 0.0034722 | OT TABLE | MAQUET | ALPHA CLASSIC PRO | JERLIN | OT | FREE MOVEMENT IN ARM ATTACHMENT | TIGHTENED THE ATTACHMENT AND RECTIFIED THE ISSUE | NO | NO | 2023-08-14 | August 2023 | 8:00 | 2023-08-14 07:45 | 2023-08-14 08:00 | 0:15 | COMPLETED | NON-CRITICAL
231 | 2023-08-14 | Aug 2023 | 11:30 | 11:40 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | JERLIN | RADIOLOGY | WORKLIST NOT OPENING FOR A LONG TIME | FOUND LOTS OF PENDING FILES IN SPOOLER. CLEARED IT AND NOW ITS WORKING GOOD. | NO | NO | 2023-08-14 | August 2023 | 11:55 | 2023-08-14 11:30 | 2023-08-14 11:55 | 0:25 | COMPLETED | SEMI-CRITICAL
232 | 2023-08-14 | Aug 2023 | 15:05 | 15:08 | 0.0020833 | IMMUNOASSY AUTO ANOLYZER | BIOMERIUX | MINIVIDAS | HAFIZ | LABORATORY | ERROR MESSAGE DISPLAYING TRAY MECHANISM FAILURE | FOUND THAT A SAMPLE WAS STRUCT IN ONE OF THE TRAY. REMOVED IT, CLEANED THE TRAY. NOW, WORKING FINE. | NO | NO | 2023-08-14 | August 2023 | 15:25 | 2023-08-14 15:05 | 2023-08-14 15:25 | 0:20 | COMPLETED | SEMI-CRITICAL
233 | 2023-08-14 | Aug 2023 | 19:53 | 19:55 | 0.0013889 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | HAFIZ | WARD | NOT WORKING | CHECKED THE MONITOR AND FOUND THAT AC POWER SUPPLY NOT TURNED ON. TURNED ON THE AC SUPPLY AND NOW WORKING GOOD. | NO | NO | 2023-08-14 | August 2023 | 20:00 | 2023-08-14 19:53 | 2023-08-14 20:00 | 0:07 | COMPLETED | SEMI-CRITICAL
234 | 2023-08-14 | Aug 2023 | 20:00 | 20:05 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | ICU | NOT WORKING | FOUND THAT THE BED IS LOCKED. UNLOCKED THE BED AND NOW, WORKING GOOD. | NO | NO | 2023-08-14 | August 2023 | 20:10 | 2023-08-14 20:00 | 2023-08-14 20:10 | 0:10 | COMPLETED | NON-CRITICAL
235 | 2023-08-14 | Aug 2023 | 20:15 | 20:20 | 0.0034722 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | HAFIZ | WARD | SPO2 PROBE NOT WORKING | FOUND THAT THE PINS ARE BEND. RECTIFIED AND CHECKED WITH THE MONITOR. NOW, WORKING GOOD. | NO | NO | 2023-08-14 | August 2023 | 20:25 | 2023-08-14 20:15 | 2023-08-14 20:25 | 0:10 | COMPLETED | SEMI-CRITICAL
236 | 2023-08-14 | Aug 2023 | 16:10 | 16:30 | 0.013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | ICU | NOT WORKING | RESET DONE, NOW IT WORKING | NO | NO | 2023-08-14 | August 2023 | 16:40 | 2023-08-14 16:10 | 2023-08-14 16:40 | 0:30 | COMPLETED | NON-CRITICAL
237 | 2023-08-15 | Aug 2023 | 14:10 | 14:20 | 0.0069444 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | RAMKIE | ICU | NOT WORKING | RESET DONE, NOW IT WORKING | NO | NO | 2023-08-15 | August 2023 | 14:50 | 2023-08-15 14:10 | 2023-08-15 14:50 | 0:40 | COMPLETED | SEMI-CRITICAL
238 | 2023-08-16 | Aug 2023 | 7:30 | 7:35 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | JERLIN | CATHLAB | AC POWER FAIL ISSUE | REPLACED PMB BOARD (STANDBY) AND RECTIFIED THE ISSUE | PMB BOARD | NO | 2023-08-16 | August 2023 | 8:20 | 2023-08-16 07:30 | 2023-08-16 08:20 | 0:50 | COMPLETED | CRITICAL
239 | 2023-08-16 | Aug 2023 | 7:15 | 7:20 | 0.0034722 | OT TABLE | MAQUET | MEERA | JERLIN | OT | REMOTE NOT WORKING | REFIXED THE REMOTE AND IT'S WORKING GOOD | NO | NO | 2023-08-16 | August 2023 | 7:25 | 2023-08-16 07:15 | 2023-08-16 07:25 | 0:10 | COMPLETED | NON-CRITICAL
240 | 2023-08-16 | Aug 2023 | 11:45 | 11:50 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SANTHOSH | ICU | REMOTE DOESNT WORK | REFIXED THE CABLE | NO | NO | 2023-08-16 | August 2023 | 12:00 | 2023-08-16 11:45 | 2023-08-16 12:00 | 0:15 | COMPLETED | NON-CRITICAL
241 | 2023-08-16 | Aug 2023 | 14:25 | 14:30 | 0.0034722 | TABLE TOP STEAM STERILIZER | GETINGE | HS 22K7 | HAFIZ | CSSD | ERROR MESSAGE DISPLAYING SERVICE | CHECKED THE MACHINE AND RUNNED WITH ONE SAMPLE. NO SUCH MESSAGES DISPLAYED. NOW, WORKING GOOD. | NO | NO | 2023-08-16 | August 2023 | 14:50 | 2023-08-16 14:25 | 2023-08-16 14:50 | 0:25 | COMPLETED | SEMI-CRITICAL
242 | 2023-08-16 | Aug 2023 | 16:35 | 16:40 | 0.0034722 | VENTILATOR, ADULT | MAQUET | SERVO-I | HAFIZ | ICU | PRE USE CHECK FAILED | FOUND THAT O2 CELL DEFECTIVE. REPLACED IT | O2 CELL | NO | 2023-08-16 | August 2023 | 16:45 | 2023-08-16 16:35 | 2023-08-16 16:45 | 0:10 | COMPLETED | CRITICAL
243 | 2023-08-16 | Aug 2023 | 18:10 | 18:20 | 0.0069444 | VENTILATOR, ADULT | MAQUET | SERVO-I | RAMKIE | ICU | INTERNOL LEAKAGE TEST FAILED | CLEANED EXP CASSETTE AND IT WORKING | NO | NO | 2023-08-16 | August 2023 | 18:30 | 2023-08-16 18:10 | 2023-08-16 18:30 | 0:20 | COMPLETED | CRITICAL
244 | 2023-08-17 | Aug 2023 | 10:00 | 10:15 | 0.010417 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | JERLIN | OT | O2 CELL ISSUE | CHANGED THE O2 SENSOR AND CHECKED. IT'S WOKING GOOD. | O2 SENSOR | NO | 2023-08-17 | August 2023 | 10:30 | 2023-08-17 10:00 | 2023-08-17 10:30 | 0:30 | COMPLETED | CRITICAL
245 | 2023-08-17 | Aug 2023 | 11:30 | 11:40 | 0.0069444 | VENTILATOR, PEDIATRIC | MINDRAY MEDICAL INDIA | SV 300 | SANTHOSH | EMERGENCY | DEVICE FAILURE 13 | RUN THE PRE-USE CHECK ,AND RECTIFIED | NO | NO | 2023-08-17 | August 2023 | 12:00 | 2023-08-17 11:30 | 2023-08-17 12:00 | 0:30 | COMPLETED | CRITICAL
246 | 2023-08-17 | Aug 2023 | 12:00 | 12:10 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | SANTHOSH | RADIOLOGY | DOESNT PUSH TO PACS | REFIXED THE CABLE & CLEARED SPOOLER | NO | NO | 2023-08-17 | August 2023 | 12:10 | 2023-08-17 12:00 | 2023-08-17 12:10 | 0:10 | COMPLETED | SEMI-CRITICAL
247 | 2023-08-17 | Aug 2023 | 20:30 | 20:35 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | WARD | ECG CABLE NOT WORKING | CHECKED THE CABLE AND FOUND THAT NOT WORKING. CUURENTLY REPLACED WITH THE ANOTHER WARDSIDE ECG CABLE. NOW, WORKING FINE. | NO | NO | 2023-08-17 | August 2023 | 20:45 | 2023-08-17 20:30 | 2023-08-17 20:45 | 0:15 | COMPLETED | SEMI-CRITICAL
248 | 2023-08-17 | Aug 2023 | 20:47 | 20:50 | 0.0020833 | PLASMA STERILIZER | JOHNSON&JOHNSON | STERRAD 100NX | HAFIZ | CSSD | ERROR - E15 FAILED | RESTARTED AND RUNNED WITH ONE SAMPLE. NOW, WORKING FINE. | NO | NO | 2023-08-17 | August 2023 | 21:00 | 2023-08-17 20:47 | 2023-08-17 21:00 | 0:13 | COMPLETED | SEMI-CRITICAL
249 | 2023-08-17 | Aug 2023 | 21:00 | 21:05 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | HAFIZ | ICU | PRINTER NOT WORKING | REFIXED THE PAPER PROPERLY. NOW, WORKING FINE. | NO | NO | 2023-08-17 | August 2023 | 21:10 | 2023-08-17 21:00 | 2023-08-17 21:10 | 0:10 | COMPLETED | SEMI-CRITICAL
250 | 2023-08-17 | Aug 2023 | 21:12 | 21:20 | 0.0055556 | GO FLO PUMP, ARTHROSCOPY | SMITH & NEPHEW | 72204968 | HAFIZ | OT | NOT WORKING DURING THE PROCEDURE | CHECKED THE MACHINE AND FOUND THAT SELF TEST NEEDED. REMOVED THE TUBE AND THEN RECONNECTED. NOW, WORKING GOOD. | NO | NO | 2023-08-17 | August 2023 | 21:27 | 2023-08-17 21:12 | 2023-08-17 21:27 | 0:15 | COMPLETED | SEMI-CRITICAL
251 | 2023-08-17 | Aug 2023 | 21:30 | 21:32 | 0.0013889 | VENTILATOR, TRANSPORT | RESMED | ASTRAL 100 | HAFIZ | WARD | PERSISTENT ALARM INDICATION | CHANGED THE SETTINGS, UNLOCKED THE SCREEN, AND NO ALARM INDICATION. SELF TEST HAS TO BE DONE AFTER USE. NOW, WORKING FINE. | NO | NO | 2023-08-17 | August 2023 | 21:40 | 2023-08-17 21:30 | 2023-08-17 21:40 | 0:10 | COMPLETED | CRITICAL
252 | 2023-08-17 | Aug 2023 | 23:10 | 23:14 | 0.0027778 | BIPAP MACHINE | PHILIPS | A40 | HAFIZ | ICU | HIGH INTERNOL O2 ERROR | REMOVED THE CIRCUITS AND REFIXED. NOW, WORKING GOOD. | NO | NO | 2023-08-17 | August 2023 | 23:20 | 2023-08-17 23:10 | 2023-08-17 23:20 | 0:10 | COMPLETED | SEMI-CRITICAL
253 | 2023-08-18 | Aug 2023 | 11:00 | 11:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | ICU | THERAPY TEST FAILED | RE-RAN OP CHECK AND EVERY TEST PASSED. | NO | NO | 2023-08-18 | August 2023 | 11:15 | 2023-08-18 11:00 | 2023-08-18 11:15 | 0:15 | COMPLETED | CRITICAL
254 | 2023-08-18 | Aug 2023 | 19:00 | 19:15 | 0.010417 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | WARD | NOT WORKING | PLUGED IN AC POWER AND RELEASED THE LOCKS IN REMOTE. NOW WORKING GOOD. | NO | NO | 2023-08-18 | August 2023 | 19:30 | 2023-08-18 19:00 | 2023-08-18 19:30 | 0:30 | COMPLETED | NON-CRITICAL
255 | 2023-08-18 | Aug 2023 | 14:00 | 14:05 | 0.0034722 | BABY WEIGHING SCALE | SECA | 354 | SANTHOSH | ICU | NOT WORKING | CHECKED BATTERY POSITION | NO | NO | 2023-08-18 | August 2023 | 14:10 | 2023-08-18 14:00 | 2023-08-18 14:10 | 0:10 | COMPLETED | NON-CRITICAL
256 | 2023-08-18 | Aug 2023 | 0:00 | 0:10 | 0.0069444 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | HAFIZ | ICU | CATRIDGE NEEDED TO BE REPLACED | A NEW CATRIDGE REPLACED | CARTRIDGE | NO | 2023-08-18 | August 2023 | 0:30 | 2023-08-18 | 2023-08-18 00:30 | 0:30 | COMPLETED | SEMI-CRITICAL
257 | 2023-08-19 | Aug 2023 | 11:20 | 11:25 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | CL CVP FAILED | LOT CHANGED AND AGAIN RUN CVP | NO | NO | 2023-08-19 | August 2023 | 12:00 | 2023-08-19 11:20 | 2023-08-19 12:00 | 0:40 | COMPLETED | SEMI-CRITICAL
258 | 2023-08-19 | Aug 2023 | 0:42 | 0:45 | 0.0020833 | USG MACHINE | MINDRAY MEDICAL INDIA | MX7 | HAFIZ | ICU | 12L PROBE NOT DETECTING | CHECKED AND FOUND THAT PROBE IS NOT FIXED PROPERLY. UNLOCKED, REMOVED, FIXED PROPERLY AND THEN LOCKED. NOW, WORKING GOOD. | NO | NO | 2023-08-19 | August 2023 | 0:50 | 2023-08-19 00:42 | 2023-08-19 00:50 | 0:08 | COMPLETED | SEMI-CRITICAL
259 | 2023-08-20 | Aug 2023 | 10:30 | 10:35 | 0.0034722 | BMI MACHINE | AKAS MEDICAL | QVS-100 | SANTHOSH | OPD | BP CUFF WALCARO DAMAGED | REPLACED BP CUFF BLADDER | BP CUFF BLADDER | NO | 2023-08-20 | August 2023 | 10:45 | 2023-08-20 10:30 | 2023-08-20 10:45 | 0:15 | COMPLETED | NON-CRITICAL
260 | 2023-08-20 | Aug 2023 | 13:20 | 13:22 | 0.0013889 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SANTHOSH | ICU | BATTERY FAIL | OP CHECK RUN STILL ISSUE PERSIST, REPLACED A BATTERY | BATTERY | NO | 2023-08-20 | August 2023 | 13:30 | 2023-08-20 13:20 | 2023-08-20 13:30 | 0:10 | COMPLETED | CRITICAL
261 | 2023-08-21 | Aug 2023 | 21:00 | 21:10 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | CARTRIDGE EXPIRED | CHANGED THE CARTRIDGE WITH NEW ONE AND CVP DONE | CARTRIDGE | NO | 2023-08-21 | August 2023 | 21:30 | 2023-08-21 21:00 | 2023-08-21 21:30 | 0:30 | COMPLETED | SEMI-CRITICAL
262 | 2023-08-21 | Aug 2023 | 19:00 | 19:15 | 0.010417 | DIALYSIS MACHINE | FRESENIUS | 4008 S NG | JERLIN | DIALYSIS | NOT WORKING | MOTOR FAULT. REPLACED WITH A NEW ONE AND RECTIFIED THE ISSUE. | MOTOR | NO | 2023-08-21 | August 2023 | 19:40 | 2023-08-21 19:00 | 2023-08-21 19:40 | 0:40 | COMPLETED | CRITICAL
263 | 2023-08-21 | Aug 2023 | 11:10 | 11:15 | 0.0034722 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | SANTHOSH | WARD | OCCLUSION ALARM | REMOVED OCCLUSION | NO | NO | 2023-08-21 | August 2023 | 11:15 | 2023-08-21 11:10 | 2023-08-21 11:15 | 0:05 | COMPLETED | NON-CRITICAL
264 | 2023-08-21 | Aug 2023 | 8:30 | 8:40 | 0.0069444 | ORTHO VISION ANALYSER | ORTHO VISION | ORTHO VISION SWIFT | HAFIZ | BLOODBANK | DAILY PROBE MAINTENONCE TEST FAILED. | FOUND PROBLEM IN SALINE TANK, FIXED IT. RESTARTED THE MACHINE, RUNNED THE DAILY PROBE MAINTENONCE TEST. NOW WORKING GOOD. | NO | NO | 2023-08-21 | August 2023 | 9:00 | 2023-08-21 08:30 | 2023-08-21 09:00 | 0:30 | COMPLETED | SEMI-CRITICAL
265 | 2023-08-21 | Aug 2023 | 10:35 | 10:40 | 0.0034722 | HBA1C ANALYSER | ADAMS A1CLITE HA8380V | ARKRAY | SOORAJ | LABORATORY | ABNORMAL DETECTION MESSAGE | RAN PRIMING AND FOUND ISUE TO BE RECTIFIED. | NO | NO | 2023-08-21 | August 2023 | 10:50 | 2023-08-21 10:35 | 2023-08-21 10:50 | 0:15 | COMPLETED | SEMI-CRITICAL
266 | 2023-08-22 | Aug 2023 | 23:30 | 23:35 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | ICU | AIR IN LINE ALARM | FOUND AIR BUBBLES IN THE IV LINE. CLEARED THE LINE AND RECTIFIED THE ISSUE | NO | NO | 2023-08-22 | August 2023 | 23:40 | 2023-08-22 23:30 | 2023-08-22 23:40 | 0:10 | COMPLETED | NON-CRITICAL
267 | 2023-08-22 | Aug 2023 | 22:00 | 22:10 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | WARD | NOT WORKING | REMOTE WAS LOCKED. RELEASED THE LOCK AND RECTIFIED THE ISSUE. | NO | NO | 2023-08-22 | August 2023 | 22:15 | 2023-08-22 22:00 | 2023-08-22 22:15 | 0:15 | COMPLETED | NON-CRITICAL
268 | 2023-08-22 | Aug 2023 | 11:30 | 11:33 | 0.0020833 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SANTHOSH | WARD | CONTROLPAD NOT WORKING | RELEASED CONTROL BOX LOCK | NO | NO | 2023-08-22 | August 2023 | 11:45 | 2023-08-22 11:30 | 2023-08-22 11:45 | 0:15 | COMPLETED | NON-CRITICAL
269 | 2023-08-22 | Aug 2023 | 8:30 | 8:35 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | HAFIZ | OT | GAS MODULE NOT WORKING | REPLACED A NEW WATER TRAP, RESTARTED THE MACHINE. NOW, WORKING FINE. | WATER TRAP | NO | 2023-08-22 | August 2023 | 8:50 | 2023-08-22 08:30 | 2023-08-22 08:50 | 0:20 | COMPLETED | SEMI-CRITICAL
270 | 2023-08-23 | Aug 2023 | 1:30 | 1:40 | 0.0069444 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | JERLIN | WARD | RESPIPAROTY RATE NOT DISPLAYED | CHANGED THE SETTINGS AND RECTIFIED THE ISSUE | NO | NO | 2023-08-23 | August 2023 | 1:50 | 2023-08-23 01:30 | 2023-08-23 01:50 | 0:20 | COMPLETED | SEMI-CRITICAL
271 | 2023-08-23 | Aug 2023 | 3:30 | 3:30 | 0 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | JERLIN | OT | CIRCUIT LEAK TEST FAILED | FOUND THAT CANISTER IS ACCUMULATED WITH LIQUID. CLEANED AND REFIXED THE CANISTER. NOW IT'S WORKING GOOD. | NO | NO | 2023-08-23 | August 2023 | 3:45 | 2023-08-23 03:30 | 2023-08-23 03:45 | 0:15 | COMPLETED | CRITICAL
272 | 2023-08-23 | Aug 2023 | 13:10 | 13:15 | 0.0034722 | EXAMINATION COUCH | MIDMARK JANOK | E122012 | SANTHOSH | OPD | WORKING SLOW WITH BEEP NOISE | FOUND BATTERY WAS LOW NOW KEEP CHARGED | NO | NO | 2023-08-23 | August 2023 | 13:25 | 2023-08-23 13:10 | 2023-08-23 13:25 | 0:15 | COMPLETED | NON-CRITICAL
273 | 2023-08-23 | Aug 2023 | 11:45 | 11:50 | 0.0034722 | USG MACHINE | MINDRAY MEDICAL INDIA | MX7 | SOORAJ | OPD | IMAGE NOT PUSING TO PACS. | PACS CONFIGURATION AND IP ADDRESS RECONFIGURED AND ISSUE RECTIFIED. | NO | NO | 2023-08-23 | August 2023 | 12:00 | 2023-08-23 11:45 | 2023-08-23 12:00 | 0:15 | COMPLETED | SEMI-CRITICAL
274 | 2023-08-23 | Aug 2023 | 16:50 | 16:55 | 0.0034722 | EEG MACHINE | NOTUS NEUROLOGY | NICOLET EEG V32 | SOORAJ | OPD | PRINT REPORT NOT PROPER. | FOUND ISSUE WITH PRINTER TONER. INFORMED TECHNICAINS TO INDENT TONER FROM CRS | NO | NO | 2023-08-23 | August 2023 | 17:05 | 2023-08-23 16:50 | 2023-08-23 17:05 | 0:15 | COMPLETED | SEMI-CRITICAL
275 | 2023-08-24 | Aug 2023 | 23:00 | 23:15 | 0.010417 | BIPAP MACHINE | PHILIPS | A40 | JERLIN | ICU | LOW CIRCUIT LEAK ALARM | REFIXED THE CIRCUIT AND CLEARED THE EXPIRATORY VALVE. ISSUE RECTIFIED | NO | NO | 2023-08-24 | August 2023 | 23:25 | 2023-08-24 23:00 | 2023-08-24 23:25 | 0:25 | COMPLETED | SEMI-CRITICAL
276 | 2023-08-24 | Aug 2023 | 1:20 | 1:30 | 0.0069444 | PATIENT MONITOR | PHILIPS | CM12 | JERLIN | WARD | HEART RATE NOT DISPLAYED | CHANGED THE SETTINGS ACCORDINGLY AND RECTIFIED THE ISSUE. | NO | NO | 2023-08-24 | August 2023 | 1:40 | 2023-08-24 01:20 | 2023-08-24 01:40 | 0:20 | COMPLETED | SEMI-CRITICAL
277 | 2023-08-24 | Aug 2023 | 2:00 | 2:10 | 0.0069444 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | JERLIN | WARD | NIBP NOT WORKING | FOUND THAT BP CUFF WAS NOT PROPERLY CONNECTED.AFTER CONNECTING IT PROPERLY, IT'S WORKING GOOD. | NO | NO | 2023-08-24 | August 2023 | 2:15 | 2023-08-24 02:00 | 2023-08-24 02:15 | 0:15 | COMPLETED | SEMI-CRITICAL
278 | 2023-08-24 | Aug 2023 | 5:00 | 5:15 | 0.010417 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | JERLIN | OT | NO FLOW SENSOR ALARM | BLOOD WAS CONTAINED IN FLOW SENSORS AND CANISTER. CLEANED AND REFIXED AND RECTIFIED THE ISSUE. | NO | NO | 2023-08-24 | August 2023 | 5:30 | 2023-08-24 05:00 | 2023-08-24 05:30 | 0:30 | COMPLETED | CRITICAL
279 | 2023-08-24 | Aug 2023 | 5:30 | 5:35 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | JERLIN | OT | ETCO2 NOT WORKING | CHANGED THE WATER TRAP AND FOUND IT'S IN GOOD WORKING CONDITION. | WATER TRAP | NO | 2023-08-24 | August 2023 | 5:45 | 2023-08-24 05:30 | 2023-08-24 05:45 | 0:15 | COMPLETED | SEMI-CRITICAL
280 | 2023-08-24 | Aug 2023 | 8:00 | 8:05 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | JERLIN | WARD | PRODUCING NOISE AND NOT WORKING | CONNECTED WITH AC POWER , REFIXED THE TEMP PROBE AND FOUND THAT IT'S WORKING GOOD. | NO | NO | 2023-08-24 | August 2023 | 8:15 | 2023-08-24 08:00 | 2023-08-24 08:15 | 0:15 | COMPLETED | SEMI-CRITICAL
281 | 2023-08-24 | Aug 2023 | 22:10 | 22:15 | 0.0034722 | FETAL HEART MONITOR | WIPRO GE HEALTH CARE | COROMETRICS 174 | JERLIN | WARD | DISPLAYS WRONG TIME | CHANGED THE TIME AND RECTIFIED THE ISSUE | NO | NO | 2023-08-24 | August 2023 | 22:25 | 2023-08-24 22:10 | 2023-08-24 22:25 | 0:15 | COMPLETED | SEMI-CRITICAL
282 | 2023-08-24 | Aug 2023 | 8:00 | 8:05 | 0.0034722 | HEART LUNG MACHINE | GETINGE | HL20-4PUMPS | HAFIZ | OT | ER20 ERROR MESSAGE | FOUND THAT A PROBLEM IN CONSOLE. TALKED TO THE COMPANY PERSON AND HE SUGGESTED THAT THE PROBLEM MIGHT BE DUE TO HUMIDITY. FOLLOWED HIS INSTRUCTIONS. REMOVED THE CONSOLE UNIT BOARD, KEPT IN NORMAL TEMPERATURE, BUT ISSUE PERSISTS. SO REPLACED A NEW CCM MODULE | CONTROL CONSOLE MODULE | NO | 2023-08-24 | August 2023 | 8:40 | 2023-08-24 08:00 | 2023-08-24 08:40 | 0:40 | COMPLETED | CRITICAL
283 | 2023-08-24 | Aug 2023 | 9:20 | 9:25 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | HAFIZ | WARD | NOT WORKING | FOUND THAT BATTERY DRAINED FULLY. CONNECTED TO AC POWER SUPPLY. NOW WORKING GOOD. | NO | NO | 2023-08-24 | August 2023 | 9:40 | 2023-08-24 09:20 | 2023-08-24 09:40 | 0:20 | COMPLETED | NON-CRITICAL
284 | 2023-08-24 | Aug 2023 | 12:10 | 12:30 | 0.013889 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | RAMKIE | ICU | OCCLUSION ERROR | ASKED TO DILITION AND CHANGE THE OCCULUSION SESNSIVITY | NO | NO | 2023-08-24 | August 2023 | 12:45 | 2023-08-24 12:10 | 2023-08-24 12:45 | 0:35 | COMPLETED | NON-CRITICAL
285 | 2023-08-24 | Aug 2023 | 14:10 | 14:20 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | RAMKIE | CATHLAB | ERROR NRTU | ECG CABLE RECONNECTED AND OP TEST PASSED | NO | NO | 2023-08-24 | August 2023 | 14:30 | 2023-08-24 14:10 | 2023-08-24 14:30 | 0:20 | COMPLETED | CRITICAL
286 | 2023-08-25 | Aug 2023 | 21:30 | 21:40 | 0.0069444 | BIPAP MACHINE | PHILIPS | A40 | JERLIN | ICU | NEED TO CHANGE THE MODE | CHANGED THE SETTINGS ACCORDINGLY | NO | NO | 2023-08-25 | August 2023 | 21:45 | 2023-08-25 21:30 | 2023-08-25 21:45 | 0:15 | COMPLETED | SEMI-CRITICAL
287 | 2023-08-25 | Aug 2023 | 14:20 | 14:30 | 0.0069444 | BMI MACHINE | SECA | 769 | SANTHOSH | DIALYSIS | SHOWS 1.19 ERROR | REFIXED POWER CABLE | NO | NO | 2023-08-25 | August 2023 | 14:40 | 2023-08-25 14:20 | 2023-08-25 14:40 | 0:20 | COMPLETED | NON-CRITICAL
288 | 2023-08-25 | Aug 2023 | 8:30 | 8:35 | 0.0034722 | BIOCHEMISTRY ANALYSER | COBAS C 501 | ROCHE | HAFIZ | LABORATORY | PROBE UP/DOWN FAILURE | RESET DONE. STILL ISSUE PERSISTS. AS PER THE COMPANY PERSON'S GUIDE, REMOVED THE PROBE, CLEANED AND REFIXED. NOW, WORKING GOOD. | NO | NO | 2023-08-25 | August 2023 | 9:00 | 2023-08-25 08:30 | 2023-08-25 09:00 | 0:30 | COMPLETED | SEMI-CRITICAL
289 | 2023-08-25 | Aug 2023 | 9:05 | 9:08 | 0.0020833 | CAMERA SYSTEM,ICG | KARL STORZ | TC200 | HAFIZ | OT | ICG MACHINE NOT WORKING | REFIXED THE CABLE PROPERLY. NOW, WORKING GOOD. | NO | NO | 2023-08-25 | August 2023 | 9:15 | 2023-08-25 09:05 | 2023-08-25 09:15 | 0:10 | COMPLETED | SEMI-CRITICAL
290 | 2023-08-25 | Aug 2023 | 13:10 | 13:20 | 0.0069444 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | RAMKIE | OPD | DISTURBANCE IN V3 | PROBLEM WITH SO REPLACED THE CABLE | CABLE CONNECTOR | NO | 2023-08-25 | August 2023 | 13:30 | 2023-08-25 13:10 | 2023-08-25 13:30 | 0:20 | COMPLETED | SEMI-CRITICAL
291 | 2023-08-25 | Aug 2023 | 10:10 | 10:15 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | SOORAJ | ICU | NOT SWITCHING ON | FOUND POWER CONNECTOR DEFECTIVE. CHANGED PLUG TOP AND ISSUE RECTIFIED. | PLUG TOP | NO | 2023-08-25 | August 2023 | 10:30 | 2023-08-25 10:10 | 2023-08-25 10:30 | 0:20 | COMPLETED | CRITICAL
292 | 2023-08-26 | Aug 2023 | 22:40 | 22:45 | 0.0034722 | BIPAP MACHINE | PHILIPS | A40 | JERLIN | EMERGENCY | NEED TO CHANGE BIPAP MACHINE SETTINGS | CHANGED THE SETTINGS ACCODING TO THE PHYSICIAN'S INFO | NO | NO | 2023-08-26 | August 2023 | 22:55 | 2023-08-26 22:40 | 2023-08-26 22:55 | 0:15 | COMPLETED | SEMI-CRITICAL
293 | 2023-08-26 | Aug 2023 | 14:00 | 14:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | RAMKIE | ICU | ERROR NOT READY TO USE | PROBLEM WITH SO REPLACED THE CABLE | ECG CABLE | NO | 2023-08-26 | August 2023 | 14:20 | 2023-08-26 14:00 | 2023-08-26 14:20 | 0:20 | COMPLETED | CRITICAL
294 | 2023-08-27 | Aug 2023 | 10:30 | 10:34 | 0.0027778 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC V2 | SANTHOSH | EMERGENCY | PRINTER OFFLINE ERROR | REMOVED PRINTER DETAILS IN UTILITY AND RESTARTED | NO | NO | 2023-08-27 | August 2023 | 10:35 | 2023-08-27 10:30 | 2023-08-27 10:35 | 0:05 | COMPLETED | SEMI-CRITICAL
295 | 2023-08-28 | Aug 2023 | 17:45 | 18:00 | 0.010417 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 9000 | JERLIN | ICU | ERROR 410 | CHANGED THE POWER SUPPLY UNIT AND RECTIFIED THE ISSUE | POWER SUPPLY BOARD | NO | 2023-08-28 | August 2023 | 18:15 | 2023-08-28 17:45 | 2023-08-28 18:15 | 0:30 | COMPLETED | NON-CRITICAL
296 | 2023-08-28 | Aug 2023 | 16:00 | 16:15 | 0.010417 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | JERLIN | CATHLAB | COOLING FAN FAILED | FOUND A PIECE OF FAN IS BROKEN. NEED TO BE REPLACED,REPLACED IT | NO | NO | 2023-08-28 | August 2023 | 16:50 | 2023-08-28 16:00 | 2023-08-28 16:50 | 0:50 | COMPLETED | CRITICAL
297 | 2023-08-28 | Aug 2023 | 19:10 | 19:15 | 0.0034722 | PRINTER,COMPUTED RADIOGRAPHY | CARESTREAM | DRY VIEW 6950 | SANTHOSH | RADIOLOGY | SERVICE DUE | RESET DONE | NO | NO | 2023-08-28 | August 2023 | 19:25 | 2023-08-28 19:10 | 2023-08-28 19:25 | 0:15 | COMPLETED | SEMI-CRITICAL
298 | 2023-08-28 | Aug 2023 | 8:00 | 8:10 | 0.0069444 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX | RAMKIE | OT | ERROR 212 | NEED TO REPLACE THE POWER BOARD | NO | POWER BOARD | 2023-08-28 | August 2023 | 8:20 | 2023-08-28 08:00 | 2023-08-28 08:20 | 0:20 | PENDING | SEMI-CRITICAL
299 | 2023-08-28 | Aug 2023 | 7:00 | 7:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | RAMKIE | OT | CIRCUIT LEAK TEST FAILED | CLEANED SOME SODA LIME AND TEST PASSED | NO | NO | 2023-08-28 | August 2023 | 7:20 | 2023-08-28 07:00 | 2023-08-28 07:20 | 0:20 | COMPLETED | CRITICAL
300 | 2023-08-28 | Aug 2023 | 20:53 | 20:55 | 0.0013889 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | SOORAJ | EMERGENCY | SPO2 NOT SHOWING | FOUND SPO2 SENSOR TO BE DEFECTIVE. REPLACED WITH A SENSOR CABLE FROM ANOTHER MONITOR. | NO | NO | 2023-08-28 | August 2023 | 21:10 | 2023-08-28 20:53 | 2023-08-28 21:10 | 0:17 | COMPLETED | SEMI-CRITICAL
301 | 2023-08-29 | Aug 2023 | 17:30 | 17:45 | 0.010417 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | CARTRIDGE EXPIRED | CHANGED THE CARTRIDGE WITH NEW ONE AND CVP DONE | CARTRIDGE | NO | 2023-08-29 | August 2023 | 18:15 | 2023-08-29 17:30 | 2023-08-29 18:15 | 0:45 | COMPLETED | SEMI-CRITICAL
302 | 2023-08-29 | Aug 2023 | 18:10 | 18:15 | 0.0034722 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | SANTHOSH | EMERGENCY | SPO2 PROBE NOT WORKING | FOUND PROBE PIN DISLOCATED,FIXED NOW | NO | NO | 2023-08-29 | August 2023 | 18:20 | 2023-08-29 18:10 | 2023-08-29 18:20 | 0:10 | COMPLETED | SEMI-CRITICAL
303 | 2023-08-29 | Aug 2023 | 8:25 | 8:30 | 0.0034722 | BIPAP MACHINE | PHILIPS | A40 | HAFIZ | ICU | NEED TO CHANGE THE MODE. | CHANGED THE SETTINGS ACCORDING TO THE PRESCIPTION. | NO | NO | 2023-08-29 | August 2023 | 8:35 | 2023-08-29 08:25 | 2023-08-29 08:35 | 0:10 | COMPLETED | SEMI-CRITICAL
304 | 2023-08-29 | Aug 2023 | 0:20 | 0:23 | 0.0020833 | BIPAP MACHINE | PHILIPS | A40 | SOORAJ | WARD | LOW MINUTE VENTILATION ALARM. | FOUND EXPIRATORY PORT CLOSED. OPENED POPT AND ISSUE RECTIFIED. | NO | NO | 2023-08-29 | August 2023 | 0:35 | 2023-08-29 00:20 | 2023-08-29 00:35 | 0:15 | COMPLETED | SEMI-CRITICAL
305 | 2023-08-29 | Aug 2023 | 1:20 | 1:25 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | FOUND LEAK IN BELLOWS. REFIXED BELLOWS AND CHECKOUT DONE, ISSUE RECTIFIED. | NO | NO | 2023-08-29 | August 2023 | 1:35 | 2023-08-29 01:20 | 2023-08-29 01:35 | 0:15 | COMPLETED | CRITICAL
306 | 2023-08-30 | Aug 2023 | 16:00 | 16:15 | 0.010417 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | DIALYSIS | NOT WORKING | CHANGED THE POWER CORD AND IT'S WORKING GGOD. | POWER CORD | NO | 2023-08-30 | August 2023 | 16:30 | 2023-08-30 16:00 | 2023-08-30 16:30 | 0:30 | COMPLETED | NON-CRITICAL
307 | 2023-08-30 | Aug 2023 | 13:00 | 13:07 | 0.0048611 | VERTICAL AUTOCLAVE | VANI SCIENTIFIC & EQUIPMENTS | NA | HAFIZ | LABORATORY | WATER LEVEL LOW ALARM | ENSURED THAT THE WATER IS FULL AND RESTARTED THE MACHINE. NOW, WORKING GOOD. | NO | NO | 2023-08-30 | August 2023 | 13:15 | 2023-08-30 13:00 | 2023-08-30 13:15 | 0:15 | COMPLETED | SEMI-CRITICAL
308 | 2023-08-30 | Aug 2023 | 13:35 | 13:40 | 0.0034722 | BIPAP MACHINE | PHILIPS | TRILOGY EV300 | HAFIZ | ICU | LOW MINUTE VOLUME ERROR | CHECKED THE MACHINE, RECTIFIED. NOW, WORKING GOOD | NO | NO | 2023-08-30 | August 2023 | 13:50 | 2023-08-30 13:35 | 2023-08-30 13:50 | 0:15 | COMPLETED | SEMI-CRITICAL
309 | 2023-08-30 | Aug 2023 | 1:55 | 2:00 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | CHANGED SODALIME AND CHECKOUT DONE, ISSUE RECTIFIED. | NO | NO | 2023-08-30 | August 2023 | 2:15 | 2023-08-30 01:55 | 2023-08-30 02:15 | 0:20 | COMPLETED | CRITICAL
310 | 2023-08-30 | Aug 2023 | 5:00 | 5:05 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | ICU | COT LOCKED | RESET DONE AND ISSUE RECTIFIED. | NO | NO | 2023-08-30 | August 2023 | 5:15 | 2023-08-30 05:00 | 2023-08-30 05:15 | 0:15 | COMPLETED | NON-CRITICAL
311 | 2023-08-31 | Aug 2023 | 8:30 | 8:35 | 0.0034722 | VENTILATOR, PEDIATRIC | MINDRAY MEDICAL INDIA | SV 300 | HAFIZ | EMERGENCY | NOT WORKING | FOUND O2 FLOW SENSOR TEST AND O2 SENSOR TEST FAILED. REPLACED O2 SENSOR AND STILL SAME RESULT. | NO | O2 SENSOR | 2023-08-31 | August 2023 | 9:00 | 2023-08-31 08:30 | 2023-08-31 09:00 | 0:30 | PENDING | CRITICAL
312 | 2023-08-31 | Aug 2023 | 10:40 | 10:45 | 0.0034722 | VENTILATOR, ADULT | MAQUET | SERVO-I | HAFIZ | ICU | INTERNAL LEAKAGE TEST AND SAFETY VALVE TEST FAILED. | REMOVED THE EXPIRATORY CASSETTE, CLEANED AND REFIXED. NOW, SAFETY VALVE TEST PASSED. INTERNOL LEAKAGE TEST STILL FAILED. | NO | NO | 2023-08-31 | August 2023 | 11:10 | 2023-08-31 10:40 | 2023-08-31 11:10 | 0:30 | COMPLETED | CRITICAL
313 | 2023-08-31 | Aug 2023 | 8:20 | 8:25 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | RAMKIE | OT | O2 VALUE VARIATION | PROBLEM WITH O2 SENSOR CABLE | NO | O2 SENSOR CABLE | 2023-08-31 | August 2023 | 8:40 | 2023-08-31 08:20 | 2023-08-31 08:40 | 0:20 | PENDING | CRITICAL
314 | 2023-08-31 | Aug 2023 | 11:20 | 11:25 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | RAMKIE | ICU | MEMORY FULL | COPIED TO SD CARD | NO | NO | 2023-08-31 | August 2023 | 11:45 | 2023-08-31 11:20 | 2023-08-31 11:45 | 0:25 | COMPLETED | SEMI-CRITICAL
315 | 2023-08-31 | Aug 2023 | 23:45 | 23:50 | 0.0034722 | DIGITAL PLATFORM SCALE | SECA | 676 | SOORAJ | DIALYSIS | ERROR 1.12 | FOUND LOOSE CONNECTOR. SECURED POWER CONNECTOR AND ISSUE RECTIFIED. | NO | NO | 2023-08-31 | August 2023 | 23:59 | 2023-08-31 23:45 | 2023-08-31 23:59 | 0:14 | COMPLETED | NON-CRITICAL
316 | 2023-09-01 | Sep 2023 | 1:00 | 1:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | SOORAJ | OT | O2 CELL ISSUE | DID CALIBRATION, BUT FAILED. SWAPED SENSOR CABLE AND FOUND CABLE TO BE DEFECTIVE. NEW CABLE REPLACED AND ISSUE RECTIFIED. | SENSOR CABLE | NO | 2023-09-01 | September 2023 | 1:15 | 2023-09-01 01:00 | 2023-09-01 01:15 | 0:15 | COMPLETED | CRITICAL
317 | 2023-09-01 | Sep 2023 | 6:35 | 6:40 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | HAFIZ | ICU | NEED TO REPLACE A CATRIDGE. | REMOVED THE OLD CATRIDGE AND INSERTED A NEW CATRIDGE (S.NO:301613052). WARMING UP. NOW WORKING GOOD. | CARTRIDGE | NO | 2023-09-01 | September 2023 | 7:10 | 2023-09-01 06:35 | 2023-09-01 07:10 | 0:35 | COMPLETED | SEMI-CRITICAL
318 | 2023-09-01 | Sep 2023 | 7:55 | 8:00 | 0.0034722 | HAEMATOLOGY ANALYSER | TRANSASIA BIOMEDICALS | XN 1000 | HAFIZ | LABORATORY | SYSTEM KEYBOARD AND SCREEN NOT WORKING. | ENABLED THE 'VALIDATE' OPTION. SETTINGS DONE. NOW, WORKING GOOD. | NO | NO | 2023-09-01 | September 2023 | 8:15 | 2023-09-01 07:55 | 2023-09-01 08:15 | 0:20 | COMPLETED | SEMI-CRITICAL
319 | 2023-09-01 | Sep 2023 | 8:25 | 8:30 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | ICU | NOT WORKING | CHECKED THE BED AND FOUND THAT BATTERY FULLY DRAINED. KEPT IN CHARGING AND NOW WORKING GOOD. | NO | NO | 2023-09-01 | September 2023 | 8:45 | 2023-09-01 08:25 | 2023-09-01 08:45 | 0:20 | COMPLETED | NON-CRITICAL
320 | 2023-09-01 | Sep 2023 | 12:30 | 12:35 | 0.0034722 | OPTICAL COHERENCE TOMOGRAPHY, 3 DIMENSIONOL | TOPCON | 3D OCT-1 | SANTHOSH | OPD | OPTHALMOLOGY REPORT DOESN’T ABLE TO PRINT | CHECKED ALL CONNECTIONS AND RESTART | NO | NO | 2023-09-01 | September 2023 | 13:00 | 2023-09-01 12:30 | 2023-09-01 13:00 | 0:30 | COMPLETED | SEMI-CRITICAL
321 | 2023-09-02 | Sep 2023 | 21:10 | 21:15 | 0.0034722 | PACEMAKER,DUAL CHAMBER | GETINGE | PACE T20 | SOORAJ | ICU | NOT SWITCHING ON | BATTERY FOUND DEFECTIVE. CHANGED BATTERY AND ISSUE RECTIFIED. | BATTERY | NO | 2023-09-02 | September 2023 | 21:30 | 2023-09-02 21:10 | 2023-09-02 21:30 | 0:20 | COMPLETED | CRITICAL
322 | 2023-09-02 | Sep 2023 | 21:45 | 21:50 | 0.0034722 | BIPAP MACHINE | PHILIPS | A40 | SOORAJ | WARD | AUTOMATICALLY SWITCHED OFF | FOUND BATTERY DEPLETED. KEPT ON CHARGE AND RAN ON BATTERY FOR 1 HR, FOUND MACHINE WORKING FINE. | NO | NO | 2023-09-02 | September 2023 | 22:45 | 2023-09-02 21:45 | 2023-09-02 22:45 | 1:00 | COMPLETED | SEMI-CRITICAL
323 | 2023-09-02 | Sep 2023 | 16:25 | 16:35 | 0.0069444 | TREATMENT UNIT | SKANRAY DENTAL | SKANLECTICA OPAL | HAFIZ | OPD | HANDPIECE NOT WORKING | REMOVED THE HANDPIECE, REFIXED AND CHECKED. NOW, WORKING FINE. | NO | NO | 2023-09-02 | September 2023 | 16:50 | 2023-09-02 16:25 | 2023-09-02 16:50 | 0:25 | COMPLETED | SEMI-CRITICAL
324 | 2023-09-02 | Sep 2023 | 17:30 | 17:35 | 0.0034722 | DEFIBRILLATOR MONITOR WITH AED | PHILIPS | EFFICIA DFM 100 | HAFIZ | OPD | ERROR CROSS SYMBOL (NOT READY TO USE) | RERUNNED THE OPERATIONAL CHECK AND PASSED THE TEST. NOW, WORKING GOOD. | NO | NO | 2023-09-02 | September 2023 | 17:50 | 2023-09-02 17:30 | 2023-09-02 17:50 | 0:20 | COMPLETED | CRITICAL
325 | 2023-09-02 | Sep 2023 | 8:50 | 9:00 | 0.0069444 | CCO MONITOR | EDWARDS | HEMOSPHERE HEM1 | RAMKIE | ICU | NOT WORKING | CHECKED AND RECONNECTED, NOT ISSUES FOUND | NO | NO | 2023-09-02 | September 2023 | 9:10 | 2023-09-02 08:50 | 2023-09-02 09:10 | 0:20 | COMPLETED | CRITICAL
326 | 2023-09-02 | Sep 2023 | 14:55 | 15:00 | 0.0034722 | PACEMAKER,DUAL CHAMBER | GETINGE | PACE T20 | RAMKIE | ICU | NOT WORKING | PROBLEM WITH CABLE | CABLE CONNECTOR | NO | 2023-09-02 | September 2023 | 15:15 | 2023-09-02 14:55 | 2023-09-02 15:15 | 0:20 | COMPLETED | CRITICAL
327 | 2023-09-03 | Sep 2023 | 3:35 | 3:40 | 0.0034722 | VENTILATOR, ADULT | MAQUET | SERVO-I | SOORAJ | ICU | INTERNOL LEAKAGE TEST FAILED | FOUND CASSETTE TO BE AN ISSUE. SWAPPED CASSETTES AND RAN PRE USE CHHECK, MACHINE WORKING FINE. | CASSETTE | NO | 2023-09-03 | September 2023 | 4:00 | 2023-09-03 03:35 | 2023-09-03 04:00 | 0:25 | COMPLETED | CRITICAL
328 | 2023-09-03 | Sep 2023 | 9:00 | 9:05 | 0.0034722 | VESSEL SEALING GENERATOR | COVIDIEN | VALLEY LAB LS10 | SANTHOSH | OT | E227 | CHECKED WITH ANOTHER EARTH PAD | EARTH PAD | NO | 2023-09-03 | September 2023 | 9:10 | 2023-09-03 09:00 | 2023-09-03 09:10 | 0:10 | COMPLETED | SEMI-CRITICAL
329 | 2023-09-04 | Sep 2023 | 11:35 | 11:40 | 0.0034722 | BABY WEIGHING SCALE | SECA | 354 | HAFIZ | OPD | NOT WORKING | REMOVED THE BATTERY, REFIXED AND CHECKED WITH THE BABY. NOW WORKING GOOD. | NO | NO | 2023-09-04 | September 2023 | 12:00 | 2023-09-04 11:35 | 2023-09-04 12:00 | 0:25 | COMPLETED | NON-CRITICAL
330 | 2023-09-04 | Sep 2023 | 12:00 | 12:05 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | OPD | WRONG DATE AND TIME IN PRINTER PAPER | NEEDFUL DONE. NOW WORKING GOOD. | NO | NO | 2023-09-04 | September 2023 | 12:20 | 2023-09-04 12:00 | 2023-09-04 12:20 | 0:20 | COMPLETED | CRITICAL
331 | 2023-09-05 | Sep 2023 | 15:45 | 15:50 | 0.0034722 | VIDEO ELECTRO ENCEPHALOGRAPHY | NOTUS NEUROLOGY | VEEG | SOORAJ | OPD | SYSTEM TOO SLOW | SOFTWARE RELOADED AND CHEKCED WORKING FINE. UNDER OBSERVATION. | NO | NO | 2023-09-05 | September 2023 | 16:30 | 2023-09-05 15:45 | 2023-09-05 16:30 | 0:45 | COMPLETED | SEMI-CRITICAL
332 | 2023-09-05 | Sep 2023 | 10:00 | 10:15 | 0.010417 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | ICU | WORKS TOO SLOW | CHANGED THE SOM BOARD AND RECTIFIED THE ISSUE | SOM BOARD | NO | 2023-09-05 | September 2023 | 10:45 | 2023-09-05 10:00 | 2023-09-05 10:45 | 0:45 | COMPLETED | CRITICAL
333 | 2023-09-05 | Sep 2023 | 22:10 | 22:20 | 0.0069444 | VENTILATOR, ADULT | MAQUET | SERVO-I | RAMKIE | EMERGENCY | TEST NOT PASSED | FOUND THE PROBLEM WITH EXP CASSETTE, SO REPLACED IT | NO | NO | 2023-09-05 | September 2023 | 22:25 | 2023-09-05 22:10 | 2023-09-05 22:25 | 0:15 | COMPLETED | CRITICAL
334 | 2023-09-05 | Sep 2023 | 22:55 | 23:05 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | RAMKIE | ICU | LEAKAGE | CHECKED AND FOUND THE LEAKAGE FROM MACHINE, FOUND THE PROBLEM WITH CATRIDGE | NO | NO | 2023-09-05 | September 2023 | 23:10 | 2023-09-05 22:55 | 2023-09-05 23:10 | 0:15 | COMPLETED | SEMI-CRITICAL
335 | 2023-09-06 | Sep 2023 | 9:10 | 9:17 | 0.0048611 | PATIENT MONITOR | PHILIPS | MX 550 | SANTHOSH | OT | NIBP MALFUCTION | REPLACED NBP MODULE BY COMPANY PERSON | NIBP MODULE | NO | 2023-09-06 | September 2023 | 9:25 | 2023-09-06 09:10 | 2023-09-06 09:25 | 0:15 | COMPLETED | SEMI-CRITICAL
336 | 2023-09-06 | Sep 2023 | 10:10 | 10:20 | 0.0069444 | CATHLAB - MONOPLANE | PHILIPS | AZURION 7M12 | RAMKIE | CATHLAB | FOOT SWITCH FLUROSCOPE NOT WORKING | FOUND BROKEN AND MADE TAG AND ADJUSTED AND WORKING WITH CONSTRAINS | NO | NO | 2023-09-06 | September 2023 | 10:40 | 2023-09-06 10:10 | 2023-09-06 10:40 | 0:30 | COMPLETED | CRITICAL
337 | 2023-09-07 | Sep 2023 | 15:01 | 15:05 | 0.0027778 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SOORAJ | ICU | SYSTEM TOO SLOW | FOUND MAIN BOARD ISSUE. REPLACED MAIN BOARD AND ISSUE RECTIFIED. | MAIN BOARD | NO | 2023-09-07 | September 2023 | 15:20 | 2023-09-07 15:01 | 2023-09-07 15:20 | 0:19 | COMPLETED | CRITICAL
338 | 2023-09-07 | Sep 2023 | 11:00 | 11:15 | 0.010417 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | RADIOLOGY | NEED TO REPLACE BATTERY | REPLACED THE OLD BATTERY WITH A NEW ONE AND IT'S WORKING GOOD. | BATTERY | NO | 2023-09-07 | September 2023 | 11:30 | 2023-09-07 11:00 | 2023-09-07 11:30 | 0:30 | COMPLETED | CRITICAL
339 | 2023-09-07 | Sep 2023 | 16:00 | 16:15 | 0.010417 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | JERLIN | OPD | V1, V3 AND V5 SHOWS ARTIFACTS. | REPLACED THE DEFECTIVE LEAD CABLE WITH A NEW ONE AND REECTIFIED THE ISSUE. | 10-LEADS ECG CABLE | NO | 2023-09-07 | September 2023 | 16:30 | 2023-09-07 16:00 | 2023-09-07 16:30 | 0:30 | COMPLETED | NON-CRITICAL
340 | 2023-09-07 | Sep 2023 | 17:00 | 17:15 | 0.010417 | VENTILATOR, UNIVERSAL | MINDRAY MEDICAL INDIA | SV 300 | JERLIN | ICU | DEVICE 13 ERROR | REPLACED PROPORTION VALVE WITH A NEW ON. | PROPORTION VALVE | NO | 2023-09-07 | September 2023 | 17:30 | 2023-09-07 17:00 | 2023-09-07 17:30 | 0:30 | COMPLETED | CRITICAL
341 | 2023-09-07 | Sep 2023 | 17:00 | 17:10 | 0.0069444 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | JERLIN | EMERGENCY | SPO2 PROBE CONNECTOR BEND | FIXED THE CONNECTOR PIN AND RECTIFIED THE ISSUE | NO | NO | 2023-09-07 | September 2023 | 17:20 | 2023-09-07 17:00 | 2023-09-07 17:20 | 0:20 | COMPLETED | SEMI-CRITICAL
342 | 2023-09-07 | Sep 2023 | 12:30 | 12:37 | 0.0048611 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SANTHOSH | WARD | CONTROL BOARD GOT DEFECT | NEED TO REPLACE THE BOARD | CONTROL PANEL | NO | 2023-09-07 | September 2023 | 12:45 | 2023-09-07 12:30 | 2023-09-07 12:45 | 0:15 | COMPLETED | NON-CRITICAL
343 | 2023-09-07 | Sep 2023 | 21:10 | 21:20 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | RAMKIE | ICU | CATRIDGE REPLACEMENT | REPLACED CATRIDGE | CARTRIDGE | NO | 2023-09-07 | September 2023 | 21:20 | 2023-09-07 21:10 | 2023-09-07 21:20 | 0:10 | COMPLETED | SEMI-CRITICAL
344 | 2023-09-08 | Sep 2023 | 18:00 | 18:05 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | SOORAJ | RADIOLOGY | NOT ABLE TO ADJUST GAIN. | TCG BOARD FOUND ISSUE. BOARD REPLACED AND ISSUE RECTIFIED. | TCG BOARD | NO | 2023-09-08 | September 2023 | 20:00 | 2023-09-08 18:00 | 2023-09-08 20:00 | 2:00 | COMPLETED | SEMI-CRITICAL
345 | 2023-09-08 | Sep 2023 | 18:30 | 18:35 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | SOORAJ | OT | PAEDIATRIC TEE PROBE NOT WORKING | FOUND SENSOR DEFECTIVE. NEW PROBE REPLACED AND GIVEN UNDER WARANTY | TEE PROBE | NO | 2023-09-08 | September 2023 | 20:00 | 2023-09-08 18:30 | 2023-09-08 20:00 | 1:30 | COMPLETED | SEMI-CRITICAL
346 | 2023-09-08 | Sep 2023 | 14:00 | 14:10 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | OT | BRAKE LEVER BROKEN | FIXED THE BRAKE LEVER AND ISSUE RECTIFIED | NO | NO | 2023-09-08 | September 2023 | 14:15 | 2023-09-08 14:00 | 2023-09-08 14:15 | 0:15 | COMPLETED | NON-CRITICAL
347 | 2023-09-08 | Sep 2023 | 14:30 | 14:40 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | WARD | BRAKE LEVER BROKEN | FIXED THE BRAKE LEVER AND ISSUE RECTIFIED | NO | NO | 2023-09-08 | September 2023 | 14:50 | 2023-09-08 14:30 | 2023-09-08 14:50 | 0:20 | COMPLETED | NON-CRITICAL
348 | 2023-09-08 | Sep 2023 | 15:00 | 15:10 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 9000 | JERLIN | ICU | ERROR MESSAGE IS DISPLAYED | POWER SUPPLY BOARD IS CHANGED AND ISSUE RECTFIED | POWER SUPPLY BOARD | NO | 2023-09-08 | September 2023 | 15:30 | 2023-09-08 15:00 | 2023-09-08 15:30 | 0:30 | COMPLETED | NON-CRITICAL
349 | 2023-09-08 | Sep 2023 | 9:10 | 9:15 | 0.0034722 | IABP MACHINE | MAQUET | CS300 | SANTHOSH | OT | HELIUM CYLINDER GOT LEAK | WASHER GOT LOOSENED ,REPLACED NEW WASHER | NO | NO | 2023-09-08 | September 2023 | 9:20 | 2023-09-08 09:10 | 2023-09-08 09:20 | 0:10 | COMPLETED | CRITICAL
350 | 2023-09-08 | Sep 2023 | 14:20 | 14:25 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 9000 | SANTHOSH | ICU | NOT WORKING | CHECKED AND POWERCORD RESET | NO | NO | 2023-09-08 | September 2023 | 14:35 | 2023-09-08 14:20 | 2023-09-08 14:35 | 0:15 | COMPLETED | NON-CRITICAL
351 | 2023-09-08 | Sep 2023 | 16:15 | 16:20 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | SANTHOSH | OT | O2 VALUE DOESN’T SHOWS PROPERLY | REPLACED NEW O2 SENSOR CABLE | O2 SENSOR CABLE | NO | 2023-09-08 | September 2023 | 16:30 | 2023-09-08 16:15 | 2023-09-08 16:30 | 0:15 | COMPLETED | CRITICAL
352 | 2023-09-09 | Sep 2023 | 10:10 | 10:15 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | JERLIN | RADIOLOGY | WORKLIST NOT LOADING | FOUND LOTS OF PENDING DATA IN SPOOLER . CLEARED AND RECTIFIED THE ISSUE | NO | NO | 2023-09-09 | September 2023 | 10:30 | 2023-09-09 10:10 | 2023-09-09 10:30 | 0:20 | COMPLETED | SEMI-CRITICAL
353 | 2023-09-09 | Sep 2023 | 13:45 | 13:50 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SANTHOSH | ENDOSCOPY | O2 CELL ISSUE | 21% AND 100% CALIBRATION DONE | NO | NO | 2023-09-09 | September 2023 | 14:00 | 2023-09-09 13:45 | 2023-09-09 14:00 | 0:15 | COMPLETED | CRITICAL
354 | 2023-09-09 | Sep 2023 | 16:10 | 16:15 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SANTHOSH | OT | CALIBRATION DUE | CALIBRATION PASSED | NO | NO | 2023-09-09 | September 2023 | 16:25 | 2023-09-09 16:10 | 2023-09-09 16:25 | 0:15 | COMPLETED | CRITICAL
355 | 2023-09-10 | Sep 2023 | 10:28 | 10:30 | 0.0013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | ICU | FUNCTIONS NOT WORKING | FOUND COT LOCKED. RESET DONE AND ISSUE RECTIFIED. | NO | NO | 2023-09-10 | September 2023 | 10:32 | 2023-09-10 10:28 | 2023-09-10 10:32 | 0:04 | COMPLETED | NON-CRITICAL
356 | 2023-09-10 | Sep 2023 | 12:32 | 12:35 | 0.0020833 | RADIOGRAPHIC UNIT | CARESTREAM | DRF ASCEND ANOLOG HF | SOORAJ | RADIOLOGY | NOT ABLE TO PUSH TO PACS | ETHERNET SWITCH WAS OFF. POWERED ON AND ISSUE RECTIFIED. | NO | NO | 2023-09-10 | September 2023 | 12:40 | 2023-09-10 12:32 | 2023-09-10 12:40 | 0:08 | COMPLETED | SEMI-CRITICAL
357 | 2023-09-10 | Sep 2023 | 15:30 | 15:35 | 0.0034722 | LIGHT SOURCE UNIT,XENON | KARL STORZ | 20133720 | SOORAJ | OT | GENERAL CHECK AFTER MOVING FROM CI | CONNECTED AND CHECKED THE SYSTEM, OBSERVED WORKING FINE. | NO | NO | 2023-09-10 | September 2023 | 16:00 | 2023-09-10 15:30 | 2023-09-10 16:00 | 0:30 | COMPLETED | SEMI-CRITICAL
358 | 2023-09-10 | Sep 2023 | 15:30 | 15:35 | 0.0034722 | MONITOR HIGH DEFINITION | KARL STORZ | TM220 | SOORAJ | OT | GENERAL CHECK AFTER MOVING FROM CI | CONNECTED AND CHECKED THE SYSTEM, OBSERVED WORKING FINE. | NO | NO | 2023-09-10 | September 2023 | 16:00 | 2023-09-10 15:30 | 2023-09-10 16:00 | 0:30 | COMPLETED | SEMI-CRITICAL
359 | 2023-09-10 | Sep 2023 | 15:30 | 15:35 | 0.0034722 | VIDEO PROCESSING SYSTEM,ENDOSCOPES | KARL STORZ | TH102 | SOORAJ | OT | GENERAL CHECK AFTER MOVING FROM CI | CONNECTED AND CHECKED THE SYSTEM, OBSERVED WORKING FINE. | NO | NO | 2023-09-10 | September 2023 | 16:00 | 2023-09-10 15:30 | 2023-09-10 16:00 | 0:30 | COMPLETED | SEMI-CRITICAL
360 | 2023-09-10 | Sep 2023 | 15:30 | 15:35 | 0.0034722 | ENDO INSUFFLATOR,50L | KARL STORZ | UI500 | SOORAJ | OT | GENERAL CHECK AFTER MOVING FROM CI | CONNECTED AND CHECKED THE SYSTEM, OBSERVED WORKING FINE. | NO | NO | 2023-09-10 | September 2023 | 16:00 | 2023-09-10 15:30 | 2023-09-10 16:00 | 0:30 | COMPLETED | SEMI-CRITICAL
361 | 2023-09-10 | Sep 2023 | 10:10 | 10:30 | 0.013889 | INCUBATOR,TRANSPORT | PHOENIX | TINC 101 | RAMKIE | ICU | NOT SENSING PROBE | FOUND THE PROBLEM WITH BOARD | NO | MAIN BOARD | 2023-09-10 | September 2023 | 10:40 | 2023-09-10 10:10 | 2023-09-10 10:40 | 0:30 | PENDING | SEMI-CRITICAL
362 | 2023-09-11 | Sep 2023 | 11:30 | 11:32 | 0.0013889 | C-MAC MONITOR | KARL STORZ | 84032ZX | SOORAJ | OT | AUTOMATIC RESTART | CHECKED AND FACTORY RESET DONE. ISSUE RECTFIED. | NO | NO | 2023-09-11 | September 2023 | 11:40 | 2023-09-11 11:30 | 2023-09-11 11:40 | 0:10 | COMPLETED | SEMI-CRITICAL
363 | 2023-09-11 | Sep 2023 | 15:15 | 15:20 | 0.0034722 | GO FLO PUMP, ARTHROSCOPY | SMITH & NEPHEW | 72204968 | SOORAJ | OT | ERROR WHILE SWITHCING ON | CHECKED AND FOUND NO ERROR WHILE USING. | NO | NO | 2023-09-11 | September 2023 | 15:30 | 2023-09-11 15:15 | 2023-09-11 15:30 | 0:15 | COMPLETED | SEMI-CRITICAL
364 | 2023-09-11 | Sep 2023 | 16:40 | 16:45 | 0.0034722 | OT TABLE | MAQUET | MEERA | SOORAJ | OT | REMOTE FUNCTIONS NOT WORKING | CHECKED ALL FUNCTIONS AND FOUND NO ISSUES. | NO | NO | 2023-09-11 | September 2023 | 16:50 | 2023-09-11 16:40 | 2023-09-11 16:50 | 0:10 | COMPLETED | NON-CRITICAL
365 | 2023-09-11 | Sep 2023 | 11:30 | 11:35 | 0.0034722 | TREATMENT UNIT | SKANRAY DENTAL | SKANLECTICA OPAL | HAFIZ | OPD | FOOT PEDAL NOT WORKING | CHECKED THE PEDAL AND FOUND THAT IT IS BROKEN. TEMPORARILY FIXED AND NOW WORKING GOOD. | NO | NO | 2023-09-11 | September 2023 | 11:45 | 2023-09-11 11:30 | 2023-09-11 11:45 | 0:15 | COMPLETED | SEMI-CRITICAL
366 | 2023-09-11 | Sep 2023 | 12:35 | 12:40 | 0.0034722 | REVERSE OSMOSIS SYSTEM | ENGINEERING ALLIANCE | 500 LPH | HAFIZ | ICU | RO NOT SWITCHING ON | CHECKED THE RO PLANT AND FOUND THAT MAIN POWER SUPPLY NOT TURNED ON. NOW TURNED ON AND WORKING GOOD. | NO | NO | 2023-09-11 | September 2023 | 12:45 | 2023-09-11 12:35 | 2023-09-11 12:45 | 0:10 | COMPLETED | SEMI-CRITICAL
367 | 2023-09-11 | Sep 2023 | 16:30 | 16:40 | 0.0069444 | ACT MACHINE | HELENO LABORATORIES | MINI II | JERLIN | ICU | S1 ERROR WHILE RUNNING SAMPLE | RERAN TEST WITH ANOTHER SAMPLE CONTAINER AND FOUND IT'S WORKING GOOD. | NO | NO | 2023-09-11 | September 2023 | 16:55 | 2023-09-11 16:30 | 2023-09-11 16:55 | 0:25 | COMPLETED | SEMI-CRITICAL
368 | 2023-09-11 | Sep 2023 | 11:00 | 11:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | OPD | PRINTER NOT WORKING | CHECKED THE PRINTER AND IT'S WORKING FINE. | NO | NO | 2023-09-11 | September 2023 | 11:20 | 2023-09-11 11:00 | 2023-09-11 11:20 | 0:20 | COMPLETED | CRITICAL
369 | 2023-09-11 | Sep 2023 | 11:30 | 11:40 | 0.0069444 | BMI MACHINE | AKAS MEDICAL | QVS-100 | JERLIN | OPD | LEAK IN PEDIATRIC BP CUFF | REPLACED WITH A NEW ONE. CHECKED AND IT'S IN GOOD CONDITION. | LARGE ADULT BP CUFF | NO | 2023-09-11 | September 2023 | 11:55 | 2023-09-11 11:30 | 2023-09-11 11:55 | 0:25 | COMPLETED | NON-CRITICAL
370 | 2023-09-11 | Sep 2023 | 12:00 | 12:10 | 0.0069444 | RADIOGRAPHIC UNIT | CARESTREAM | DRF ASCEND ANOLOG HF | JERLIN | RADIOLOGY | WORKLIST NOT LOADING | TOO MUCH OF DATA IN SPOOLER . AFTER CLEARING IT WORKS GOOD. | NO | NO | 2023-09-11 | September 2023 | 12:25 | 2023-09-11 12:00 | 2023-09-11 12:25 | 0:25 | COMPLETED | SEMI-CRITICAL
371 | 2023-09-11 | Sep 2023 | 14:00 | 14:15 | 0.010417 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | ICU | REMOTE LOCKED | LOCK RELEASED AND ISSUE RECTIFIED | NO | NO | 2023-09-11 | September 2023 | 14:20 | 2023-09-11 14:00 | 2023-09-11 14:20 | 0:20 | COMPLETED | NON-CRITICAL
372 | 2023-09-11 | Sep 2023 | 7:30 | 7:35 | 0.0034722 | TRANSCRANIAL DOPPLER | VIASONIC | DOLPHIN 4D | JERLIN | ICU | WORKS TOO SLOW | FOUND THAT LOTS OF DATA ACCUMULATED IN STORAGE . CLEARED AND RECTIFIED THE ISSUE | NO | NO | 2023-09-11 | September 2023 | 7:50 | 2023-09-11 07:30 | 2023-09-11 07:50 | 0:20 | COMPLETED | SEMI-CRITICAL
373 | 2023-09-12 | Sep 2023 | 7:45 | 7:50 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | SODA LIME WASHER CLEANED AND SODA LIME QUANTITY REDUCED. RAN CHECKOUT AND PASSED. | NO | NO | 2023-09-12 | September 2023 | 8:00 | 2023-09-12 07:45 | 2023-09-12 08:00 | 0:15 | COMPLETED | CRITICAL
374 | 2023-09-12 | Sep 2023 | 9:00 | 9:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | SODA LIME WASHER CLEANED AND SODA LIME QUANTITY REDUCED. RAN CHECKOUT AND PASSED. | NO | NO | 2023-09-12 | September 2023 | 9:20 | 2023-09-12 09:00 | 2023-09-12 09:20 | 0:20 | COMPLETED | CRITICAL
375 | 2023-09-12 | Sep 2023 | 10:50 | 10:55 | 0.0034722 | IABP MACHINE | MAQUET | CS300 | SOORAJ | ICU | LEAK IN HELIUM CYLINDER | FOUND WASHER DEFECTIVE. REPLACED NEW WASHER | WASHER | NO | 2023-09-12 | September 2023 | 11:10 | 2023-09-12 10:50 | 2023-09-12 11:10 | 0:20 | COMPLETED | CRITICAL
376 | 2023-09-12 | Sep 2023 | 11:30 | 11:35 | 0.0034722 | IABP MACHINE | MAQUET | CS300 | SOORAJ | OT | HELIUM CYLINDER EMPTY | NEW HELIUM CYCLINDER ISSUED | HELIUM CYLINDER | NO | 2023-09-12 | September 2023 | 11:43 | 2023-09-12 11:30 | 2023-09-12 11:43 | 0:13 | COMPLETED | CRITICAL
377 | 2023-09-12 | Sep 2023 | 14:10 | 14:11 | 0.00069444 | DEFIBRILLATOR MONITOR | WIPRO GE HEALTH CARE | CARESTATION 650 | SOORAJ | OT | DEFIB NOT BOOTIG UP ON BATTERY | FOUND BATTERY LOOSENED. BATTERY TAKEN OUT AND REINSRTED PROPERLY, ISSUE RECTIFIED. | NO | NO | 2023-09-12 | September 2023 | 14:20 | 2023-09-12 14:10 | 2023-09-12 14:20 | 0:10 | COMPLETED | CRITICAL
378 | 2023-09-12 | Sep 2023 | 16:30 | 16:35 | 0.0034722 | LIGHT SOURCE UNIT,XENON | KARL STORZ | 20133720 | SOORAJ | OT | HYSTEROSCOPY IMAGE NOT CLEAR | CLEANED LIGHT SOURCE SCOPE AMD LENS. CONNECTED AND CHECKED, FOUND IMAGE TO BE CLEAR AND FINE. | NO | NO | 2023-09-12 | September 2023 | 16:50 | 2023-09-12 16:30 | 2023-09-12 16:50 | 0:20 | COMPLETED | SEMI-CRITICAL
379 | 2023-09-12 | Sep 2023 | 16:30 | 16:35 | 0.0034722 | MONITOR HIGH DEFINITION | KARL STORZ | TM220 | SOORAJ | OT | HYSTEROSCOPY IMAGE NOT CLEAR | CLEANED LIGHT SOURCE SCOPE AMD LENS. CONNECTED AND CHECKED, FOUND IMAGE TO BE CLEAR AND FINE. | NO | NO | 2023-09-12 | September 2023 | 16:50 | 2023-09-12 16:30 | 2023-09-12 16:50 | 0:20 | COMPLETED | SEMI-CRITICAL
380 | 2023-09-12 | Sep 2023 | 16:30 | 16:35 | 0.0034722 | VIDEO PROCESSING SYSTEM,ENDOSCOPES | KARL STORZ | TH102 | SOORAJ | OT | HYSTEROSCOPY IMAGE NOT CLEAR | CLEANED LIGHT SOURCE SCOPE AMD LENS. CONNECTED AND CHECKED, FOUND IMAGE TO BE CLEAR AND FINE. | NO | NO | 2023-09-12 | September 2023 | 16:50 | 2023-09-12 16:30 | 2023-09-12 16:50 | 0:20 | COMPLETED | SEMI-CRITICAL
381 | 2023-09-12 | Sep 2023 | 17:30 | 17:35 | 0.0034722 | NITRIC OXIDE DELIVERY SYSTEM | WIPRO GE HEALTH CARE | SLE | SOORAJ | ICU | REPLACE SENSOR MESSAGE | NEW NO & NO2 SENSORS REPLACED. | YENITIC OXIDE & NIROUS DI OXIDE SENSOR | NO | 2023-09-12 | September 2023 | 18:00 | 2023-09-12 17:30 | 2023-09-12 18:00 | 0:30 | COMPLETED | CRITICAL
382 | 2023-09-12 | Sep 2023 | 10:00 | 10:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | WARD | THERAPY KNOB TEST FAILED | CONNECTED THE PADDLES TO THE MACHINE AND RAN TEST . PASSED | NO | NO | 2023-09-12 | September 2023 | 10:20 | 2023-09-12 10:00 | 2023-09-12 10:20 | 0:20 | COMPLETED | CRITICAL
383 | 2023-09-12 | Sep 2023 | 14:00 | 14:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | JERLIN | CATHLAB | SPO2 PROBE NOT WORKING | FIXED THE PROBE ANFD ISSUE RECTIFIED | NO | NO | 2023-09-12 | September 2023 | 14:10 | 2023-09-12 14:00 | 2023-09-12 14:10 | 0:10 | COMPLETED | SEMI-CRITICAL
384 | 2023-09-12 | Sep 2023 | 22:00 | 22:10 | 0.0069444 | BIPAP MACHINE | PHILIPS | A40 | SANTHOSH | WARD | BATTERY LOW ISSUE | CHECKED WITH ANOTHER PORT ISSUE RECTIFIED | NO | NO | 2023-09-12 | September 2023 | 22:15 | 2023-09-12 22:00 | 2023-09-12 22:15 | 0:15 | COMPLETED | SEMI-CRITICAL
385 | 2023-09-13 | Sep 2023 | 12:00 | 12:05 | 0.0034722 | NITRIC OXIDE DELIVERY SYSTEM | WIPRO GE HEALTH CARE | SLE | SOORAJ | ICU | SENSOR CALIBRATION | NO & NO2 SENSORS CALIBRATED AND HANDEDOVER TO TECH | NO | NO | 2023-09-13 | September 2023 | 12:20 | 2023-09-13 12:00 | 2023-09-13 12:20 | 0:20 | COMPLETED | CRITICAL
386 | 2023-09-13 | Sep 2023 | 0:10 | 0:15 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | SANTHOSH | DIALYSIS | CANT TURN OFF WITH POWER BUTTON | ONLY TURN OFF WITH SETTINGS | NO | NO | 2023-09-13 | September 2023 | 0:20 | 2023-09-13 00:10 | 2023-09-13 00:20 | 0:10 | COMPLETED | SEMI-CRITICAL
387 | 2023-09-13 | Sep 2023 | 22:30 | 22:35 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | SANTHOSH | ICU | CATRIDGE GET EMPTY | NEW CATRIDGE REPLACED | CARTRIDGE | NO | 2023-09-13 | September 2023 | 22:40 | 2023-09-13 22:30 | 2023-09-13 22:40 | 0:10 | COMPLETED | SEMI-CRITICAL
388 | 2023-09-13 | Sep 2023 | 22:45 | 22:50 | 0.0034722 | EXAMINATION COUCH | MIDMARK JANOK | E122012 | SANTHOSH | OPD | COUCH NOT WORKING | PROPERLY FIXED POWER CORD ISSUE RECTIFIED | NO | NO | 2023-09-13 | September 2023 | 23:00 | 2023-09-13 22:45 | 2023-09-13 23:00 | 0:15 | COMPLETED | NON-CRITICAL
389 | 2023-09-14 | Sep 2023 | 7:30 | 7:35 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | SOORAJ | OT | ANGULATION NOT WORKING. | RESTARTED THE SYSTEM AND CHECKED, ISSUE RECTIFIED. | NO | NO | 2023-09-14 | September 2023 | 7:45 | 2023-09-14 07:30 | 2023-09-14 07:45 | 0:15 | COMPLETED | SEMI-CRITICAL
390 | 2023-09-14 | Sep 2023 | 9:01 | 9:05 | 0.0027778 | OT TABLE | MAQUET | MEERA | SOORAJ | OT | OIL LEAK REPORTED | PERFORMED BASIC TABLE FUNCTIONS AND LEAK STOPPED. ISSUE RECTIFIED TEMPORARILY. | NO | NO | 2023-09-14 | September 2023 | 9:20 | 2023-09-14 09:01 | 2023-09-14 09:20 | 0:19 | COMPLETED | NON-CRITICAL
391 | 2023-09-14 | Sep 2023 | 15:20 | 15:25 | 0.0034722 | ULTRASONIC ASPIRATOR | STRYKER | UST-2001 | SOORAJ | OT | SUCTION NOT WORKING | FOUND INTERNAL SUCTION TUBE LOOSELY CONNECTED. RECONNECTED AND ISSUE RECTIFIED. | NO | NO | 2023-09-14 | September 2023 | 17:00 | 2023-09-14 15:20 | 2023-09-14 17:00 | 1:40 | COMPLETED | SEMI-CRITICAL
392 | 2023-09-14 | Sep 2023 | 17:40 | 17:45 | 0.0034722 | TREATMENT UNIT | SKANRAY DENTAL | SKANLECTICA OPAL | HAFIZ | OPD | NOT WORKING | CLEANED THE BOARD WITH IPA. REFIXED THE COMMUNICATION BOARD ALONG WITH THE COMPANY PERSON. NOW WORKING GOOD. | NO | NO | 2023-09-14 | September 2023 | 17:55 | 2023-09-14 17:40 | 2023-09-14 17:55 | 0:15 | COMPLETED | SEMI-CRITICAL
393 | 2023-09-14 | Sep 2023 | 18:10 | 18:16 | 0.0041667 | PRESSURE INJECTOR | LF OPTI VONTAGE | ILLUMENO NEO | HAFIZ | CATHLAB | PLUNGER STUCK DURING THE PROCEDURE | RESTARTED THE MACHINE, CHECKED WITH THE DEMO SYRINGE. NOW WORKING GOOD. | NO | NO | 2023-09-14 | September 2023 | 18:30 | 2023-09-14 18:10 | 2023-09-14 18:30 | 0:20 | COMPLETED | SEMI-CRITICAL
394 | 2023-09-14 | Sep 2023 | 18:40 | 18:42 | 0.0013889 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | HAFIZ | CATHLAB | CALIBRATION DUE | REPLACED THE O2 SENSOR, O2 CELL CALIBRATION DONE. NOW WORKING GOOD. | O2 SENSOR | NO | 2023-09-14 | September 2023 | 19:00 | 2023-09-14 18:40 | 2023-09-14 19:00 | 0:20 | COMPLETED | CRITICAL
395 | 2023-09-14 | Sep 2023 | 23:10 | 23:11 | 0.00069444 | HEATER COOLER | GETINGE | HCU40 | SANTHOSH | OT | PLUG TOP STTRUCKED WITH PENDENT POWER PORT | TAKEN OUT FROM THE PENDENT AND REPLACED WITH NEW PLUG TOP | PLUG TOP | NO | 2023-09-14 | September 2023 | 23:20 | 2023-09-14 23:10 | 2023-09-14 23:20 | 0:10 | COMPLETED | CRITICAL
396 | 2023-09-15 | Sep 2023 | 8:30 | 8:35 | 0.0034722 | HEATER COOLER | GETINGE | HCU40 | SOORAJ | OT | NOT SWITCHING ON | FOUND CONNECTORS INSIDE PLUG TOP LOOSELY CONNECTED. REFIXED AND ISSUE RECTIFIED. | NO | NO | 2023-09-15 | September 2023 | 9:00 | 2023-09-15 08:30 | 2023-09-15 09:00 | 0:30 | COMPLETED | SEMI-CRITICAL
397 | 2023-09-15 | Sep 2023 | 9:02 | 9:05 | 0.0020833 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | SOORAJ | OT | O2 SUPPLY LOW ALARM | FOUND HOSE NOT CONNECTED PROPERLY. REFIXED AND ISSUE RECTIFIED. | NO | NO | 2023-09-15 | September 2023 | 9:10 | 2023-09-15 09:02 | 2023-09-15 09:10 | 0:08 | COMPLETED | CRITICAL
398 | 2023-09-15 | Sep 2023 | 12:00 | 12:02 | 0.0013889 | PATIENT MONITOR | PHILIPS | CM12 | SOORAJ | DIALYSIS | AUTOMATIC RESTART | COMMUNICATION BOARD ISSUE. BOARD REPLACED AND ISSUE RECTIFIED. | COMMUNICATION BOARD | NO | 2023-09-16 | September 2023 | 12:30 | 2023-09-15 12:00 | 2023-09-16 12:30 | 24:30 | COMPLETED | SEMI-CRITICAL
399 | 2023-09-15 | Sep 2023 | 11:50 | 11:55 | 0.0034722 | VIDEO PROCESSING SYSTEM,ENDOSCOPES | OLYMPUS | CV-190 | HAFIZ | OT | VIDEO NOT DISPLAYING DURING THE PROCEDURE | REFIXED THE CABLE AND NOW WORKING FINE. | NO | NO | 2023-09-15 | September 2023 | 12:00 | 2023-09-15 11:50 | 2023-09-15 12:00 | 0:10 | COMPLETED | SEMI-CRITICAL
400 | 2023-09-15 | Sep 2023 | 12:10 | 12:15 | 0.0034722 | BIPAP MACHINE | PHILIPS | A40 | HAFIZ | ICU | FILTER IS NOT PRESENT IN THE MACHINE | REFIXED THE FILTER IN THE BIPAP MACHINE. NOW WORKING GOOD. | NO | NO | 2023-09-15 | September 2023 | 12:25 | 2023-09-15 12:10 | 2023-09-15 12:25 | 0:15 | COMPLETED | SEMI-CRITICAL
401 | 2023-09-16 | Sep 2023 | 14:30 | 14:35 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | 94097862 | SOORAJ | ICU | CARTRIDGE EMPTY | NEW CARTRIDGE REPLACED AND CVP DONE. | CARTRIDGE | NO | 2023-09-16 | September 2023 | 15:30 | 2023-09-16 14:30 | 2023-09-16 15:30 | 1:00 | COMPLETED | SEMI-CRITICAL
402 | 2023-09-16 | Sep 2023 | 14:00 | 14:30 | 0.020833 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | JERLIN | CATHLAB | DEFECTIVE HOSE | REPLACED IT WITH A WORKING ONE | NIBP HOSE | NO | 2023-09-16 | September 2023 | 14:40 | 2023-09-16 14:00 | 2023-09-16 14:40 | 0:40 | COMPLETED | SEMI-CRITICAL
403 | 2023-09-16 | Sep 2023 | 8:00 | 8:10 | 0.0069444 | DIGITAL VISION CHART | DELL | D1918H | JERLIN | OPD | REMOTE NOT WORKING | REFIXED THE BATTERY AND ISSUE RECTIFIED | NO | NO | 2023-09-16 | September 2023 | 8:20 | 2023-09-16 08:00 | 2023-09-16 08:20 | 0:20 | COMPLETED | SEMI-CRITICAL
404 | 2023-09-16 | Sep 2023 | 20:10 | 20:15 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | CATRIDGE REAGENT GOT LEAKED | REPLACED ANOTHER CATRIDGE | CARTRIDGE | NO | 2023-09-16 | September 2023 | 20:30 | 2023-09-16 20:10 | 2023-09-16 20:30 | 0:20 | COMPLETED | SEMI-CRITICAL
405 | 2023-09-17 | Sep 2023 | 9:00 | 9:05 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | EMERGENCY | OPERATIONAL CHECK FAILED | FOUND THAT ECG CABLE MALFUNCTION. REMOVED THE ECG CABLE, CLEANED AND REFIXED. NOW, OPERATIONAL CHECK PASSED AND WORKING GOOD. | NO | NO | 2023-09-17 | September 2023 | 9:20 | 2023-09-17 09:00 | 2023-09-17 09:20 | 0:20 | COMPLETED | CRITICAL
406 | 2023-09-17 | Sep 2023 | 10:40 | 10:45 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | ICU | OPERATIONAL CHECK FAILED | FOUND THAT ECG CABLE MALFUNCTION. REMOVED THE ECG CABLE, CLEANED AND REFIXED. NOW, OPERATIONAL CHECK PASSED AND WORKING GOOD. | NO | NO | 2023-09-17 | September 2023 | 10:55 | 2023-09-17 10:40 | 2023-09-17 10:55 | 0:15 | COMPLETED | CRITICAL
407 | 2023-09-17 | Sep 2023 | 11:00 | 11:02 | 0.0013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | ICU | NOT WORKING | CHECKED THE CONNECTIONS AND FOUND FINE. RESET THE MACHINE AND WORKING GOOD. | NO | NO | 2023-09-17 | September 2023 | 11:10 | 2023-09-17 11:00 | 2023-09-17 11:10 | 0:10 | COMPLETED | NON-CRITICAL
408 | 2023-09-17 | Sep 2023 | 11:10 | 11:14 | 0.0027778 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | DIALYSIS | EQUIPMENT DIABLED, SYSTEM FAILURE | FOUND THAT ECG CABLE MALFUNCTION. REMOVED THE ECG CABLE, CLEANED AND REFIXED. NOW, OPERATIONAL CHECK PASSED AND WORKING GOOD. | NO | NO | 2023-09-17 | September 2023 | 11:25 | 2023-09-17 11:10 | 2023-09-17 11:25 | 0:15 | COMPLETED | CRITICAL
409 | 2023-09-17 | Sep 2023 | 11:30 | 11:35 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | EMERGENCY | OPERATIONAL CHECK FAILED | FOUND THAT ECG CABLE MALFUNCTION. REMOVED THE ECG CABLE, CLEANED AND REFIXED. NOW, OPERATIONAL CHECK PASSED AND WORKING GOOD. | NO | NO | 2023-09-17 | September 2023 | 11:45 | 2023-09-17 11:30 | 2023-09-17 11:45 | 0:15 | COMPLETED | CRITICAL
410 | 2023-09-17 | Sep 2023 | 11:30 | 11:35 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | CARTRIDGE NEEDS TO BE REPLACED. | NEW CARTRIDGE LOADED AND CVP TEST DONE. READY TO USE. | CARTRIDGE | NO | 2023-09-17 | September 2023 | 12:30 | 2023-09-17 11:30 | 2023-09-17 12:30 | 1:00 | COMPLETED | SEMI-CRITICAL
411 | 2023-09-17 | Sep 2023 | 15:25 | 15:30 | 0.0034722 | COOLING CENTRIFUGE | HETTICH | ROTO-SILENTA 630RS | HAFIZ | BLOODBANK | NOT WORKING | RESTARTED THE MACHINE, RESET DONE. NOW WORKING GOOD. | NO | NO | 2023-09-17 | September 2023 | 15:40 | 2023-09-17 15:25 | 2023-09-17 15:40 | 0:15 | COMPLETED | SEMI-CRITICAL
412 | 2023-09-17 | Sep 2023 | 3:30 | 3:33 | 0.0020833 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | SANTHOSH | WARD | GETTING UNWANTED NOSE WHILE RUNNING | REPLACED ANOTHER ROTOR UNIT | NO | NO | 2023-09-17 | September 2023 | 3:40 | 2023-09-17 03:30 | 2023-09-17 03:40 | 0:10 | COMPLETED | NON-CRITICAL
413 | 2023-09-18 | Sep 2023 | 10:10 | 10:15 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | SANTHOSH | EMERGENCY | BULP ELECTRODE GOT BROKEN | REPLACED NEW BULB ELECTRODE | BULB ELECTRODE | NO | 2023-09-18 | September 2023 | 10:20 | 2023-09-18 10:10 | 2023-09-18 10:20 | 0:10 | COMPLETED | SEMI-CRITICAL
414 | 2023-09-18 | Sep 2023 | 10:30 | 10:35 | 0.0034722 | BABY WEIGHING SCALE WITH INFANTOMETER | SECA | 334 | SANTHOSH | ICU | NOT WORKING | BATTERY GOT DRIED INFORMED TO REPLACE BATTERY | NO | NO | 2023-09-18 | September 2023 | 10:40 | 2023-09-18 10:30 | 2023-09-18 10:40 | 0:10 | COMPLETED | NON-CRITICAL
415 | 2023-09-18 | Sep 2023 | 12:30 | 12:40 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | SANTHOSH | OPD | TGC BOEARD NOT WORKING | REPLACED NEW TGC BOARD BY COMPANY PERSON | TGC BOARD | NO | 2023-09-18 | September 2023 | 13:00 | 2023-09-18 12:30 | 2023-09-18 13:00 | 0:30 | COMPLETED | SEMI-CRITICAL
416 | 2023-09-18 | Sep 2023 | 17:00 | 17:15 | 0.010417 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S7 | SANTHOSH | RADIOLOGY | ABDOMEN PROBE GOT ARTIFACTS WHILE USINE | REPLACED NEW ABDOMEN PROBE | ABDOMEN PROBE | NO | 2023-09-18 | September 2023 | 17:25 | 2023-09-18 17:00 | 2023-09-18 17:25 | 0:25 | COMPLETED | SEMI-CRITICAL
417 | 2023-09-18 | Sep 2023 | 14:10 | 14:30 | 0.013889 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | RAMKIE | ICU | EXTRA FLUID | CHANGED THE DOOR COMPARTMENT | DOOR ASSEMBLY | NO | 2023-09-18 | September 2023 | 14:40 | 2023-09-18 14:10 | 2023-09-18 14:40 | 0:30 | COMPLETED | NON-CRITICAL
418 | 2023-09-18 | Sep 2023 | 15:10 | 15:40 | 0.020833 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | RAMKIE | WARD | NOT WORKING | RESET DONE | NO | NO | 2023-09-18 | September 2023 | 16:00 | 2023-09-18 15:10 | 2023-09-18 16:00 | 0:50 | COMPLETED | NON-CRITICAL
419 | 2023-09-19 | Sep 2023 | 7:35 | 7:40 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SOORAJ | WARD | SERVICE INDICATION BLINKING | FOUND PREOP CHECK NOT DONE. PREOP CHECK CARRIED OUT AND ISSUE RECTIFIED. | NO | NO | 2023-09-19 | September 2023 | 7:50 | 2023-09-19 07:35 | 2023-09-19 07:50 | 0:15 | COMPLETED | CRITICAL
420 | 2023-09-19 | Sep 2023 | 7:50 | 7:55 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SANTHOSH | EMERGENCY | ECG CHECK FAILED. | ECG CABLE LOOSE CONTACT. REFIXED AND PREOP CHECK DONE. ISSUE RECTIFIED. | NO | NO | 2023-09-19 | September 2023 | 8:00 | 2023-09-19 07:50 | 2023-09-19 08:00 | 0:10 | COMPLETED | CRITICAL
421 | 2023-09-19 | Sep 2023 | 21:10 | 21:12 | 0.0013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | HAFIZ | WARD | NOT WORKING | CHECKED THE COT AND FOUND THAT POWER SUPPLY END REMOVED. REFIXED AND NOW WORKING GOOD. | NO | NO | 2023-09-19 | September 2023 | 21:20 | 2023-09-19 21:10 | 2023-09-19 21:20 | 0:10 | COMPLETED | NON-CRITICAL
422 | 2023-09-19 | Sep 2023 | 22:10 | 22:12 | 0.0013889 | PULSE OXIMETER,PORTABLE | COVIDIEN | NELLCOR | HAFIZ | WARD | NOT WORKING | FOUND THAT POWER CABLE WAS LOOSE. REFIXED THE CABLE, CHECKED WITH THE SISTER AND NOW WORKING GOOD. | NO | NO | 2023-09-19 | September 2023 | 22:20 | 2023-09-19 22:10 | 2023-09-19 22:20 | 0:10 | COMPLETED | SEMI-CRITICAL
423 | 2023-09-19 | Sep 2023 | 16:20 | 16:25 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | RAMKIE | ICU | OP CHECK FAIL | PROBLEM WITH BATTERY, NEED REPLACEMNT | NO | BATTERY | 2023-09-19 | September 2023 | 16:30 | 2023-09-19 16:20 | 2023-09-19 16:30 | 0:10 | PENDING | CRITICAL
424 | 2023-09-20 | Sep 2023 | 0:50 | 0:55 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | HAFIZ | OT | CALIBRATION DUE | CALIBRATION DONE. NOW WORKING GOOD. | NO | NO | 2023-09-20 | September 2023 | 1:05 | 2023-09-20 00:50 | 2023-09-20 01:05 | 0:15 | COMPLETED | CRITICAL
425 | 2023-09-20 | Sep 2023 | 1:00 | 1:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | HAFIZ | OT | CALIBRATION DUE | CALIBRATION DONE. NOW WORKING GOOD. | NO | NO | 2023-09-20 | September 2023 | 1:15 | 2023-09-20 01:00 | 2023-09-20 01:15 | 0:15 | COMPLETED | CRITICAL
426 | 2023-09-20 | Sep 2023 | 1:35 | 1:40 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | HAFIZ | OT | CALIBRATION DUE | CALIBRATION DONE. NOW WORKING GOOD. | NO | NO | 2023-09-20 | September 2023 | 1:55 | 2023-09-20 01:35 | 2023-09-20 01:55 | 0:20 | COMPLETED | CRITICAL
427 | 2023-09-20 | Sep 2023 | 22:45 | 22:50 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | HAFIZ | ICU | NEED TO REPLACE A CATRIDGE. | REMOVED THE OLD CARTRIDGE AND REPLACED A NEW CARTRIDGE (S.NO; 301651435). READY TO USE. | CARTRIDGE | NO | 2023-09-20 | September 2023 | 23:00 | 2023-09-20 22:45 | 2023-09-20 23:00 | 0:15 | COMPLETED | SEMI-CRITICAL
428 | 2023-09-20 | Sep 2023 | 8:15 | 8:20 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | JERLIN | OT | ETCO2 MODULE NOT WORKING | WATER TRAP CHANGED AND RECTIFIED THE ISSUE | WATER TRAP | NO | 2023-09-20 | September 2023 | 8:30 | 2023-09-20 08:15 | 2023-09-20 08:30 | 0:15 | COMPLETED | SEMI-CRITICAL
429 | 2023-09-20 | Sep 2023 | 15:30 | 15:35 | 0.0034722 | GO FLO PUMP, ARTHROSCOPY | SMITH & NEPHEW | 72204968 | JERLIN | OT | NOT WORKING | CHECKED THE PUMP AND FOUND THAT IT'S IN GOOD WORKING CONDITION | NO | NO | 2023-09-20 | September 2023 | 15:50 | 2023-09-20 15:30 | 2023-09-20 15:50 | 0:20 | COMPLETED | SEMI-CRITICAL
430 | 2023-09-20 | Sep 2023 | 17:10 | 17:15 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SANTHOSH | EMERGENCY | REMOTE NOT WORKING | RESET DONE ,WORKING GOOD | NO | NO | 2023-09-20 | September 2023 | 17:20 | 2023-09-20 17:10 | 2023-09-20 17:20 | 0:10 | COMPLETED | NON-CRITICAL
431 | 2023-09-20 | Sep 2023 | 12:20 | 12:30 | 0.0069444 | INTRA ORAL XRAY | SKANRAY DENTAL | INTRA SKANDC | RAMKIE | OPD | REMOTE NOT WORKING | RECONNECTED THE RIBBON CABLE CONNECTOR AND NOW IT WORKING | NO | NO | 2023-09-20 | September 2023 | 12:30 | 2023-09-20 12:20 | 2023-09-20 12:30 | 0:10 | COMPLETED | SEMI-CRITICAL
432 | 2023-09-20 | Sep 2023 | 13:40 | 13:50 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | RAMKIE | WARD | NOT WORKING | FOUND THE POWER CORD GOT REMOVED FORM CTRL BOX, RECONNECTED AND RECTIFIED | NO | NO | 2023-09-20 | September 2023 | 14:00 | 2023-09-20 13:40 | 2023-09-20 14:00 | 0:20 | COMPLETED | NON-CRITICAL
433 | 2023-09-21 | Sep 2023 | 14:35 | 14:40 | 0.0034722 | CATHLAB - MONOPLANE | PHILIPS | AZURION 7M12 | SOORAJ | CATHLAB | STENT IMAGE NOT CLEAR | SETTINGS CHANGED AND ISSUE RECTIFIED. | NO | NO | 2023-09-21 | September 2023 | 15:00 | 2023-09-21 14:35 | 2023-09-21 15:00 | 0:25 | COMPLETED | CRITICAL
434 | 2023-09-21 | Sep 2023 | 14:45 | 14:45 | 0 | CATHLAB - BIPLANE | PHILIPS | AZURION 7B20 | SOORAJ | CATHLAB | EXAMINATION LIGHT ADJUSTMENT ISSUE | FOUND CONNECTOR BROKEN AND STUCK ON THE INTERFACE PANEL. WIRE ADJUSTMENT DONE AND WORKING. WORKING WITH CONSTRAINT. | NO | NO | 2023-09-21 | September 2023 | 15:00 | 2023-09-21 14:45 | 2023-09-21 15:00 | 0:15 | COMPLETED | CRITICAL
435 | 2023-09-21 | Sep 2023 | 14:55 | 14:55 | 0 | IVUS MACHINE | PHILIPS | CORE SERIES S5 | SOORAJ | CATHLAB | PACS CONNECTION REQUIRED. | CONFIGURATION DONE. ETHERNET SOCKET TO BE PROVIDED BY IT TEAM. | NO | NO | 2023-09-21 | September 2023 | 15:15 | 2023-09-21 14:55 | 2023-09-21 15:15 | 0:20 | COMPLETED | SEMI-CRITICAL
436 | 2023-09-21 | Sep 2023 | 16:45 | 16:50 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | ICU | BREAK AXEL BROKEN | BREAK AXEL SCREW REFIXED AND ISSUE RECTIFIED. | NO | NO | 2023-09-21 | September 2023 | 17:50 | 2023-09-21 16:45 | 2023-09-21 17:50 | 1:05 | COMPLETED | NON-CRITICAL
437 | 2023-09-21 | Sep 2023 | 20:10 | 20:12 | 0.0013889 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | HAFIZ | WARD | NOT WORKING | RESTARTED THE MACHINE. NOW WORKING GOOD. | NO | NO | 2023-09-21 | September 2023 | 20:20 | 2023-09-21 20:10 | 2023-09-21 20:20 | 0:10 | COMPLETED | SEMI-CRITICAL
438 | 2023-09-21 | Sep 2023 | 23:25 | 23:28 | 0.0020833 | PATIENT MONITOR | PHILIPS | MX 450 | HAFIZ | ICU | UNOBLE TO CHANGE THE NIBP SETTINGS. | CHECKED THE MACHINE, CHANGED THE SETTINGS AS NEEDED. NOW WORKING GOOD. | NO | NO | 2023-09-21 | September 2023 | 23:35 | 2023-09-21 23:25 | 2023-09-21 23:35 | 0:10 | COMPLETED | SEMI-CRITICAL
439 | 2023-09-21 | Sep 2023 | 0:20 | 0:23 | 0.0020833 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | HAFIZ | WARD | NOT WORKING | CHECKED THE MACHINE AND FOUND THAT BATTERY FULLY DRAINED. CONNECTED TO AC POWER SUPPLY. SWITCHED ON THE MACHINE. NOW WORKING GOOD. | NO | NO | 2023-09-21 | September 2023 | 0:30 | 2023-09-21 00:20 | 2023-09-21 00:30 | 0:10 | COMPLETED | SEMI-CRITICAL
440 | 2023-09-21 | Sep 2023 | 1:20 | 1:23 | 0.0020833 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | HAFIZ | WARD | NOT SWITCHING ON | CHECKED THE MACHINE AND FOUND THAT CABLE WAS REMOVED. REFIXED THE CABLE AT BACK. CHECKED AND NOW WORKING GOOD. | NO | NO | 2023-09-21 | September 2023 | 1:30 | 2023-09-21 01:20 | 2023-09-21 01:30 | 0:10 | COMPLETED | SEMI-CRITICAL
441 | 2023-09-21 | Sep 2023 | 14:00 | 14:10 | 0.0069444 | RADIANT HEAT WARMER | WIPRO GE HEALTH CARE | GIRAFFE | JERLIN | ICU | TEMPERATURE PROBE NOT DETECTED | REFIXED THE TEMPERATURE PROBE AND ISSUE RECTIFIED. | NO | NO | 2023-09-21 | September 2023 | 14:30 | 2023-09-21 14:00 | 2023-09-21 14:30 | 0:30 | COMPLETED | SEMI-CRITICAL
442 | 2023-09-21 | Sep 2023 | 13:00 | 13:10 | 0.0069444 | PATIENT MONITOR | PHILIPS | CM12 | JERLIN | ICU | BATTERY COMMUNICATION ERROR | REMOVED AND REFIXED THE BATTERY. ISSUE RECTIFIED. | NO | NO | 2023-09-21 | September 2023 | 13:15 | 2023-09-21 13:00 | 2023-09-21 13:15 | 0:15 | COMPLETED | SEMI-CRITICAL
443 | 2023-09-21 | Sep 2023 | 13:30 | 13:40 | 0.0069444 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | ICU | CLAMP BROKEN | FIXED THE CLAMP AND RECTIFED TH E ISSUE | NO | NO | 2023-09-21 | September 2023 | 13:55 | 2023-09-21 13:30 | 2023-09-21 13:55 | 0:25 | COMPLETED | NON-CRITICAL
444 | 2023-09-21 | Sep 2023 | 8:00 | 8:05 | 0.0034722 | OT TABLE | MAQUET | MEERA | JERLIN | OT | REMOTE NOT WORKING | FIXED IN ANOTHER PORT AND IT'S WORKING. | NO | NO | 2023-09-21 | September 2023 | 8:15 | 2023-09-21 08:00 | 2023-09-21 08:15 | 0:15 | COMPLETED | NON-CRITICAL
445 | 2023-09-21 | Sep 2023 | 15:00 | 15:10 | 0.0069444 | OT TABLE | MAQUET | ALPHA MAX | JERLIN | OT | COLUMN ALIGNMENT CHANGED | ALIGNMENT FIXED AND ISSUE RECTIFIED | NO | NO | 2023-09-21 | September 2023 | 15:30 | 2023-09-21 15:00 | 2023-09-21 15:30 | 0:30 | COMPLETED | NON-CRITICAL
446 | 2023-09-21 | Sep 2023 | 15:30 | 15:35 | 0.0034722 | OT TABLE | MAQUET | ALPHA MAX | JERLIN | OT | COLUMN ALIGNMENT CHANGED | ALIGNMENT FIXED AND ISSUE RECTIFIED | NO | NO | 2023-09-21 | September 2023 | 15:55 | 2023-09-21 15:30 | 2023-09-21 15:55 | 0:25 | COMPLETED | NON-CRITICAL
447 | 2023-09-21 | Sep 2023 | 17:30 | 17:35 | 0.0034722 | AUTOMATED ELISA PLATE WASHER | TRANSASIA | ERBA LISA WASH II | SANTHOSH | LABORATORY | TRAY MOVEMENT NOT OK | OBSERVED OPTICS OCCLUDED,THEN CLEANED TEST RUN | NO | NO | 2023-09-21 | September 2023 | 17:45 | 2023-09-21 17:30 | 2023-09-21 17:45 | 0:15 | COMPLETED | SEMI-CRITICAL
448 | 2023-09-21 | Sep 2023 | 17:00 | 17:05 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | SANTHOSH | OPD | CANT PUSHING TO PACS | CHECKED LAN CABLE AND DELETED SPOOLER, NOW ITS FINE | NO | NO | 2023-09-21 | September 2023 | 17:10 | 2023-09-21 17:00 | 2023-09-21 17:10 | 0:10 | COMPLETED | SEMI-CRITICAL
449 | 2023-09-21 | Sep 2023 | 18:00 | 18:10 | 0.0069444 | PRESSURE INJECTOR | LF OPTI VONTAGE | ILLUMENO NEO | RAMKIE | CATHLAB | ERROR 2003 DURING CASE | COMP PERSON CHECKED AND FOUND THE PROBLEM WITH INJECTOR HEAD, FOR NOW RECEIVED THE STANDBY MACHINE | NO | NO | 2023-09-21 | September 2023 | 18:30 | 2023-09-21 18:00 | 2023-09-21 18:30 | 0:30 | COMPLETED | SEMI-CRITICAL
450 | 2023-09-22 | Sep 2023 | 11:30 | 11:35 | 0.0034722 | FETAL HEART MONITOR | WIPRO GE HEALTH CARE | COROMETRICS 259 | SOORAJ | WARD | NOISE OBSERVED WHILE CONNECTING US PROBE. | CLEANED AND CHECKED, WORKING FINE. | NO | NO | 2023-09-22 | September 2023 | 11:50 | 2023-09-22 11:30 | 2023-09-22 11:50 | 0:20 | COMPLETED | SEMI-CRITICAL
451 | 2023-09-22 | Sep 2023 | 21:30 | 21:37 | 0.0048611 | BIPAP MACHINE | PHILIPS | A40 | HAFIZ | WARD | LOW CIRCUIT LEAK ALARM | OPENED THE O2 FLOW VALVE. NOW WORKING GOOD. | NO | NO | 2023-09-22 | September 2023 | 21:45 | 2023-09-22 21:30 | 2023-09-22 21:45 | 0:15 | COMPLETED | SEMI-CRITICAL
452 | 2023-09-22 | Sep 2023 | 13:30 | 13:45 | 0.010417 | TOURNIQUET,DUAL CUFF | VBM MEDZINTECHNIK | 9000 | JERLIN | OT | LEAKAGE IN CUFF | CHECKED AND FOUND IT'S IN GOOD WORKING CONDITION | NO | NO | 2023-09-22 | September 2023 | 14:00 | 2023-09-22 13:30 | 2023-09-22 14:00 | 0:30 | COMPLETED | SEMI-CRITICAL
453 | 2023-09-22 | Sep 2023 | 17:30 | 17:40 | 0.0069444 | CORE CONSOLE WITH HANDPIECE | STRYKER | 5400-052-000 | RAMKIE | OT | ATTACHMENT FIXING PROBLEM | CHECKED AND FOUND THE ATTACHMENT NOT HOLDING, SENT TO COMPANY SERVICE( LONG ST AND ANGLED ATTACHEMTN) | NO | NO | 2023-09-22 | September 2023 | 18:00 | 2023-09-22 17:30 | 2023-09-22 18:00 | 0:30 | COMPLETED | SEMI-CRITICAL
454 | 2023-09-23 | Sep 2023 | 15:30 | 15:40 | 0.0069444 | MICRO DRILL,DEBRIDING SYSTEM | MEDTRONICS | EC 300 | JERLIN | OT | WOBBLING ISSUE DURING CASE | PERFORATOR HAS BEEN OVER USED | NO | NO | 2023-09-23 | September 2023 | 15:50 | 2023-09-23 15:30 | 2023-09-23 15:50 | 0:20 | COMPLETED | SEMI-CRITICAL
455 | 2023-09-23 | Sep 2023 | 17:30 | 17:40 | 0.0069444 | SHAKER ,VDRL | REMI | VDS-300 | RAMKIE | LABORATORY | NOT WORKING | CHECKED AND CONNECTED THE BELT AND ALSO SUSPECTING THE PROBLEM WITH SENSOR | NO | NO | 2023-09-23 | September 2023 | 18:00 | 2023-09-23 17:30 | 2023-09-23 18:00 | 0:30 | COMPLETED | SEMI-CRITICAL
456 | 2023-09-24 | Sep 2023 | 15:00 | 15:10 | 0.0069444 | THROMBOELASTOMETER | FRAMER HEMOLOGIC | 5000 | JERLIN | ICU | MACHINE NOT WORKING | REFIXED THE COMMUNICATION CABLES AND RECTIFIED THE ISSUE | NO | NO | 2023-09-24 | September 2023 | 15:25 | 2023-09-24 15:00 | 2023-09-24 15:25 | 0:25 | COMPLETED | SEMI-CRITICAL
457 | 2023-09-24 | Sep 2023 | 7:30 | 7:30 | 0 | PATIENT MONITOR | PHILIPS | MX 550 | SANTHOSH | OT | SPO2PROBE NOT WORKING | REPLACED NEW PROBE | SPO2 PROBE | NO | 2023-09-24 | September 2023 | 7:45 | 2023-09-24 07:30 | 2023-09-24 07:45 | 0:15 | COMPLETED | SEMI-CRITICAL
458 | 2023-09-25 | Sep 2023 | 19:40 | 19:45 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SOORAJ | ICU | CARTRIDGE EMPTY | NEW CARTRIDGE REPLACED AND CVP DONE. | CARTRIDGE | NO | 2023-09-25 | September 2023 | 20:45 | 2023-09-25 19:40 | 2023-09-25 20:45 | 1:05 | COMPLETED | SEMI-CRITICAL
459 | 2023-09-25 | Sep 2023 | 13:10 | 13:15 | 0.0034722 | ELISA MICROPLATE READER | TRANSASIA | LISA SCAN EM | HAFIZ | LABORATORY | SELF TEST NOT PASSED | FOUND THAT TRAY MOVEMENT TEST FAILED. CLEANED THE TRAY, RESTARTED THE MACHINE. NOW TEST PASSED AND WORKING GOOD. | NO | NO | 2023-09-25 | September 2023 | 13:25 | 2023-09-25 13:10 | 2023-09-25 13:25 | 0:15 | COMPLETED | SEMI-CRITICAL
460 | 2023-09-25 | Sep 2023 | 7:50 | 7:55 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | SANTHOSH | OT | CIRCUIT LEAK TEST FAILED | REDUCED SOME EXECIVE SODALIME NOW ITS <100ML | NO | NO | 2023-09-25 | September 2023 | 8:00 | 2023-09-25 07:50 | 2023-09-25 08:00 | 0:10 | COMPLETED | CRITICAL
461 | 2023-09-25 | Sep 2023 | 17:10 | 17:20 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | RAMKIE | RADIOLOGY | PROBE NOT DETECTING | RECONNECTED ALL PROBES, RECTIFIED | NO | NO | 2023-09-25 | September 2023 | 17:40 | 2023-09-25 17:10 | 2023-09-25 17:40 | 0:30 | COMPLETED | SEMI-CRITICAL
462 | 2023-09-25 | Sep 2023 | 13:10 | 13:20 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | RAMKIE | RADIOLOGY | UNOBLE TO MOVE PACS | ETHERNET CABLE DETACHED FROM WALL POINT, RECTIFIED | NO | NO | 2023-09-25 | September 2023 | 13:30 | 2023-09-25 13:10 | 2023-09-25 13:30 | 0:20 | COMPLETED | SEMI-CRITICAL
463 | 2023-09-26 | Sep 2023 | 13:50 | 13:55 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | WARD | NOT WORKING | CHECKED THE COT AND FOUND THAT POWER SUPPLY END REMOVED. REFIXED AND NOW WORKING GOOD. | NO | NO | 2023-09-26 | September 2023 | 14:05 | 2023-09-26 13:50 | 2023-09-26 14:05 | 0:15 | COMPLETED | NON-CRITICAL
464 | 2023-09-26 | Sep 2023 | 19:00 | 19:15 | 0.010417 | NERVE STIMULATOR | INMED EQUIPMENTS | NS-100 | JERLIN | OT | CONNECTOR DISCONNECTED FROM ELECTRODE CABLE | SOLDERED THE CONNECTOR , CHECKED IT'S WORKING GOOD. | NO | NO | 2023-09-26 | September 2023 | 19:40 | 2023-09-26 19:00 | 2023-09-26 19:40 | 0:40 | COMPLETED | SEMI-CRITICAL
465 | 2023-09-26 | Sep 2023 | 8:52 | 8:55 | 0.0020833 | TOURNIQUET,DUAL CUFF | VBM MEDZINTECHNIK | 9000 | SANTHOSH | OT | REPORTED GET LEAK WHILE IN CASE | CHECKED 300 PRESSURE WITH 24&32 CUFFS NO LEAK FOUND | NO | NO | 2023-09-26 | September 2023 | 9:02 | 2023-09-26 08:52 | 2023-09-26 09:02 | 0:10 | COMPLETED | SEMI-CRITICAL
466 | 2023-09-26 | Sep 2023 | 16:20 | 16:30 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | RAMKIE | RADIOLOGY | OP CHECK FAIL | REPLACED THE NEW BATTERY, NOW IT GOT PASSED | NO | NO | 2023-09-26 | September 2023 | 16:45 | 2023-09-26 16:20 | 2023-09-26 16:45 | 0:25 | COMPLETED | CRITICAL
467 | 2023-09-27 | Sep 2023 | 20:05 | 20:10 | 0.0034722 | RADIOGRAPHIC UNIT | CARESTREAM | DRF ASCEND ANOLOG HF | HAFIZ | OPD | HANDLE WAS REMOVED | APPLIED THE TEFLON TAPE AND FIXED INTO THE MACHINE. NOW WORKING GOOD. | NO | NO | 2023-09-27 | September 2023 | 20:30 | 2023-09-27 20:05 | 2023-09-27 20:30 | 0:25 | COMPLETED | SEMI-CRITICAL
468 | 2023-09-27 | Sep 2023 | 16:30 | 16:40 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | JERLIN | OT | O2 CELL FAILED | RERAN CALIBRATION 21% & 100% PASSED. ISSUE RECTIFIED | NO | NO | 2023-09-27 | September 2023 | 16:55 | 2023-09-27 16:30 | 2023-09-27 16:55 | 0:25 | COMPLETED | CRITICAL
469 | 2023-09-27 | Sep 2023 | 10:30 | 10:35 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | SANTHOSH | ICU | INTERNOL STORAGE GOT FULL | CLEARED SOME OLD IMAGES | NO | NO | 2023-09-27 | September 2023 | 10:45 | 2023-09-27 10:30 | 2023-09-27 10:45 | 0:15 | COMPLETED | SEMI-CRITICAL
470 | 2023-09-27 | Sep 2023 | 17:30 | 17:35 | 0.0034722 | TOURNIQUET,DUAL CUFF | VBM MEDZINTECHNIK | 9000 | SANTHOSH | OT | GOT LEAK WHILE IN CASAE | FOUND LEAK IN CONNECTOR FIXED PROPERLY,ISSUE RECTIFIED | NO | NO | 2023-09-27 | September 2023 | 17:40 | 2023-09-27 17:30 | 2023-09-27 17:40 | 0:10 | COMPLETED | SEMI-CRITICAL
471 | 2023-09-28 | Sep 2023 | 13:30 | 13:35 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | SOORAJ | OT | PAEDIATRIC TEE PROBE NOT WORKING | FOUND ISSUE WITH NEW PROBE. REPALCEMENT TO BE GIVEN | NO | TEE PROBE | 2023-09-28 | September 2023 | 14:10 | 2023-09-28 13:30 | 2023-09-28 14:10 | 0:40 | PENDING | SEMI-CRITICAL
472 | 2023-09-28 | Sep 2023 | 17:55 | 18:00 | 0.0034722 | OT TABLE | MAQUET | MEERA | SOORAJ | OT | REMOTE DISPLAY HAVING ISSUES | DISPLAY REPLACED AND CHECKED, FOUND WORKING | REMOTE DISPLAY | NO | 2023-09-28 | September 2023 | 18:30 | 2023-09-28 17:55 | 2023-09-28 18:30 | 0:35 | COMPLETED | NON-CRITICAL
473 | 2023-09-28 | Sep 2023 | 22:00 | 22:10 | 0.0069444 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | WARD | IV POLE CLAMP BROKEN | REAR CASE CHANGED AND ISSU E RECTIFIED | REAR PANEL | NO | 2023-09-28 | September 2023 | 22:25 | 2023-09-28 22:00 | 2023-09-28 22:25 | 0:25 | COMPLETED | NON-CRITICAL
474 | 2023-09-28 | Sep 2023 | 23:00 | 23:10 | 0.0069444 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | OPD | DOOR BROKEN | CHANGED THE DOOR AND ISSUE RECTIFIED | DOOR ASSEMBLY | NO | 2023-09-28 | September 2023 | 23:30 | 2023-09-28 23:00 | 2023-09-28 23:30 | 0:30 | COMPLETED | NON-CRITICAL
475 | 2023-09-28 | Sep 2023 | 17:30 | 17:35 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | SANTHOSH | EMERGENCY | NOT GETTING ECG WAVE | REFIXED ECG CABLE AND CLEANED ITS CONNECTOR END | NO | NO | 2023-09-28 | September 2023 | 17:40 | 2023-09-28 17:30 | 2023-09-28 17:40 | 0:10 | COMPLETED | SEMI-CRITICAL
476 | 2023-09-28 | Sep 2023 | 18:10 | 18:30 | 0.013889 | ELECTROTHERAPHY WITHOUT LASER | CHATTANOOGA | INTELECT ADVANCED | RAMKIE | PHYSIOTHERAPY | NOT SWITCHING ON | CHECKED AND RECONNECTED, NOT ISSUES FOUND, IT WORKING | NO | NO | 2023-09-28 | September 2023 | 18:40 | 2023-09-28 18:10 | 2023-09-28 18:40 | 0:30 | COMPLETED | SEMI-CRITICAL
477 | 2023-09-29 | Sep 2023 | 14:15 | 14:22 | 0.0048611 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | HAFIZ | ICU | FIO2 VALUE NOT REACHED | ADJUSTED THE O2 CELL. NOW WORKING GOOD. | NO | NO | 2023-09-29 | September 2023 | 14:35 | 2023-09-29 14:15 | 2023-09-29 14:35 | 0:20 | COMPLETED | CRITICAL
478 | 2023-09-29 | Sep 2023 | 23:00 | 23:30 | 0.020833 | FLUID WARMER SYSTEM | MATRIX | MM 5 | JERLIN | OT | LEAKAGE IN THE BLANKET . | SEALED THE HOLES AND ISSUE RECTIFIED. | NO | NO | 2023-09-29 | September 2023 | 23:50 | 2023-09-29 23:00 | 2023-09-29 23:50 | 0:50 | COMPLETED | SEMI-CRITICAL
479 | 2023-09-29 | Sep 2023 | 7:30 | 7:35 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SANTHOSH | OT | CANNOT PRESSURE CIRCUIT | FOUND LEAK IN PATIENT CIRCUIT CHECKED WITH ANOTHER ONE THEN PASSED VENT TEST | NO | NO | 2023-09-29 | September 2023 | 7:40 | 2023-09-29 07:30 | 2023-09-29 07:40 | 0:10 | COMPLETED | CRITICAL
480 | 2023-09-29 | Sep 2023 | 10:30 | 10:35 | 0.0034722 | CEREBRAL OXIMETRY SYSTEM | COVIDIEN | SOMANETICS | SANTHOSH | OT | ONLY WORKS ONE SIDE AT TIME,IF BOTH SIDE OF SENSOR CONNECTS ,DOESN’T SHOW VALUES | CHECKED CABLES AND SENSOR AND RECTIFIED | NO | NO | 2023-09-29 | September 2023 | 10:40 | 2023-09-29 10:30 | 2023-09-29 10:40 | 0:10 | COMPLETED | SEMI-CRITICAL
481 | 2023-09-30 | Sep 2023 | 5:30 | 5:40 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | JERLIN | OT | CIRCUIT LEAK TEST FAILED | REFIXED THE CIRCUIT AND THE SODA LIME CHAMBER . RERAN THE TEST AND ISSUE RECTIFIED. | NO | NO | 2023-09-30 | September 2023 | 5:55 | 2023-09-30 05:30 | 2023-09-30 05:55 | 0:25 | COMPLETED | CRITICAL
482 | 2023-09-30 | Sep 2023 | 6:00 | 6:00 | 0 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | JERLIN | OT | O2 CELL ISSUE | CLEANED THE SENSOR AND THE CABLE , RERAN THE TEST . CALIBRATION PASSED AND ISSUE RECTIFIED. | NO | NO | 2023-09-30 | September 2023 | 6:15 | 2023-09-30 06:00 | 2023-09-30 06:15 | 0:15 | COMPLETED | CRITICAL
483 | 2023-09-30 | Sep 2023 | 6:15 | 6:15 | 0 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | JERLIN | OT | CALIBRATION FAILED | CLEANED THE O2 SENSOR, RERAN THE CALIBRATION TEST. TEST PASSED AND ISSUE RECTIFIED. | NO | NO | 2023-09-30 | September 2023 | 6:30 | 2023-09-30 06:15 | 2023-09-30 06:30 | 0:15 | COMPLETED | CRITICAL
484 | 2023-10-01 | Oct 2023 | 10:10 | 10:30 | 0.013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | EMERGENCY | NOT WORKING | CHECKED AND POWER INSERTED PROPERLY AND RECTIFIED | NO | NO | 2023-10-01 | October 2023 | 10:40 | 2023-10-01 10:10 | 2023-10-01 10:40 | 0:30 | COMPLETED | NON-CRITICAL
485 | 2023-10-02 | Oct 2023 | 13:30 | 13:40 | 0.0069444 | QUANTUM MACHINE | SMITH & NEPHEW | RF12000 | SOORAJ | OT | KEY FUNCTION NOT WORKING | PROBLEM IDENTIFIED WITH DISPLAY. NEW DISPLAY MODULE CHANGED AND CHECKED, WORKING FINE. | DISPLAY MODULE | NO | 2023-10-02 | October 2023 | 14:00 | 2023-10-02 13:30 | 2023-10-02 14:00 | 0:30 | COMPLETED | SEMI-CRITICAL
486 | 2023-10-03 | Oct 2023 | 12:45 | 12:48 | 0.0020833 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | OPD | PRINTER NOT WORKING | REMOVED THE OLD INK PRINTER PAPER, REPLACED IT WITH THE NEW PRINTER PAPER. NOW WORKING GOOD. | NO | NO | 2023-10-03 | October 2023 | 13:00 | 2023-10-03 12:45 | 2023-10-03 13:00 | 0:15 | COMPLETED | CRITICAL
487 | 2023-10-03 | Oct 2023 | 14:00 | 14:05 | 0.0034722 | INCUBATOR,TRANSPORT | PHOENIX | TINC 101 | HAFIZ | ICU | NOT WORKING | REPLACED THE BOARD AND NOW WORKING GOOD. | NO | NO | 2023-10-03 | October 2023 | 14:30 | 2023-10-03 14:00 | 2023-10-03 14:30 | 0:30 | COMPLETED | SEMI-CRITICAL
488 | 2023-10-03 | Oct 2023 | 14:25 | 14:30 | 0.0034722 | OXYGEN THERAPY UNIT | FISHER AND PAYKEL | AIRVO | HAFIZ | ICU | LEAK ALARM | CHECKED THE CIRCUITS, FINE. RESTARTED THE MACHINE. NOW, WORKING GOOD. | NO | NO | 2023-10-03 | October 2023 | 14:45 | 2023-10-03 14:25 | 2023-10-03 14:45 | 0:20 | COMPLETED | SEMI-CRITICAL
489 | 2023-10-03 | Oct 2023 | 7:10 | 7:15 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SOORAJ | OT | NOT SWITCHING ON | FOUND BATTERY DRAINED COMPLETELY. CONNECTED ON MAINS AND FOUND WORKING. | NO | NO | 2023-10-03 | October 2023 | 7:25 | 2023-10-03 07:10 | 2023-10-03 07:25 | 0:15 | COMPLETED | CRITICAL
490 | 2023-10-03 | Oct 2023 | 7:25 | 7:25 | 0 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | CANNISTER NOT FIXED PROPERLY. REFIXED AND ISSUE RECTIFIED. | NO | NO | 2023-10-03 | October 2023 | 7:35 | 2023-10-03 07:25 | 2023-10-03 07:35 | 0:10 | COMPLETED | CRITICAL
491 | 2023-10-03 | Oct 2023 | 10:20 | 10:25 | 0.0034722 | LENS INTEGRATED SYSTEM,ARTHROSCOPY | SMITH & NEPHEW | 72204354 | SOORAJ | OT | LAMP NOT WORKING. | CHECKED AND FOUND LAMP GOT DEFECTIVE. | NO | LAMP | 2023-10-03 | October 2023 | 11:00 | 2023-10-03 10:20 | 2023-10-03 11:00 | 0:40 | PENDING | SEMI-CRITICAL
492 | 2023-10-03 | Oct 2023 | 12:10 | 12:20 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | VOLUSON E10 | RAMKIE | OPD | C1 6 PROBE ARTIFACTS | NEED TO REPLACE THE C1 6 PROBE | NO | NO | 2023-10-03 | October 2023 | 12:40 | 2023-10-03 12:10 | 2023-10-03 12:40 | 0:30 | COMPLETED | SEMI-CRITICAL
493 | 2023-10-04 | Oct 2023 | 23:40 | 23:45 | 0.0034722 | OT TABLE | MAQUET | MEERA | SANTHOSH | OT | REMOTE SCREENGAURD PEELED OFF | GLUED PROPERLY ISSUE RECTIFIED | NO | NO | 2023-10-04 | October 2023 | 23:50 | 2023-10-04 23:40 | 2023-10-04 23:50 | 0:10 | COMPLETED | NON-CRITICAL
494 | 2023-10-04 | Oct 2023 | 16:30 | 16:35 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | HAFIZ | ICU | NEEDS TO REPLACE A NEW CARTRIDGE | REMOVED THE OLD CARTRIDGE AND REPLACED A NEW CARTRIDGE (S.NO: 301651529). WARMING UP. NOW READY TO USE. | CARTRIDGE | NO | 2023-10-04 | October 2023 | 17:30 | 2023-10-04 16:30 | 2023-10-04 17:30 | 1:00 | COMPLETED | SEMI-CRITICAL
495 | 2023-10-04 | Oct 2023 | 17:55 | 18:00 | 0.0034722 | EXAMINATION COUCH | MIDMARK JANOK | E122012 | HAFIZ | OPD | NOT WORKING | CHECKED THE MACHINE, WORKING GOOD. | NO | NO | 2023-10-04 | October 2023 | 18:15 | 2023-10-04 17:55 | 2023-10-04 18:15 | 0:20 | COMPLETED | NON-CRITICAL
496 | 2023-10-04 | Oct 2023 | 7:50 | 7:55 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | ICU | COT FUNCTION NOT WORKING. | FOUND BATTERY DEFECTIVE. REMOVED BATTERY CONNECTION AND RESET DONE. WORKING ON MAINS. | NO | BATTERY | 2023-10-04 | October 2023 | 8:15 | 2023-10-04 07:50 | 2023-10-04 08:15 | 0:25 | PENDING | NON-CRITICAL
497 | 2023-10-04 | Oct 2023 | 14:05 | 14:07 | 0.0013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SOORAJ | OT | BUTTON FUNCTIONS LOCKED. | FOUND ISSUE WITH CONTROL. | NO | CONTROL BOX | 2023-10-04 | October 2023 | 14:15 | 2023-10-04 14:05 | 2023-10-04 14:15 | 0:10 | PENDING | NON-CRITICAL
498 | 2023-10-04 | Oct 2023 | 15:20 | 15:40 | 0.013889 | RT POLYMERASE CHAIN REACTION | CO DX | MIC 9 | RAMKIE | LABORATORY | NOT WORKING | FOUND THE PROBLEM WITH POWER ADAPTER | NO | POWER ADAPTOR | 2023-10-04 | October 2023 | 15:55 | 2023-10-04 15:20 | 2023-10-04 15:55 | 0:35 | PENDING | SEMI-CRITICAL
499 | 2023-10-05 | Oct 2023 | 12:10 | 12:15 | 0.0034722 | FLUID WARMER SYSTEM | MATRIX | MN05 | SOORAJ | OT | BLANKET NOT WARMING UP. | FOUND WATER LEVEL LOW. FILLED WATER IN TANK AND CHECKED, ISSUE RECTIFIED. | NO | NO | 2023-10-05 | October 2023 | 12:30 | 2023-10-05 12:10 | 2023-10-05 12:30 | 0:20 | COMPLETED | SEMI-CRITICAL
500 | 2023-10-05 | Oct 2023 | 16:00 | 16:10 | 0.0069444 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | JERLIN | OPD | CABLE GRABBER NOT WORKING | FIXED THE GRABBER AND RECTIFIED THE ISSUE. | NO | NO | 2023-10-05 | October 2023 | 16:20 | 2023-10-05 16:00 | 2023-10-05 16:20 | 0:20 | COMPLETED | NON-CRITICAL
501 | 2023-10-05 | Oct 2023 | 15:30 | 15:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | WARD | NOT WORKING | REMOTE WAS LOCKED. UNLOCKED THE REMOTE AND RECTIFIED THE ISSUE. | NO | NO | 2023-10-05 | October 2023 | 15:45 | 2023-10-05 15:30 | 2023-10-05 15:45 | 0:15 | COMPLETED | NON-CRITICAL
502 | 2023-10-06 | Oct 2023 | 13:30 | 13:35 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | AUTOMATIC CATRIDGE EJECTION | REPLACED BY NEW CATRIDGE | NO | NO | 2023-10-06 | October 2023 | 13:45 | 2023-10-06 13:30 | 2023-10-06 13:45 | 0:15 | COMPLETED | SEMI-CRITICAL
503 | 2023-10-06 | Oct 2023 | 13:35 | 13:40 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | HAFIZ | ICU | NO+ RANGE DIFFERS | RUNNED THE CVP TEST 1 AND 2. NOW, NO+ IS IN NORMAL RANGE (127 MMOL/L) | NO | NO | 2023-10-06 | October 2023 | 13:50 | 2023-10-06 13:35 | 2023-10-06 13:50 | 0:15 | COMPLETED | SEMI-CRITICAL
504 | 2023-10-06 | Oct 2023 | 17:00 | 17:10 | 0.0069444 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX 8 | HAFIZ | OT | RF PCBA BOARD ISSUE | REPLACED ANOTHER POWERCORD. CHECKED THE MACHINE AND NOW WORKING GOOD. | POWERCORD | NO | 2023-10-06 | October 2023 | 17:30 | 2023-10-06 17:00 | 2023-10-06 17:30 | 0:30 | COMPLETED | SEMI-CRITICAL
505 | 2023-10-06 | Oct 2023 | 8:00 | 8:02 | 0.0013889 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | REDUCED SODA LIME QUANTITY AND CHECKED, ISSUE RECTIFIED. | NO | NO | 2023-10-06 | October 2023 | 8:10 | 2023-10-06 08:00 | 2023-10-06 08:10 | 0:10 | COMPLETED | CRITICAL
506 | 2023-10-06 | Oct 2023 | 8:10 | 8:10 | 0 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SOORAJ | OT | VAPORISER CHECK FAILED. | VAPORISER FIXED PROPERLY AND CHECKOUT DONE. ISSUE RECTIFIED. | NO | NO | 2023-10-06 | October 2023 | 8:20 | 2023-10-06 08:10 | 2023-10-06 08:20 | 0:10 | COMPLETED | CRITICAL
507 | 2023-10-06 | Oct 2023 | 14:30 | 14:40 | 0.0069444 | URINE ANALYSER | ROCHE | COBAS U 411 | RAMKIE | BLOODBANK | HIV TEST NOT POSSIBLE | CHECKED AND AMBIENT TEMP NOT GOOD, INFORMED COMPANY PERSON | NO | NO | 2023-10-06 | October 2023 | 14:55 | 2023-10-06 14:30 | 2023-10-06 14:55 | 0:25 | PENDING | SEMI-CRITICAL
508 | 2023-10-06 | Oct 2023 | 17:30 | 17:40 | 0.0069444 | PATIENT MONITOR | PHILIPS | CM12 | JERLIN | WARD | NBP NOT WORKING | FOUND LEAK IN THE CUFF AND REPLACED IT. | BP CUFF | NO | 2023-10-06 | October 2023 | 17:55 | 2023-10-06 17:30 | 2023-10-06 17:55 | 0:25 | COMPLETED | SEMI-CRITICAL
509 | 2023-10-07 | Oct 2023 | 22:15 | 22:20 | 0.0034722 | BABY WEIGHING SCALE WITH INFANTOMETER | SECA | 334 | HAFIZ | ICU | NOT WORKING | CHECKED WITH THE MACHINE. BATTERY NEEDS TO BE REPLACED. | NO | NO | 2023-10-07 | October 2023 | 22:30 | 2023-10-07 22:15 | 2023-10-07 22:30 | 0:15 | COMPLETED | NON-CRITICAL
510 | 2023-10-07 | Oct 2023 | 11:30 | 11:40 | 0.0069444 | ORTHOPANTOMOGRAPHY | SKANRAY DENTAL | MYRAY HYPERION X5 | RAMKIE | OPD | UNABLE TO USE | APP CORRUTED, COMPANY PERSONG REINSTALLED AND RECTIFIED | NO | NO | 2023-10-07 | October 2023 | 12:55 | 2023-10-07 11:30 | 2023-10-07 12:55 | 1:25 | COMPLETED | SEMI-CRITICAL
511 | 2023-10-09 | Oct 2023 | 17:20 | 17:27 | 0.0048611 | VENTILATOR, ADULT | MAQUET | SERVO-I | SANTHOSH | ICU | EXPIRATORY CASSETE TECHNICAL ERROR | REFIXED AND RESTARTED ISSUE RECTIFIED | NO | NO | 2023-10-09 | October 2023 | 17:35 | 2023-10-09 17:20 | 2023-10-09 17:35 | 0:15 | COMPLETED | CRITICAL
512 | 2023-10-09 | Oct 2023 | 6:50 | 6:55 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | HAFIZ | OT | REVERSE EXP. FLOW. CHECK VALVES OK?' ALARM | REMOVED THE CANISTER, REFIXED. RESTARTED THE MACHINE. PASSED THE FULL TEST. NOW WORKING GOOD. | NO | NO | 2023-10-09 | October 2023 | 7:10 | 2023-10-09 06:50 | 2023-10-09 07:10 | 0:20 | COMPLETED | CRITICAL
513 | 2023-10-09 | Oct 2023 | 9:15 | 9:18 | 0.0020833 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | HAFIZ | OPD | NOT WORKING | CLEARED THE SPOOLER, NOW WORKING GOOD. | NO | NO | 2023-10-09 | October 2023 | 9:25 | 2023-10-09 09:15 | 2023-10-09 09:25 | 0:10 | COMPLETED | SEMI-CRITICAL
514 | 2023-10-09 | Oct 2023 | 2:00 | 2:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | SOORAJ | OT | REVERSE FLOW ALARM | FLOW SENSORS DEFECTIVE. CHANGED NEW FLOW SENSORS AND ISSUE RECTIFIED. | FLOW SENSOR | NO | 2023-10-09 | October 2023 | 2:30 | 2023-10-09 02:00 | 2023-10-09 02:30 | 0:30 | COMPLETED | CRITICAL
515 | 2023-10-09 | Oct 2023 | 3:00 | 3:05 | 0.0034722 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | SOORAJ | WARD | SPO2 PARAMETER NOT SHOWING. | FOUND PINS BENT ON SPO2 CABLE. STRAIGHTENED THE PINS AND ISSUE RECTIFIED. | NO | NO | 2023-10-09 | October 2023 | 3:20 | 2023-10-09 03:00 | 2023-10-09 03:20 | 0:20 | COMPLETED | SEMI-CRITICAL
516 | 2023-10-09 | Oct 2023 | 6:20 | 6:25 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | FOUND ISSUE WITH PATIENT CIRCUIT. CHANGED CIRCUIT AND CHECKOUT PASSED. | NO | NO | 2023-10-09 | October 2023 | 6:35 | 2023-10-09 06:20 | 2023-10-09 06:35 | 0:15 | COMPLETED | CRITICAL
517 | 2023-10-09 | Oct 2023 | 6:25 | 6:25 | 0 | OT TABLE | MAQUET | MEERA | SOORAJ | OT | REMOTE DISPLAY CAME OFF | FIXED WITH FASTENER AND ISSUE RECTIFIED. | NO | NO | 2023-10-09 | October 2023 | 6:30 | 2023-10-09 06:25 | 2023-10-09 06:30 | 0:05 | COMPLETED | NON-CRITICAL
518 | 2023-10-09 | Oct 2023 | 9:30 | 9:40 | 0.0069444 | POLYMERASE CHAIN REACTION ANOLYSER | CEPHEID | GX 1V/2 | RAMKIE | LABORATORY | NOT WORKING SHOWS ERROR | FOUND THE PROBLEM WITH VOLTAGE FLUXTUATION | NO | NO | 2023-10-09 | October 2023 | 9:50 | 2023-10-09 09:30 | 2023-10-09 09:50 | 0:20 | COMPLETED | SEMI-CRITICAL
519 | 2023-10-09 | Oct 2023 | 14:30 | 14:40 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | VOLUSON S8 | JERLIN | OPD | PRINTER NOT WORKING | CHECKED AND IT'S WORKING GOOD. | NO | NO | 2023-10-09 | October 2023 | 14:50 | 2023-10-09 14:30 | 2023-10-09 14:50 | 0:20 | COMPLETED | SEMI-CRITICAL
520 | 2023-10-09 | Oct 2023 | 15:00 | 15:15 | 0.010417 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | JERLIN | RADIOLOGY | WORKS TOO SLOW | CLEARED SPOOLER AND IT'S WORKING GOOD. | NO | NO | 2023-10-09 | October 2023 | 15:50 | 2023-10-09 15:00 | 2023-10-09 15:50 | 0:50 | COMPLETED | SEMI-CRITICAL
521 | 2023-10-09 | Oct 2023 | 16:30 | 16:40 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | JERLIN | ICU | TGC PROBLEM | CHECKED AND IT"S WORKING FINE. | NO | NO | 2023-10-09 | October 2023 | 16:50 | 2023-10-09 16:30 | 2023-10-09 16:50 | 0:20 | COMPLETED | SEMI-CRITICAL
522 | 2023-10-09 | Oct 2023 | 15:20 | 15:20 | 0 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | JERLIN | RADIOLOGY | ABDOMEN PROBE NOT DETECTED | PROBE WAS NOT FIXED PROPERLY, FIXED IT AND ISSUE RECTIFIED | NO | NO | 2023-10-09 | October 2023 | 15:35 | 2023-10-09 15:20 | 2023-10-09 15:35 | 0:15 | COMPLETED | SEMI-CRITICAL
523 | 2023-10-10 | Oct 2023 | 9:10 | 9:13 | 0.0020833 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | HAFIZ | OT | LOW Fi02 ALARM | CHECKED THE CIRCUITS. RERUN THE CHECKOUT. NOW, PROBLEM RECTIFIED, READY TO USE. | NO | NO | 2023-10-10 | October 2023 | 9:30 | 2023-10-10 09:10 | 2023-10-10 09:30 | 0:20 | COMPLETED | CRITICAL
524 | 2023-10-10 | Oct 2023 | 18:10 | 18:14 | 0.0027778 | RADIANT HEAT WARMER | WIPRO GE HEALTH CARE | LULLABY | HAFIZ | ICU | LIGHT NOT WHEN WARMER IS ON | REMOVED THE CABLE FROM THE BOARD, RECONNECTED FIRMLY ALONG WITH RAMKIE SIR. NOW WORKING GOOD. | NO | NO | 2023-10-10 | October 2023 | 18:45 | 2023-10-10 18:10 | 2023-10-10 18:45 | 0:35 | COMPLETED | SEMI-CRITICAL
525 | 2023-10-10 | Oct 2023 | 1:18 | 1:20 | 0.0013889 | VENTILATOR, PEDIATRIC | MINDRAY MEDICAL INDIA | SV 300 | SOORAJ | ICU | BATTERY NOT CHARGING MORE THAN 40% | FOUND BATTERY TO BE DEFECTIVE | NO | BATTERY | 2023-10-10 | October 2023 | 1:25 | 2023-10-10 01:18 | 2023-10-10 01:25 | 0:07 | PENDING | CRITICAL
526 | 2023-10-10 | Oct 2023 | 11:30 | 11:40 | 0.0069444 | EXAMINATION COUCH | MIDMARK JANAK | E122012 | RAMKIE | OPD | HEAD POSITION MOT OPERATEABLE | GAS SPRING PROBLEM | NO | GAS SPRING | 2023-10-10 | October 2023 | 13:20 | 2023-10-10 11:30 | 2023-10-10 13:20 | 1:50 | PENDING | NON-CRITICAL
527 | 2023-10-10 | Oct 2023 | 15:30 | 15:40 | 0.0069444 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | ICU | CLAMP BROKEN | FIXED THE CLAMP AND ISSUE RECTIFIED. | NO | NO | 2023-10-10 | October 2023 | 15:55 | 2023-10-10 15:30 | 2023-10-10 15:55 | 0:25 | COMPLETED | NON-CRITICAL
528 | 2023-10-11 | Oct 2023 | 20:20 | 20:25 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SANTHOSH | WARD | DOWN KEY NOT WORKING | FOUND CPR VALVE OPEN CLOSE AND ISSUE RECIFIED | NO | NO | 2023-10-11 | October 2023 | 20:35 | 2023-10-11 20:20 | 2023-10-11 20:35 | 0:15 | COMPLETED | NON-CRITICAL
529 | 2023-10-11 | Oct 2023 | 8:16 | 8:20 | 0.0027778 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | HAFIZ | ICU | NOT WORKING | CHECKED THE MACHINE AND FOUND THAT INTERFACE CABLE GOT DETACHED AT THE BACK. FIXED IT. NOW WORKING GOOD. | NO | NO | 2023-10-11 | October 2023 | 8:26 | 2023-10-11 08:16 | 2023-10-11 08:26 | 0:10 | COMPLETED | SEMI-CRITICAL
530 | 2023-10-11 | Oct 2023 | 8:50 | 8:55 | 0.0034722 | PACEMAKER,DUAL CHAMBER | GETINGE | PACE T20 | HAFIZ | ICU | NOT WORKING | CHECKED THE MACHINE AND FOUND THAT AC INDICATION PRESENT BUT BATTERY INDICATION IS NOT. FURTHER, BATTERY END FULLY RUSTED. REMOVED THE BATTERY, CLEARED AND REFIXED THE BATTERY. NOW, CHECKED AND WORKING FINE. | NO | NO | 2023-10-11 | October 2023 | 9:10 | 2023-10-11 08:50 | 2023-10-11 09:10 | 0:20 | COMPLETED | CRITICAL
531 | 2023-10-11 | Oct 2023 | 9:30 | 9:35 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | ICU | DOOR BROKEN | REMOVED THE BROKEN DOOR, REFIXED WITH ANOTHER DOOR, CHECKED AND WORKING GOOD. | DOOR ASSEMBLY | NO | 2023-10-11 | October 2023 | 9:50 | 2023-10-11 09:30 | 2023-10-11 09:50 | 0:20 | COMPLETED | NON-CRITICAL
532 | 2023-10-11 | Oct 2023 | 9:50 | 9:55 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | ICU | KNOB BROKEN | REMOVED THE BROKEN KNOB, RFIXED WITH ANOTHER NOT WORKING INFUSION PUMP. CHECKED AND WORKING | KNOB | NO | 2023-10-11 | October 2023 | 10:10 | 2023-10-11 09:50 | 2023-10-11 10:10 | 0:20 | COMPLETED | NON-CRITICAL
533 | 2023-10-11 | Oct 2023 | 10:30 | 10:35 | 0.0034722 | BABY WEIGHING SCALE WITH INFANTOMETER | SECA | 334 | HAFIZ | ICU | NOT WORKING | CHECKED WITH THE MACHINE. BATTERY NEEDS TO BE REPLACED. REPLACED AND NOW WORKING GOOD. | NO | NO | 2023-10-11 | October 2023 | 10:40 | 2023-10-11 10:30 | 2023-10-11 10:40 | 0:10 | COMPLETED | NON-CRITICAL
534 | 2023-10-11 | Oct 2023 | 11:10 | 11:12 | 0.0013889 | VENTILATOR, ADULT | MAQUET | SERVO-I | HAFIZ | ICU | TECHNICAL ERROR MESSAGE | REMOVED THE EXPIRATORY CASSETTE, CLEANED AND REFIXED. REMOVED THE FILTER, CLEANED AND REFIXED. NOW, PASSED THE PRE-USE CHECK AND WORKING GOOD. | NO | NO | 2023-10-11 | October 2023 | 11:30 | 2023-10-11 11:10 | 2023-10-11 11:30 | 0:20 | COMPLETED | CRITICAL
535 | 2023-10-11 | Oct 2023 | 15:30 | 15:40 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | VOLUSON E10 | JERLIN | OPD | MACHINE NOT WORKING | RESTARTED THE MACHINE AND NOW IT'S WORKING GOOD. | NO | NO | 2023-10-11 | October 2023 | 15:55 | 2023-10-11 15:30 | 2023-10-11 15:55 | 0:25 | COMPLETED | SEMI-CRITICAL
536 | 2023-10-12 | Oct 2023 | 7:00 | 7:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | CHECKED THE CIRCUITS, REFIXED. CHECKOUT PASSED. NOW WORKING FINE. | NO | NO | 2023-10-12 | October 2023 | 7:20 | 2023-10-12 07:00 | 2023-10-12 07:20 | 0:20 | COMPLETED | CRITICAL
537 | 2023-10-12 | Oct 2023 | 18:00 | 18:05 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID E95 | JERLIN | OPD | MACHINE NOT WORKING | RESTARTED THE MACHINE AND NOW IT'S WORKING GOOD. | NO | NO | 2023-10-12 | October 2023 | 18:15 | 2023-10-12 18:00 | 2023-10-12 18:15 | 0:15 | COMPLETED | SEMI-CRITICAL
538 | 2023-10-12 | Oct 2023 | 17:30 | 17:40 | 0.0069444 | BLOOD GAS ANALYSER | GEM 3500 | INSTRUMENTATION LABORATORY | JERLIN | ICU | SCREEN WENT BLANK | RESTARTED THE MACHINE AND NOW IT'S WORKING GOOD. | NO | NO | 2023-10-12 | October 2023 | 17:55 | 2023-10-12 17:30 | 2023-10-12 17:55 | 0:25 | COMPLETED | SEMI-CRITICAL
539 | 2023-10-13 | Oct 2023 | 18:15 | 18:20 | 0.0034722 | OT TABLE | MAQUET | MEERA | SANTHOSH | OT | DISPLAY NOT WORKING | REPLACED NEW DISPLAY | REMOTE DISPLAY | NO | 2023-10-13 | October 2023 | 18:30 | 2023-10-13 18:15 | 2023-10-13 18:30 | 0:15 | COMPLETED | NON-CRITICAL
540 | 2023-10-13 | Oct 2023 | 18:30 | 18:35 | 0.0034722 | OT TABLE | MAQUET | MEERA | SANTHOSH | OT | DISPLAY NOT WORKING | REPLACED NEW DISPLAY | REMOTE DISPLAY | NO | 2023-10-13 | October 2023 | 18:45 | 2023-10-13 18:30 | 2023-10-13 18:45 | 0:15 | COMPLETED | NON-CRITICAL
541 | 2023-10-13 | Oct 2023 | 18:45 | 18:50 | 0.0034722 | OT TABLE | MAQUET | ALPHA MAX | SANTHOSH | OT | REMOTE NOT WORKING | SERVICED BY COMPANY PERSON | NO | NO | 2023-10-13 | October 2023 | 19:00 | 2023-10-13 18:45 | 2023-10-13 19:00 | 0:15 | COMPLETED | NON-CRITICAL
542 | 2023-10-13 | Oct 2023 | 12:30 | 12:32 | 0.0013889 | TRANSCRANIAL DOPPLER | VIASONIC | DOLPHIN 4D | HAFIZ | OT | 16PW PROBE NOT DETECTING | CHECKED THE MACHINE, ENOBLED THE INTRAOPERATIVE MODE, CHANGED IT TO THE INTRAOPERATIVE MODE FROM INTRACRANIAL MODE. NOW WORKING GOOD. | NO | NO | 2023-10-13 | October 2023 | 12:40 | 2023-10-13 12:30 | 2023-10-13 12:40 | 0:10 | COMPLETED | SEMI-CRITICAL
543 | 2023-10-13 | Oct 2023 | 12:40 | 12:42 | 0.0013889 | CAMERA SYSTEM | KARL STORZ | TC200 | HAFIZ | OT | VIDEO NOT CLEAR DURING THE PROCEDURE | REMOVED THE LIGHT SOURCE, CLEANED AND REFIXED. INCREASED THE INTENSITY, ADJUSTED THE BRIGHTNESS. NOW WORKING FINE. | NO | NO | 2023-10-13 | October 2023 | 12:50 | 2023-10-13 12:40 | 2023-10-13 12:50 | 0:10 | COMPLETED | SEMI-CRITICAL
544 | 2023-10-13 | Oct 2023 | 15:00 | 15:05 | 0.0034722 | MICROSCOPE OPERATING, SURGICAL | LEICA MICROSYSTEMS | LEICA OHX | HAFIZ | OT | XENON LAMP KIT HOURS LIMIT EXCEED | REMOVED THE OLD XENON LAMP, REFIXED A NEW XENON LAMP (MODEL:J2019, S.NO:23CRA5728) | NO | NO | 2023-10-13 | October 2023 | 15:30 | 2023-10-13 15:00 | 2023-10-13 15:30 | 0:30 | COMPLETED | SEMI-CRITICAL
545 | 2023-10-13 | Oct 2023 | 6:45 | 6:50 | 0.0034722 | THROMBOELASTOMETER | FRAMER HEMOLOGIC | HEMOLOGIC | SOORAJ | ICU | TEMPERATURE VALUES BLINKING. | CLEANED TEMPERATURE STRIPS AND RESTARTED THE SYSTEM, ISSUE RECTIFIED. | NO | NO | 2023-10-13 | October 2023 | 7:00 | 2023-10-13 06:45 | 2023-10-13 07:00 | 0:15 | COMPLETED | SEMI-CRITICAL
546 | 2023-10-13 | Oct 2023 | 16:10 | 16:30 | 0.013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | ICU | NOT WORKING | RESET DONE, NOW IT WORKING | NO | NO | 2023-10-13 | October 2023 | 16:40 | 2023-10-13 16:10 | 2023-10-13 16:40 | 0:30 | COMPLETED | NON-CRITICAL
547 | 2023-10-14 | Oct 2023 | 16:15 | 16:20 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SANTHOSH | CATHLAB | CIRCUIT LEAK TEST FAILED | TEST LUNG GOT TORE,CHECKED WITH ANOTHER ONE | NO | NO | 2023-10-14 | October 2023 | 16:30 | 2023-10-14 16:15 | 2023-10-14 16:30 | 0:15 | COMPLETED | CRITICAL
548 | 2023-10-14 | Oct 2023 | 17:30 | 17:35 | 0.0034722 | MRI MACHINE | PHILIPS | INGENIA ELITON S | RAMKIE | RADIOLOGY | TABLE MOVEMENT PROBLEM | CHECKED AND IT GOT OVERLAPPED WITH RAIL, SO RECTIFIED AND IT WORKING | NO | NO | 2023-10-14 | October 2023 | 17:55 | 2023-10-14 17:30 | 2023-10-14 17:55 | 0:25 | COMPLETED | CRITICAL
549 | 2023-10-15 | Oct 2023 | 13:45 | 13:50 | 0.0034722 | DEFIBRILLATOR MONITOR WITH AED | PHILIPS | EFFICIA DFM 100 | SANTHOSH | ICU | ECG MALFUNCTION | RUN OP CHECK ISSUE RECTIFIED | NO | NO | 2023-10-15 | October 2023 | 14:00 | 2023-10-15 13:45 | 2023-10-15 14:00 | 0:15 | COMPLETED | CRITICAL
550 | 2023-10-15 | Oct 2023 | 13:30 | 13:35 | 0.0034722 | CT MACHINE | PHILIPS | INGENUITY | SANTHOSH | RADIOLOGY | CIRS ERROR | RESTARTED GANTRY AND | NO | NO | 2023-10-15 | October 2023 | 13:40 | 2023-10-15 13:30 | 2023-10-15 13:40 | 0:10 | COMPLETED | CRITICAL
551 | 2023-10-15 | Oct 2023 | 16:40 | 16:45 | 0.0034722 | BIOCHEMISTRY ANALYSER | COBAS C311 | ROCHE | SANTHOSH | LABORATORY | CELL BLANK OUT OF LIMIT | RUN CELL BLANK MEASUREMENT ISSUE RECTIFIED | NO | NO | 2023-10-15 | October 2023 | 16:50 | 2023-10-15 16:40 | 2023-10-15 16:50 | 0:10 | COMPLETED | SEMI-CRITICAL
552 | 2023-10-16 | Oct 2023 | 13:30 | 13:35 | 0.0034722 | HEAD LIGHT SOURCE, XENON | SUNOPTIC | TITANX450S | SANTHOSH | OT | XENON LAMP REACHED ITS LIMIT | REPLACED NEW LAMP | XENON LAMP | NO | 2023-10-16 | October 2023 | 13:40 | 2023-10-16 13:30 | 2023-10-16 13:40 | 0:10 | COMPLETED | SEMI-CRITICAL
553 | 2023-10-16 | Oct 2023 | 15:10 | 15:15 | 0.0034722 | CATHLAB - BIPLANE | PHILIPS | AZURION 7B20 | SANTHOSH | CATHLAB | MONITOR REMOTE NOT WORKING | REPLACED NEW BATTERIES | BATTERY | NO | 2023-10-16 | October 2023 | 15:20 | 2023-10-16 15:10 | 2023-10-16 15:20 | 0:10 | COMPLETED | CRITICAL
554 | 2023-10-16 | Oct 2023 | 22:00 | 22:10 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | ICU | NOT WORKING | RESET DONE, STILL SMAE RESULT. UNPLUGED FROM AC SUPPLY. KEPT FOR SOME TIME. CHECKED AND WORKING GOOD. | NO | NO | 2023-10-16 | October 2023 | 22:30 | 2023-10-16 22:00 | 2023-10-16 22:30 | 0:30 | COMPLETED | NON-CRITICAL
555 | 2023-10-16 | Oct 2023 | 16:20 | 16:25 | 0.0034722 | RADIOGRAPHIC UNIT | CARESTREAM | DRF ASCEND ANOLOG HF | SOORAJ | RADIOLOGY | HANDLE ASSEMBLY BROKEN | NEW HANDLE ASEMBLY CHANGED AND ISSUE RECTIFIED. | HANDLE ASSEMBLY | NO | 2023-10-16 | October 2023 | 17:00 | 2023-10-16 16:20 | 2023-10-16 17:00 | 0:40 | COMPLETED | SEMI-CRITICAL
556 | 2023-10-16 | Oct 2023 | 16:30 | 16:40 | 0.0069444 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | RAMKIE | OT | NOT WORKING | FOUND THE WIRE DISCONNECTED, SO RECTIFIED | NO | NO | 2023-10-16 | October 2023 | 16:55 | 2023-10-16 16:30 | 2023-10-16 16:55 | 0:25 | COMPLETED | SEMI-CRITICAL
557 | 2023-10-17 | Oct 2023 | 17:35 | 17:40 | 0.0034722 | VIDEO PROCESSING SYSTEM,ENDOSCOPES | FUJIFILM CORPORATION | VP-3500HD | SANTHOSH | ICU | DISPLAY NOT WORKING | PROPERLY FIXED INTERFACE CABLE ISSUE RECTIFIED | NO | NO | 2023-10-17 | October 2023 | 17:50 | 2023-10-17 17:35 | 2023-10-17 17:50 | 0:15 | COMPLETED | SEMI-CRITICAL
558 | 2023-10-17 | Oct 2023 | 7:10 | 7:12 | 0.0013889 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | RERUN THE CHECKOUT. PASSED AND NOW WORKING GOOD. | NO | NO | 2023-10-17 | October 2023 | 7:30 | 2023-10-17 07:10 | 2023-10-17 07:30 | 0:20 | COMPLETED | CRITICAL
559 | 2023-10-17 | Oct 2023 | 22:10 | 22:15 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | ICU | PRINTER NOT WORKING | CHECKED THE MACHINE AND FOUND THAT PAPER WAS TOO OLD. CROSSCHECKED WITH ANOTHER DEFIBRILLATOR PAPER, WORKING GOOD. ASKED THEM TO CHANGE THE CURRENT DEFIBRILLATOT PAPER. | NO | NO | 2023-10-17 | October 2023 | 22:30 | 2023-10-17 22:10 | 2023-10-17 22:30 | 0:20 | COMPLETED | CRITICAL
560 | 2023-10-17 | Oct 2023 | 22:35 | 22:40 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | HAFIZ | WARD | NOT WORKING | CHECKED THE MACHINE AND FOUND THAT AC SUPPLY NOT CONNECTED. BATTERY FULLY DRAINED. CONNECTED TO AC POWER SOURCE. NOW, WORKING GOOD. | NO | NO | 2023-10-17 | October 2023 | 22:50 | 2023-10-17 22:35 | 2023-10-17 22:50 | 0:15 | COMPLETED | NON-CRITICAL
561 | 2023-10-17 | Oct 2023 | 9:30 | 9:40 | 0.0069444 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | RAMKIE | OT | NOT WORKING | FOUND THE WIRE DISCONNECTED, SO RECTIFIED | NO | NO | 2023-10-17 | October 2023 | 10:00 | 2023-10-17 09:30 | 2023-10-17 10:00 | 0:30 | COMPLETED | SEMI-CRITICAL
562 | 2023-10-17 | Oct 2023 | 14:00 | 14:05 | 0.0034722 | LENS INTEGRATED SYSTEM,ARTHROSCOPY | SMITH & NEPHEW | 72204354 | JERLIN | OT | DISPLAY IWAS TOO SMALL | ADJUSTED THE SETTINGS AND ISSUE RECTIFIED. | NO | NO | 2023-10-17 | October 2023 | 14:15 | 2023-10-17 14:00 | 2023-10-17 14:15 | 0:15 | COMPLETED | SEMI-CRITICAL
563 | 2023-10-17 | Oct 2023 | 15:00 | 15:10 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | JERLIN | ENDOSCOPY | CIRCUIT LEAK TEST FAILED | REMOVED A DEFECTIVE FILTER AND PERFORMED CHECKOUT. IT PASSESD. | NO | NO | 2023-10-17 | October 2023 | 15:25 | 2023-10-17 15:00 | 2023-10-17 15:25 | 0:25 | COMPLETED | CRITICAL
564 | 2023-10-17 | Oct 2023 | 7:30 | 7:40 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | JERLIN | CATHLAB | CIRCUIT LEAK TEST FAILED | REFIXED THE VAPOURIZER AND ISSUE ECTIFIED. | NO | NO | 2023-10-17 | October 2023 | 7:55 | 2023-10-17 07:30 | 2023-10-17 07:55 | 0:25 | COMPLETED | CRITICAL
565 | 2023-10-17 | Oct 2023 | 13:00 | 13:10 | 0.0069444 | VENTILATOR, ADULT | MAQUET | SERVO-I | JERLIN | ICU | TECHNICAL ERROR | REFIXED THE CASSETTE AND IT'S WORKING GOOD. | NO | NO | 2023-10-17 | October 2023 | 13:25 | 2023-10-17 13:00 | 2023-10-17 13:25 | 0:25 | COMPLETED | CRITICAL
566 | 2023-10-17 | Oct 2023 | 7:45 | 7:45 | 0 | PRESSURE INJECTOR | LF OPTI VONTAGE | ILLUMENO NEO | JERLIN | CATHLAB | NOT WORKING PROPERLY | CONNECTED A CATHETER AND CHECKED, IT'S WORKING GOOD. | NO | NO | 2023-10-17 | October 2023 | 7:55 | 2023-10-17 07:45 | 2023-10-17 07:55 | 0:10 | COMPLETED | SEMI-CRITICAL
567 | 2023-10-18 | Oct 2023 | 0:35 | 0:40 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | HAFIZ | ICU | O2 CONCENTRATION LOW | ADJUSTED THE O2 CELL CONCENTRATION. AND NOW WORKING GOOD. | NO | NO | 2023-10-18 | October 2023 | 0:50 | 2023-10-18 00:35 | 2023-10-18 00:50 | 0:15 | COMPLETED | CRITICAL
568 | 2023-10-18 | Oct 2023 | 21:00 | 21:05 | 0.0034722 | BIPAP MACHINE | RESMED | STELLAR 150 | HAFIZ | ICU | NOT ABLE TO CHANGE THE MODE | CHECKED THE MACHINE AND FOUND THAT THE SCREEN IS LOCKED. UNLOCKED IT, CHANGED THE SETTINGS ACCORDINGLY. NOW WORKING GOOD. | NO | NO | 2023-10-18 | October 2023 | 21:15 | 2023-10-18 21:00 | 2023-10-18 21:15 | 0:15 | COMPLETED | SEMI-CRITICAL
569 | 2023-10-18 | Oct 2023 | 14:45 | 14:48 | 0.0020833 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | SOORAJ | ICU | CARTRIDGE EMPTY | NEW CARTRDIGE REPLACED AND ISSUE RECTIFIED. | CARTRIDGE | NO | 2023-10-18 | October 2023 | 14:50 | 2023-10-18 14:45 | 2023-10-18 14:50 | 0:05 | COMPLETED | SEMI-CRITICAL
570 | 2023-10-18 | Oct 2023 | 17:25 | 17:30 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | WARD | COT LOCKED | RESET DONE AND ISSUE RECTIFIED. | NO | NO | 2023-10-18 | October 2023 | 17:35 | 2023-10-18 17:25 | 2023-10-18 17:35 | 0:10 | COMPLETED | NON-CRITICAL
571 | 2023-10-18 | Oct 2023 | 20:00 | 20:05 | 0.0034722 | WASHER DISINFECTOR | GETINGE | 46-5-403 | SOORAJ | CSSD | NOT ABLE TO CLOSE DOOR. | LIMIT SWITCH DID NOT ACTIVATE. CLEANED CONTACTS AND CHECKED, ISSUE RECTIFIED. | NO | NO | 2023-10-18 | October 2023 | 20:20 | 2023-10-18 20:00 | 2023-10-18 20:20 | 0:20 | COMPLETED | SEMI-CRITICAL
572 | 2023-10-18 | Oct 2023 | 12:00 | 12:00 | 0 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | ICU | BREAK PEDAL ISSUE | TEMPORARILY FIXED THE PANEL. NOW READY TO USE. | NO | NO | 2023-10-18 | October 2023 | 12:20 | 2023-10-18 12:00 | 2023-10-18 12:20 | 0:20 | COMPLETED | NON-CRITICAL
573 | 2023-10-18 | Oct 2023 | 8:45 | 8:55 | 0.0069444 | QUANTUM MACHINE | SMITH & NEPHEW | RF12000 | JERLIN | OT | DISPLAY NOT WORKING | CONNECTED THE CONNECTOR PROPERLY AND IT'S WORKING GOOD. | NO | NO | 2023-10-18 | October 2023 | 9:05 | 2023-10-18 08:45 | 2023-10-18 09:05 | 0:20 | COMPLETED | SEMI-CRITICAL
574 | 2023-10-18 | Oct 2023 | 8:15 | 8:20 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | JERLIN | ICU | NOT CONNECTING TO CENTRAL MONITOR | CHECKED WITH ANOTHER LAN CABLE, THEN IT WAS CONNECTED TO THE CENTRAL MONITOR. | NO | NO | 2023-10-18 | October 2023 | 8:35 | 2023-10-18 08:15 | 2023-10-18 08:35 | 0:20 | COMPLETED | SEMI-CRITICAL
575 | 2023-10-19 | Oct 2023 | 12:50 | 12:55 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | SANTHOSH | ICU | PACS NOT SENDING | CONFIGURED IP ,NOW PUSHING TO PACS | NO | NO | 2023-10-19 | October 2023 | 13:05 | 2023-10-19 12:50 | 2023-10-19 13:05 | 0:15 | COMPLETED | SEMI-CRITICAL
576 | 2023-10-19 | Oct 2023 | 10:40 | 11:00 | 0.013889 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | SANTHOSH | ICU | CANT REACH SIMV 18 | ADJUSTED BREATH RATE ISSUE RECTIFIED | NO | NO | 2023-10-19 | October 2023 | 11:20 | 2023-10-19 10:40 | 2023-10-19 11:20 | 0:40 | COMPLETED | CRITICAL
577 | 2023-10-19 | Oct 2023 | 12:00 | 12:00 | 0 | VENTILATOR, ADULT | MAQUET | SERVO-I | RAMKIE | ICU | Pre use check test failed | CLEANED THE DIAPHRAGM AND REFIXED THE EXPIRATORY CASSEETTEE. Test passed | NO | NO | 2023-10-19 | October 2023 | 12:20 | 2023-10-19 12:00 | 2023-10-19 12:20 | 0:20 | COMPLETED | CRITICAL
578 | 2023-10-19 | Oct 2023 | 7:15 | 7:20 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | JERLIN | OT | J PROBE NOT WORKING | REFIXED THE PROBE AND FOUND IT'S IN GOOD WORKING CONDITION. | NO | NO | 2023-10-19 | October 2023 | 7:35 | 2023-10-19 07:15 | 2023-10-19 07:35 | 0:20 | COMPLETED | SEMI-CRITICAL
579 | 2023-10-19 | Oct 2023 | 8:20 | 8:30 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | JERLIN | OT | CIRCUIT LEAK TEST FAILED | CHANGED THE CIRCUIT AND RAN THE TEST , IT PASSED. | NO | NO | 2023-10-19 | October 2023 | 8:40 | 2023-10-19 08:20 | 2023-10-19 08:40 | 0:20 | COMPLETED | CRITICAL
580 | 2023-10-19 | Oct 2023 | 8:15 | 8:30 | 0.010417 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX | JERLIN | OT | NOT WORKING PROPERLY | ISSUE RECTIFIED | NO | NO | 2023-10-19 | October 2023 | 8:40 | 2023-10-19 08:15 | 2023-10-19 08:40 | 0:25 | COMPLETED | SEMI-CRITICAL
581 | 2023-10-19 | Oct 2023 | 11:30 | 11:40 | 0.0069444 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX 8 | JERLIN | OT | ERROR 213 | RESTARTED THE MACHINE AND NOW IT'S WORKING GOOD. | NO | NO | 2023-10-19 | October 2023 | 11:55 | 2023-10-19 11:30 | 2023-10-19 11:55 | 0:25 | COMPLETED | SEMI-CRITICAL
582 | 2023-10-20 | Oct 2023 | 10:50 | 10:55 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | SANTHOSH | ICU | CATRIDGE EJECTED AUTOMATICALLLY | REPLACED ANOTHER CATRIDGE | CARTRIDGE | NO | 2023-10-20 | October 2023 | 11:05 | 2023-10-20 10:50 | 2023-10-20 11:05 | 0:15 | COMPLETED | SEMI-CRITICAL
583 | 2023-10-20 | Oct 2023 | 12:00 | 12:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SANTHOSH | OT | NOT SWITCHING ON | THEN ISSUE RECTIFIED | NO | NO | 2023-10-20 | October 2023 | 12:10 | 2023-10-20 12:00 | 2023-10-20 12:10 | 0:10 | COMPLETED | CRITICAL
584 | 2023-10-20 | Oct 2023 | 0:55 | 1:00 | 0.0034722 | STEAM STERILIZER | GETINGE | GSS67H102E | HAFIZ | CSSD | VACCUM TIMEOUT ALARM | RERUN THE BOWIE-DICK TEST, OBSERVED AND NOW WORKING GOOD. | NO | NO | 2023-10-20 | October 2023 | 1:25 | 2023-10-20 00:55 | 2023-10-20 01:25 | 0:30 | COMPLETED | SEMI-CRITICAL
585 | 2023-10-20 | Oct 2023 | 3:00 | 3:05 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | HAFIZ | ICU | NEED TO REPLACE A CARTRIDGE | REMOVED THE OLD CARTRIDGE AND REPLACED IT WITH A NEW CARTRIDGE (LOT;230818). NOW, WARMING UP | NO | NO | 2023-10-20 | October 2023 | 3:14 | 2023-10-20 03:00 | 2023-10-20 03:14 | 0:14 | COMPLETED | SEMI-CRITICAL
586 | 2023-10-20 | Oct 2023 | 22:25 | 22:30 | 0.0034722 | VENTILATOR, ADULT | MAQUET | SERVO-I | HAFIZ | ICU | TECHNICAL ERROR AT EXPIRATORY CASSETTE | REMOVED THE OLD EXPIRATORY CASETTE (S.NO: 293141), REFIXED ANOTHER EXP. CASSETTE (S.NO:293337) FROM THE DEPT. PASSED THE PRE-USE CHECK AND NOW WORKING GOOD. | NO | NO | 2023-10-20 | October 2023 | 22:50 | 2023-10-20 22:25 | 2023-10-20 22:50 | 0:25 | COMPLETED | CRITICAL
587 | 2023-10-20 | Oct 2023 | 14:30 | 14:30 | 0 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | ICU | NOT WORKING | REMOTE RESET DONE | NO | NO | 2023-10-20 | October 2023 | 14:40 | 2023-10-20 14:30 | 2023-10-20 14:40 | 0:10 | COMPLETED | NON-CRITICAL
588 | 2023-10-20 | Oct 2023 | 16:30 | 16:40 | 0.0069444 | POLYMERASE CHAIN REACTION ANOLYSER | CEPHEID | GX 1V/2 | RAMKIE | LABORATORY | NOT WORKING | FOUND THE WIRE DISCCONECTION, SO RECTIFIED | NO | NO | 2023-10-20 | October 2023 | 16:55 | 2023-10-20 16:30 | 2023-10-20 16:55 | 0:25 | COMPLETED | SEMI-CRITICAL
589 | 2023-10-21 | Oct 2023 | 6:05 | 6:10 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | WARD | BED LOCKED, BUTTON NOT WORKING | UNLOCKED THE COT, CHECKED AND NOW WORKING GOOD. | NO | NO | 2023-10-21 | October 2023 | 6:15 | 2023-10-21 06:05 | 2023-10-21 06:15 | 0:10 | COMPLETED | NON-CRITICAL
590 | 2023-10-21 | Oct 2023 | 9:30 | 9:40 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | JERLIN | OT | NOT SWITCHING ON | CONNECTED IN ANOTHER SOCKET AND IT'S WORKING GOOD. | NO | NO | 2023-10-21 | October 2023 | 9:50 | 2023-10-21 09:30 | 2023-10-21 09:50 | 0:20 | COMPLETED | CRITICAL
591 | 2023-10-23 | Oct 2023 | 13:30 | 13:35 | 0.0034722 | AUTOMATED MICROBIAL ANALYSER | BIOMERIUX | VITEK 2 COMPACT 30 | HAFIZ | LABORATORY | NOT WORKING | CHECKED THE MACHINE AND FOUND THAT CONNECTION PORT WAS UNPLUGGED. CONNECTED BACK AND CHECKED. WORKING GOOD. | NO | NO | 2023-10-23 | October 2023 | 13:45 | 2023-10-23 13:30 | 2023-10-23 13:45 | 0:15 | COMPLETED | SEMI-CRITICAL
592 | 2023-10-23 | Oct 2023 | 13:50 | 13:54 | 0.0027778 | XENON LIGHT SOURCE,ENDOSCOPE | FUJIFILM CORPORATION | XL-4450 | HAFIZ | OT | VIDEO NOT DISPLAYED DURING THE PROCEDURE. | CHECKED AND FOUND THAT VIDEO INTERFACE CABLE WAS NOT PROPERLY FIXED. FIXED PROPERLY. NOW, ISSUE SOLVED AND WORKING GOOD. | NO | NO | 2023-10-23 | October 2023 | 14:04 | 2023-10-23 13:50 | 2023-10-23 14:04 | 0:14 | COMPLETED | SEMI-CRITICAL
593 | 2023-10-24 | Oct 2023 | 10:10 | 10:15 | 0.0034722 | HARMONIC GENERATOR | LOTUS | UBERSONIC TM | SANTHOSH | OT | CONNECTER PIN GET OUT | FIXED PROPERLY ISSUE RECTIFIED | NO | NO | 2023-10-24 | October 2023 | 10:20 | 2023-10-24 10:10 | 2023-10-24 10:20 | 0:10 | COMPLETED | SEMI-CRITICAL
594 | 2023-10-24 | Oct 2023 | 17:30 | 17:35 | 0.0034722 | STEAM STERILIZER | GETINGE | GSS67H102E | RAMKIE | CSSD | STERILIZER DOOR NOT OPENING | OBSERVED THE PRESSURE TO BE REACHED ABOVE 9 BAR. REDUCED THE PRESSURE AND KEPT IT UNDER 8 BAR. ISSUE RECTIFIED. | NO | NO | 2023-10-24 | October 2023 | 17:50 | 2023-10-24 17:30 | 2023-10-24 17:50 | 0:20 | COMPLETED | SEMI-CRITICAL
595 | 2023-10-24 | Oct 2023 | 22:30 | 22:35 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | JERLIN | OT | CALIBRATION FAILED | O2 SENSOR CLEANED AND REFIXED | O2 SENSOR | NO | 2023-10-24 | October 2023 | 22:45 | 2023-10-24 22:30 | 2023-10-24 22:45 | 0:15 | COMPLETED | CRITICAL
596 | 2023-10-24 | Oct 2023 | 2:00 | 2:10 | 0.0069444 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | JERLIN | ICU | NOT WORKING | CHANGED THE HOSE AND ISSUE RECTIFIED | WARMER HOSE | NO | 2023-10-24 | October 2023 | 2:20 | 2023-10-24 02:00 | 2023-10-24 02:20 | 0:20 | COMPLETED | SEMI-CRITICAL
597 | 2023-10-24 | Oct 2023 | 1:30 | 1:40 | 0.0069444 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | JERLIN | ICU | NOT WORKING | CONNECTED THE SENSOR CABLE PROPERLY AND IT'S WORKING GOOD. | NO | NO | 2023-10-24 | October 2023 | 1:55 | 2023-10-24 01:30 | 2023-10-24 01:55 | 0:25 | COMPLETED | SEMI-CRITICAL
598 | 2023-10-25 | Oct 2023 | 14:35 | 14:40 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | HAFIZ | OPD | MACHINE HANGING DURING THE PROCEDURE | CLEARED THE SPOOLER, NOW WORKING GOOD. | NO | NO | 2023-10-25 | October 2023 | 14:50 | 2023-10-25 14:35 | 2023-10-25 14:50 | 0:15 | COMPLETED | SEMI-CRITICAL
599 | 2023-10-25 | Oct 2023 | 20:30 | 20:45 | 0.010417 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | JERLIN | ICU | NOT SWITCHING ON | IT WAS REMOVED FROM THE POWER SOCKET. FIXED IT PROPERLY | NO | NO | 2023-10-25 | October 2023 | 20:55 | 2023-10-25 20:30 | 2023-10-25 20:55 | 0:25 | COMPLETED | SEMI-CRITICAL
600 | 2023-10-25 | Oct 2023 | 10:30 | 10:40 | 0.0069444 | BIPAP MACHINE | PHILIPS | A40 | JERLIN | WARD | CIRCUIT DISCONNECTED ALARM | CONNECTED THE CIRCUIT PROPERLY AND ISSUE RECTIFIED. | NO | NO | 2023-10-25 | October 2023 | 10:55 | 2023-10-25 10:30 | 2023-10-25 10:55 | 0:25 | COMPLETED | SEMI-CRITICAL
601 | 2023-10-25 | Oct 2023 | 23:00 | 23:15 | 0.010417 | BIPAP MACHINE | PHILIPS | A40 | JERLIN | WARD | LOW LEAK ALARM | ADJUSTED THE OUTLET PORT ACCORDINGLY, AND ISSUE RECTIFIED. | NO | NO | 2023-10-25 | October 2023 | 23:30 | 2023-10-25 23:00 | 2023-10-25 23:30 | 0:30 | COMPLETED | SEMI-CRITICAL
602 | 2023-10-25 | Oct 2023 | 23:00 | 23:35 | 0.024306 | PATIENT MONITOR | PHILIPS | MX 450 | JERLIN | ICU | STAND IS LOOSELY CONNECTED | FIXED THE STAND PROPERLY AND ISSUE RECTIFIED. | NO | NO | 2023-10-25 | October 2023 | 23:45 | 2023-10-25 23:00 | 2023-10-25 23:45 | 0:45 | COMPLETED | SEMI-CRITICAL
603 | 2023-10-26 | Oct 2023 | 18:30 | 18:35 | 0.0034722 | VENTILATOR, ADULT | MAQUET | SERVI I | RAMKIE | ICU | AIR LEAKAGE | DISCONNECTED AND RECONNECTED THE AIRWAAY CIRCUIT. ISSUE RECTIFIED | NO | NO | 2023-10-26 | October 2023 | 18:45 | 2023-10-26 18:30 | 2023-10-26 18:45 | 0:15 | COMPLETED | CRITICAL
604 | 2023-10-26 | Oct 2023 | 1:00 | 1:15 | 0.010417 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | JERLIN | OT | CIRCUIT LEAK TEST FAILED | RERAN TEST AND IT PASSED. | NO | NO | 2023-10-26 | October 2023 | 1:25 | 2023-10-26 01:00 | 2023-10-26 01:25 | 0:25 | COMPLETED | CRITICAL
605 | 2023-10-26 | Oct 2023 | 21:30 | 21:40 | 0.0069444 | BIPAP MACHINE | PHILIPS | A40 | JERLIN | WARD | CIRCUIT DISCONNECTED ALARM | REFIXED THE CIRCUIT AND ISSUE RECTIFIED | NO | NO | 2023-10-26 | October 2023 | 21:55 | 2023-10-26 21:30 | 2023-10-26 21:55 | 0:25 | COMPLETED | SEMI-CRITICAL
606 | 2023-10-27 | Oct 2023 | 8:00 | 8:05 | 0.0034722 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX 8 | JERLIN | OT | E232 ERROR | RESTARTED SELF TEST DONE ISSUE RECTIFIED. | NO | NO | 2023-10-27 | October 2023 | 8:10 | 2023-10-27 08:00 | 2023-10-27 08:10 | 0:10 | COMPLETED | SEMI-CRITICAL
607 | 2023-10-27 | Oct 2023 | 9:30 | 9:35 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SANTHOSH | OT | O2 CELL ISSUE | NEW O2 CELL REPLACED | O2 CELL | NO | 2023-10-27 | October 2023 | 9:40 | 2023-10-27 09:30 | 2023-10-27 09:40 | 0:10 | COMPLETED | CRITICAL
608 | 2023-10-27 | Oct 2023 | 22:30 | 22:40 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 450 | JERLIN | EMERGENCY | IBP NOT WORKING | ZEROED AND CONNECTED AGAIN. IT'S WORKING GOOD. | NO | NO | 2023-10-27 | October 2023 | 22:55 | 2023-10-27 22:30 | 2023-10-27 22:55 | 0:25 | COMPLETED | SEMI-CRITICAL
609 | 2023-10-28 | Oct 2023 | 10:10 | 10:14 | 0.0027778 | VENTILATOR, UNIVERSAL | MINDRAY | SV 300 | HAFIZ | ICU | HIGH LEAKAGE, VENTILATION DISABLED | REMOVED THE EXPIRATORY VALVE AND THE FILTER. FOUND CLOT INSIDE. REMOVED IT, AND PASSED THE SYSTEM CHECK AND NOW WORKING GOOD. | NO | NO | 2023-10-28 | October 2023 | 10:30 | 2023-10-28 10:10 | 2023-10-28 10:30 | 0:20 | COMPLETED | CRITICAL
610 | 2023-10-28 | Oct 2023 | 10:30 | 10:32 | 0.0013889 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | HAFIZ | ICU | HIGH LEAKAGE | REMOVED THE GAS SUPPLY AND REFIXED. NOW, PASSED THE PRE-USE CHECK AND WORKING GOOD. | NO | NO | 2023-10-28 | October 2023 | 10:50 | 2023-10-28 10:30 | 2023-10-28 10:50 | 0:20 | COMPLETED | CRITICAL
611 | 2023-10-28 | Oct 2023 | 10:50 | 10:55 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | HAFIZ | ICU | NEEDS TO INSTALL A NEW CARTRIDGE | REMOVED THE OLD CARTRIDGE AND REPLACED IT WITH A NEW CARTRIDGE. NOW, CVP TESTS DONE AND READY TO USE. | NO | NO | 2023-10-28 | October 2023 | 23:50 | 2023-10-28 10:50 | 2023-10-28 23:50 | 13:00 | COMPLETED | SEMI-CRITICAL
612 | 2023-10-28 | Oct 2023 | 17:10 | 17:13 | 0.0020833 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | WARD | EQUIPMENT DISABLED, SYSTEM FAILURE | FOUND THAT THERAPY KNOB TEST FAILED. REMOVED THE KNOB, AND REFIXED. NOW, PASSED THE OPERATIONAL CHECK AND WORKING GOOD. | NO | NO | 2023-10-28 | October 2023 | 17:25 | 2023-10-28 17:10 | 2023-10-28 17:25 | 0:15 | COMPLETED | CRITICAL
613 | 2023-10-28 | Oct 2023 | 16:30 | 16:35 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | RAMKIE | OT | GAS MODULE STANDBY LIGHT GLOWING CONTINUOUSLY AND EPCO2 NOT DETECTING | REMOVED AND REINSERTED GAS MODULE. STANDBY ISSUE RECTIFIED. REPLACED WATER TRAP AND OSBERVED EPCO2 VALUE ON DISPLAY. | GAS MODULE | NO | 2023-10-28 | October 2023 | 16:45 | 2023-10-28 16:30 | 2023-10-28 16:45 | 0:15 | COMPLETED | SEMI-CRITICAL
614 | 2023-10-28 | Oct 2023 | 23:30 | 23:40 | 0.0069444 | MRI MACHINE | PHILIPS | INGENIA ELITON S | JERLIN | RADIOLOGY | NOT WORKING | RESTARTED THE MACHINE AND NOW IT'S WORKING GOOD. | NO | NO | 2023-10-28 | October 2023 | 23:55 | 2023-10-28 23:30 | 2023-10-28 23:55 | 0:25 | COMPLETED | CRITICAL
615 | 2023-10-28 | Oct 2023 | 22:30 | 22:40 | 0.0069444 | VENTILATOR, PEDIATRIC | MINDRAY MEDICAL INDIA | SV 300 | JERLIN | EMERGENCY | NEED TO DO PRE-USE CHECK | CHECKED AND IT'S WORKING GOOD. | NO | NO | 2023-10-28 | October 2023 | 22:55 | 2023-10-28 22:30 | 2023-10-28 22:55 | 0:25 | COMPLETED | CRITICAL
616 | 2023-10-30 | Oct 2023 | 21:30 | 21:32 | 0.0013889 | BIPAP MACHINE | RESMED | STELLAR 150 | HAFIZ | WARD | ALARM INDICATION | UNLOCKED THE MACHINE, CHANGES DONE AND WORKING GOOD. | NO | NO | 2023-10-30 | October 2023 | 21:40 | 2023-10-30 21:30 | 2023-10-30 21:40 | 0:10 | COMPLETED | SEMI-CRITICAL
617 | 2023-10-30 | Oct 2023 | 12:00 | 12:06 | 0.0041667 | BABY WEIGHING SCALE | SECA | 354 | HAFIZ | ICU | NOT WORKING | CHECKED THE MACHINE, AND FOUND THAT BATTERY NOT FIXED PROPERLY. REFIXED IT AND CHECKED. NOW, WORKING GOOD. | NO | NO | 2023-10-30 | October 2023 | 12:15 | 2023-10-30 12:00 | 2023-10-30 12:15 | 0:15 | COMPLETED | NON-CRITICAL
618 | 2023-10-30 | Oct 2023 | 7:10 | 7:15 | 0.0034722 | PACEMAKER,DUAL CHAMBER | GETINGE | PACE T20 | SOORAJ | ICU | NOT SWITCHING ON | FOUND BATTERY DEFECTIVE. CHANGED NEW BATTERIES AND ISSUE RECTIFIED. | BATTERY | NO | 2023-10-30 | October 2023 | 7:25 | 2023-10-30 07:10 | 2023-10-30 07:25 | 0:15 | COMPLETED | CRITICAL
619 | 2023-10-30 | Oct 2023 | 11:45 | 11:50 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | SOORAJ | OT | NOT SWITCHING ON | RESTARTED THE SYSTEM AND ISSUE RECTIFIED. | NO | NO | 2023-10-30 | October 2023 | 12:00 | 2023-10-30 11:45 | 2023-10-30 12:00 | 0:15 | COMPLETED | CRITICAL
620 | 2023-10-30 | Oct 2023 | 16:45 | 16:50 | 0.0034722 | TABLE TOP STEAM STERILIZER | GETINGE | HS 22K7 | RAMKIE | OT | WATER QA ERROR | REPLACED WATER IN THE TANK AND RESTARTED THE UNIT. ISSUE RECTIFIED AND WORKING FINE. | NO | NO | 2023-10-30 | October 2023 | 17:00 | 2023-10-30 16:45 | 2023-10-30 17:00 | 0:15 | COMPLETED | SEMI-CRITICAL
621 | 2023-10-30 | Oct 2023 | 17:35 | 17:40 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | ICU | PRINTER NOT WORKING | FIXED THE PAPER PROPERLY AND ISSUE RECTIFIED | NO | NO | 2023-10-30 | October 2023 | 17:55 | 2023-10-30 17:35 | 2023-10-30 17:55 | 0:20 | COMPLETED | CRITICAL
622 | 2023-10-30 | Oct 2023 | 18:00 | 18:05 | 0.0034722 | USG MACHINE |  | LOGIC E-R7 | JERLIN | OT | ECHO PROBE NOT WORKING | REFIXED THE PROBE AND FOUND IT'S IN GOOD WORKING CONDITION. | NO | NO | 2023-10-30 | October 2023 | 18:10 | 2023-10-30 18:00 | 2023-10-30 18:10 | 0:10 | COMPLETED | SEMI-CRITICAL
623 | 2023-10-30 | Oct 2023 | 15:00 | 15:15 | 0.010417 | PATIENT MONITOR | PHILIPS | CM12 | JERLIN | WARD | MONITOR NOT WORKING | RESTARTED THE MACHINE AND NOW IT'S WORKING GOOD. | NO | NO | 2023-10-30 | October 2023 | 15:20 | 2023-10-30 15:00 | 2023-10-30 15:20 | 0:20 | COMPLETED | SEMI-CRITICAL
624 | 2023-10-31 | Oct 2023 | 1:10 | 1:15 | 0.0034722 | VENTILATOR, ADULT | MAQUET | SERVO-I | SANTHOSH | ICU | INTERNOL LEAK | CHECKOUT DONE ISSUE RECTIFIED | NO | NO | 2023-10-31 | October 2023 | 1:20 | 2023-10-31 01:10 | 2023-10-31 01:20 | 0:10 | COMPLETED | CRITICAL
625 | 2023-10-31 | Oct 2023 | 17:10 | 17:15 | 0.0034722 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | SANTHOSH | WARD | BP NOT READ | FIXED CABLE PROPERLY | NO | NO | 2023-10-31 | October 2023 | 17:20 | 2023-10-31 17:10 | 2023-10-31 17:20 | 0:10 | COMPLETED | SEMI-CRITICAL
626 | 2023-10-31 | Oct 2023 | 17:30 | 17:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SANTHOSH | OT | REPORTED NOT WORKING | CHECKED CONNECTIONS AND RESET | NO | NO | 2023-10-31 | October 2023 | 17:45 | 2023-10-31 17:30 | 2023-10-31 17:45 | 0:15 | COMPLETED | NON-CRITICAL
627 | 2023-10-31 | Oct 2023 | 9:00 | 9:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | NO FLOW SENSOR ALARM | REFIXED FLOW SENSORS AND CHECKOUT DONE. ISSUE RECTIFIED. | NO | NO | 2023-10-31 | October 2023 | 9:20 | 2023-10-31 09:00 | 2023-10-31 09:20 | 0:20 | COMPLETED | CRITICAL
628 | 2023-10-31 | Oct 2023 | 10:00 | 10:10 | 0.0069444 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | THANISH | ICU | AIR FILTER BLOCKAGE | CLEANED THE AIR FILTR, ISSUE RECTIFIED | NO | NO | 2023-10-31 | October 2023 | 10:30 | 2023-10-31 10:00 | 2023-10-31 10:30 | 0:30 | COMPLETED | SEMI-CRITICAL
629 | 2023-10-31 | Oct 2023 | 10:30 | 10:34 | 0.0027778 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | THANISH | ICU | WARMER HOSE ISSUE | SOLDERED WIRE AND RECTIFIED THE WARMER HOSE | NO | NO | 2023-10-31 | October 2023 | 11:00 | 2023-10-31 10:30 | 2023-10-31 11:00 | 0:30 | COMPLETED | SEMI-CRITICAL
630 | 2023-10-31 | Oct 2023 | 11:00 | 11:10 | 0.0069444 | BIPAP MACHINE | PHILIPS | A40 | THANISH | WARD | LANGUAGE CHANGED | SETTINGS CHANGED, ISSUE RECTIFIED | NO | NO | 2023-10-31 | October 2023 | 11:30 | 2023-10-31 11:00 | 2023-10-31 11:30 | 0:30 | COMPLETED | SEMI-CRITICAL
631 | 2023-11-01 | Nov 2023 | 9:15 | 9:20 | 0.0034722 | VENTILATOR, PEDIATRIC | MINDRAY MEDICAL INDIA | SV 300 | THANISH | ICU | O2 CELL NOT WORKING | CLEANED THE O2 CELL AND PERFORMED THE 02 CELL CALIBRATION.WORKING FINE. | NO | NO | 2023-11-01 | November 2023 | 9:30 | 2023-11-01 09:15 | 2023-11-01 09:30 | 0:15 | COMPLETED | CRITICAL
632 | 2023-11-01 | Nov 2023 | 12:20 | 12:26 | 0.0041667 | BMI MACHINE | WELCH ALLYN | TANITA | THANISH | OPD | HEIGHT SCALE NOT WORKING | RE-ATTACHED THE HEIGHT SCALE.NOW WORKING FINE. | NO | NO | 2023-11-01 | November 2023 | 12:40 | 2023-11-01 12:20 | 2023-11-01 12:40 | 0:20 | COMPLETED | NON-CRITICAL
633 | 2023-11-01 | Nov 2023 | 12:50 | 12:55 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | THANISH | ICU | WARMER NOT WORKING | AIR FILTER BLCOKAGE ISSUE/CLEANED THE AIR FILTER.WORKING FINE. | NO | NO | 2023-11-01 | November 2023 | 13:15 | 2023-11-01 12:50 | 2023-11-01 13:15 | 0:25 | COMPLETED | SEMI-CRITICAL
634 | 2023-11-01 | Nov 2023 | 14:10 | 14:15 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | THANISH | ICU | WARMER NOT WORKING | AIR FILTER BLCOKAGE ISSUE/CLEANED THE AIR FILTER.WORKING FINE. | NO | NO | 2023-11-01 | November 2023 | 14:35 | 2023-11-01 14:10 | 2023-11-01 14:35 | 0:25 | COMPLETED | SEMI-CRITICAL
635 | 2023-11-01 | Nov 2023 | 14:40 | 14:45 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | THANISH | WARD | COT NOT WORKING | RESETTED THE REMOTE AND CONTROL BOX.WORKING FINE. | NO | NO | 2023-11-01 | November 2023 | 14:50 | 2023-11-01 14:40 | 2023-11-01 14:50 | 0:10 | COMPLETED | NON-CRITICAL
636 | 2023-11-01 | Nov 2023 | 15:05 | 15:10 | 0.0034722 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | THANISH | WARD | SPO2 CONNECTOR DAMAGED | RECTIFIED THE SPO2 CONNECTOR ISSUE .WORKING FINE. | NO | NO | 2023-11-01 | November 2023 | 15:40 | 2023-11-01 15:05 | 2023-11-01 15:40 | 0:35 | COMPLETED | SEMI-CRITICAL
637 | 2023-11-01 | Nov 2023 | 7:55 | 8:00 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | REMOVED SODA LIME AND DONE PRE USE CHECK PASSED | NO | NO | 2023-11-01 | November 2023 | 8:15 | 2023-11-01 07:55 | 2023-11-01 08:15 | 0:20 | COMPLETED | CRITICAL
638 | 2023-11-01 | Nov 2023 | 10:20 | 10:25 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | SOORAJ | OT | GAS MODULE NOT WORKING | FOUND ISSUE DUE TO HUMIDITY. REMOVED GAS MODULE AND KEPT IN ROOM TEMPERATURE FOR 4 HRS. REFIXED AND ISSUE RECTIFIED. | NO | NO | 2023-11-01 | November 2023 | 14:30 | 2023-11-01 10:20 | 2023-11-01 14:30 | 4:10 | COMPLETED | SEMI-CRITICAL
639 | 2023-11-01 | Nov 2023 | 13:50 | 13:55 | 0.0034722 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | HAFIZ | OPD | ARTIFACT ISSUE | CHECKED THE MACHINE AND FOUND ARTIFACTS IN 1ST AND 2ND ROUND. RESTARTED THE MACHINE, AND REFIXED THE CONNECTIONS. OBSERVED WITH A PATIENT AND WORKING FINE. | NO | NO | 2023-11-01 | November 2023 | 14:20 | 2023-11-01 13:50 | 2023-11-01 14:20 | 0:30 | COMPLETED | NON-CRITICAL
640 | 2023-11-01 | Nov 2023 | 11:20 | 11:40 | 0.013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | OT | NOT WORKING | RESET DONE | NO | NO | 2023-11-01 | November 2023 | 11:50 | 2023-11-01 11:20 | 2023-11-01 11:50 | 0:30 | COMPLETED | NON-CRITICAL
641 | 2023-11-01 | Nov 2023 | 1:10 | 1:15 | 0.0034722 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | SANTHOSH | WARD | BP NOT READING | CHECKED AND OBSERVED LOOSENED BP HOSE FIXED PROPERLY | NO | NO | 2023-11-01 | November 2023 | 1:20 | 2023-11-01 01:10 | 2023-11-01 01:20 | 0:10 | COMPLETED | SEMI-CRITICAL
642 | 2023-11-01 | Nov 2023 | 7:48 | 7:55 | 0.0048611 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SANTHOSH | OT | NOT WORKING | CHECKED AND OBSERVED SUSPECT MACHINE ACCUATER GOT DEFECT | NO | ACTUATOR | 2023-11-01 | November 2023 | 8:00 | 2023-11-01 07:48 | 2023-11-01 08:00 | 0:12 | PENDING | NON-CRITICAL
643 | 2023-11-02 | Nov 2023 | 10:00 | 10:15 | 0.010417 | AUTOMATED MICROBIAL ANALYSER | BIOMERIUX | VITEK 2 COMPACT 30 | JERLIN | LABORATORY | HOST COMMUNICATION PROBLEM | REFIXED THE CABLE AND RECTIFIED THE ISSUE | NO | NO | 2023-11-02 | November 2023 | 10:25 | 2023-11-02 10:00 | 2023-11-02 10:25 | 0:25 | COMPLETED | SEMI-CRITICAL
644 | 2023-11-02 | Nov 2023 | 18:40 | 18:45 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | JERLIN | ICU | TEMP NOT DISPLAYED | FOUND ISSUE WITH THE TEMP PROBE AND CHECKED WITH ANOTHER PROBE AND IT'S WORKING GOOD. | NO | NO | 2023-11-02 | November 2023 | 19:00 | 2023-11-02 18:40 | 2023-11-02 19:00 | 0:20 | COMPLETED | SEMI-CRITICAL
645 | 2023-11-02 | Nov 2023 | 9:03 | 9:05 | 0.0013889 | OT TABLE | MAQUET | ALPHA MAX | SOORAJ | OT | REMOTE NOT WORKING. | CHANGED TO ANOTHER PORT AND CHECKED, WORKING FINE. | NO | NO | 2023-11-02 | November 2023 | 9:10 | 2023-11-02 09:03 | 2023-11-02 09:10 | 0:07 | COMPLETED | NON-CRITICAL
646 | 2023-11-02 | Nov 2023 | 11:00 | 11:05 | 0.0034722 | C-MAC MONITOR | KARL STORZ | 84032ZX | SOORAJ | OT | AUTOMATCIALLY SWITCH OFF | FOUN CONNECTOR NOT FIT PROPERLY WHILE USING. FIXED PROPERLY AND ISSUE RECTIFIED | NO | NO | 2023-11-02 | November 2023 | 11:15 | 2023-11-02 11:00 | 2023-11-02 11:15 | 0:15 | COMPLETED | SEMI-CRITICAL
647 | 2023-11-02 | Nov 2023 | 9:50 | 9:55 | 0.0034722 | PULSE OXIMETER,PORTABLE | MASIMO | RAD-97 | HAFIZ | ICU | NOT WORKING | REMOVED THE POWERCHORD AND REFIXED. RESTARTED THE MACHINE AND NOW WORKING GOOD. | NO | NO | 2023-11-02 | November 2023 | 10:15 | 2023-11-02 09:50 | 2023-11-02 10:15 | 0:25 | COMPLETED | SEMI-CRITICAL
648 | 2023-11-02 | Nov 2023 | 10:20 | 10:22 | 0.0013889 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | ICU | DATE AND TME CHANGED | CHANGES DONE, RERUN THE OPERATIONAL CHECK AND WORKING GOOD. | NO | NO | 2023-11-02 | November 2023 | 10:30 | 2023-11-02 10:20 | 2023-11-02 10:30 | 0:10 | COMPLETED | CRITICAL
649 | 2023-11-02 | Nov 2023 | 13:10 | 13:15 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | HCT UNOVAIL | ISSUE RECTIFIED BY COMPANY PERSON | NO | NO | 2023-11-02 | November 2023 | 13:20 | 2023-11-02 13:10 | 2023-11-02 13:20 | 0:10 | COMPLETED | SEMI-CRITICAL
650 | 2023-11-03 | Nov 2023 | 9:15 | 9:20 | 0.0034722 | PRINTER,COMPUTED RADIOGRAPHY | CARESTREAM | DRY VIEW 6950 | THANISH | RADIOLOGY | FILTER ISSUE | RECTIFIED THE FILTER ISSUE.WORKING FINE. | NO | NO | 2023-11-03 | November 2023 | 9:30 | 2023-11-03 09:15 | 2023-11-03 09:30 | 0:15 | COMPLETED | SEMI-CRITICAL
651 | 2023-11-03 | Nov 2023 | 19:00 | 19:05 | 0.0034722 | URINE ANALYSER | COBAS U 411 | ROCHE | JERLIN | LABORATORY | TEMPERATURE ERROR | RESTARTED THE MACHINE AND ISSUE RECTIFIED | NO | NO | 2023-11-03 | November 2023 | 19:15 | 2023-11-03 19:00 | 2023-11-03 19:15 | 0:15 | COMPLETED | SEMI-CRITICAL
652 | 2023-11-03 | Nov 2023 | 20:00 | 20:10 | 0.0069444 | BIPAP MACHINE | PHILIPS | A40 | JERLIN | ICU | NOT WORKING | CHECKED AND ITS WORKING GOOD | NO | NO | 2023-11-03 | November 2023 | 20:20 | 2023-11-03 20:00 | 2023-11-03 20:20 | 0:20 | COMPLETED | SEMI-CRITICAL
653 | 2023-11-03 | Nov 2023 | 10:05 | 10:10 | 0.0034722 | MRI MACHINE | PHILIPS | INGENIA ELITON S | SOORAJ | RADIOLOGY | VERTICAL LINE ON CONSOLE DISPLAY. | FOUND ISSUE DUE TO STORAGE FULL. CLEARED DATA AND ISSUE RECTIFIED. | NO | NO | 2023-11-03 | November 2023 | 10:30 | 2023-11-03 10:05 | 2023-11-03 10:30 | 0:25 | COMPLETED | CRITICAL
654 | 2023-11-03 | Nov 2023 | 12:30 | 12:40 | 0.0069444 | IVUS MACHINE | PHILIPS | CORE SERIES S5 | RAMKIE | CATHLAB | INTERMITTENTLY SLAVE MONITOR SIGNOL GETS LOST | FOUND THE PROBLEM WITH IVUS CONNECTING PORT PROBLEM, NEED REPLACEMENT | NO | NO | 2023-11-03 | November 2023 | 12:55 | 2023-11-03 12:30 | 2023-11-03 12:55 | 0:25 | COMPLETED | SEMI-CRITICAL
655 | 2023-11-04 | Nov 2023 | 11:15 | 11:20 | 0.0034722 | PULSE OXIMETER,PORTABLE | COVIDIEN | NELLCOR | HAFIZ | WARD | NOT WORKING | REMOVED THE EXTENSION CABLE, CLEANED AND REFIXED. CHECKED AND NOW WORKING GOOD. | NO | NO | 2023-11-04 | November 2023 | 11:30 | 2023-11-04 11:15 | 2023-11-04 11:30 | 0:15 | COMPLETED | SEMI-CRITICAL
656 | 2023-11-04 | Nov 2023 | 16:50 | 16:55 | 0.0034722 | DEFIBRILLATOR MONITOR WITH AED | PHILIPS | EFFICIA DFM 100 | HAFIZ | ICU | PRINTER ISSUE | CHECKED AND FOUND ISSUE IN PRINTER PAPER. REPLACED WITH ANOTHER PAPER. REPLACED WITH ANOTHER PAPER. CHECKED AND WORKING GOOD. | NO | NO | 2023-11-04 | November 2023 | 17:10 | 2023-11-04 16:50 | 2023-11-04 17:10 | 0:20 | COMPLETED | CRITICAL
657 | 2023-11-04 | Nov 2023 | 11:00 | 11:10 | 0.0069444 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | RAMKIE | OT | BUBBLE SENSOR BROKEN | REPLACED BUBBLE SENSOR | BUBBLE SENSOR | NO | 2023-11-04 | November 2023 | 11:40 | 2023-11-04 11:00 | 2023-11-04 11:40 | 0:40 | COMPLETED | NON-CRITICAL
658 | 2023-11-05 | Nov 2023 | 11:00 | 11:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | ICU | PRINTER NOT WORKING | PAPER WAS NOT FIXED PROPERLY, FIXED AND ISSUE RECTIFIED | NO | NO | 2023-11-05 | November 2023 | 11:20 | 2023-11-05 11:00 | 2023-11-05 11:20 | 0:20 | COMPLETED | CRITICAL
659 | 2023-11-05 | Nov 2023 | 20:25 | 20:30 | 0.0034722 | STEAM STERILIZER | GETINGE | GSS67H102E | SOORAJ | CSSD | NO WATER ALARM | RESET DONE AND ISSUE RECTIFIED. | NO | NO | 2023-11-05 | November 2023 | 20:40 | 2023-11-05 20:25 | 2023-11-05 20:40 | 0:15 | COMPLETED | SEMI-CRITICAL
660 | 2023-11-05 | Nov 2023 | 22:25 | 22:30 | 0.0034722 | STEAM STERILIZER | GETINGE | GSS67H102E | SOORAJ | CSSD | AIR IN LINE ALARM | FOUND TUBE NNOT INSERTED PROPERLY. REFIXED AND ISSUE RECTIFIED. | NO | NO | 2023-11-05 | November 2023 | 22:45 | 2023-11-05 22:25 | 2023-11-05 22:45 | 0:20 | COMPLETED | SEMI-CRITICAL
661 | 2023-11-06 | Nov 2023 | 10:10 | 10:15 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | THANISH | OT | MONITOR SWITCHING OFF AUTOMATICALLY | CLEANED THE MAIN BOARD AND RECONNECTED THE CABLE.WORKING FINE. | NO | NO | 2023-11-06 | November 2023 | 10:40 | 2023-11-06 10:10 | 2023-11-06 10:40 | 0:30 | COMPLETED | SEMI-CRITICAL
662 | 2023-11-06 | Nov 2023 | 9:25 | 9:30 | 0.0034722 | LENS INTEGRATED SYSTEM,ARTHROSCOPY | SMITH & NEPHEW | 72204354 | THANISH | OT | VIDEO NOT RECORDING | FORMATTED THE PENDRIVE.WORKING FINE. | NO | NO | 2023-11-06 | November 2023 | 9:40 | 2023-11-06 09:25 | 2023-11-06 09:40 | 0:15 | COMPLETED | SEMI-CRITICAL
663 | 2023-11-06 | Nov 2023 | 11:05 | 11:10 | 0.0034722 | MICROSCOPE OPERATING, SURGICAL | LEICA MICROSYSTEMS | LEICA OHX | THANISH | OT | VIDEO NOT RECORDING | DELETED THE OLD PATIENT DATA.WORKING FINE. | NO | NO | 2023-11-06 | November 2023 | 11:30 | 2023-11-06 11:05 | 2023-11-06 11:30 | 0:25 | COMPLETED | SEMI-CRITICAL
664 | 2023-11-06 | Nov 2023 | 3:30 | 3:35 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SOORAJ | ICU | PRINT NOT WORKING. | SENSOR CLEANED AND ISSUE RECTIFIED. | NO | NO | 2023-11-06 | November 2023 | 3:45 | 2023-11-06 03:30 | 2023-11-06 03:45 | 0:15 | COMPLETED | CRITICAL
665 | 2023-11-06 | Nov 2023 | 7:10 | 7:15 | 0.0034722 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX 8 | HAFIZ | OT | ES232 ERROR | REMOVED THE POWERCHORD AND REFIXED. RESTARTED THE MACHINE AND NOW WORKING GOOD. | NO | NO | 2023-11-06 | November 2023 | 7:30 | 2023-11-06 07:10 | 2023-11-06 07:30 | 0:20 | COMPLETED | SEMI-CRITICAL
666 | 2023-11-06 | Nov 2023 | 7:25 | 7:30 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | HAFIZ | OT | NOT WORKING | RESTARTED THE MACHINE, WORKING GOOD. STILL, HOSE FULLY DAMAGED. REPLACED A NEW HOSE. | NO | NO | 2023-11-06 | November 2023 | 7:45 | 2023-11-06 07:25 | 2023-11-06 07:45 | 0:20 | COMPLETED | SEMI-CRITICAL
667 | 2023-11-06 | Nov 2023 | 12:00 | 12:05 | 0.0034722 | ECHO MACHINE | GE HEALTHCARE | VIVID S70 | HAFIZ | CATHLAB | SLAVE MONITOR DISPLAY ISSUE | FOUND ISSUE IN DVI CABLE, REPLACED ANOTHER DVI CABLE. ISSUE RECTIFIED | DVI CABLE | NO | 2023-11-06 | November 2023 | 12:30 | 2023-11-06 12:00 | 2023-11-06 12:30 | 0:30 | COMPLETED | SEMI-CRITICAL
668 | 2023-11-06 | Nov 2023 | 13:20 | 13:24 | 0.0027778 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | HAFIZ | ICU | AIRWAY PRESSURE HIGH ALARM | FOUND ISSUE IN CIRCUIT. REFIXED, AND NOW WORKING GOOD. | NO | NO | 2023-11-06 | November 2023 | 13:35 | 2023-11-06 13:20 | 2023-11-06 13:35 | 0:15 | COMPLETED | CRITICAL
669 | 2023-11-06 | Nov 2023 | 14:30 | 14:35 | 0.0034722 | TREATMENT UNIT | SKANRAY DENTAL | SKANLECTICA OPAL | HAFIZ | OPD | NOT WORKING | REMOVED THE INTERFACE CABLE, REFIXED. NOW WORKING GOOD. | NO | NO | 2023-11-06 | November 2023 | 14:50 | 2023-11-06 14:30 | 2023-11-06 14:50 | 0:20 | COMPLETED | SEMI-CRITICAL
670 | 2023-11-06 | Nov 2023 | 16:30 | 16:35 | 0.0034722 | ELECTROTHERAPHY WITHOUT LASER | CHATTANOOGA | INTELECT ADVANCED | SANTHOSH | PHYSIOTHERAPY | POWER CORD PORT TORE | REPLACED BY COMPANY PERSON | POWERCORD | NO | 2023-11-06 | November 2023 | 16:40 | 2023-11-06 16:30 | 2023-11-06 16:40 | 0:10 | COMPLETED | SEMI-CRITICAL
671 | 2023-11-07 | Nov 2023 | 0:55 | 1:00 | 0.0034722 | PLASMA STERILIZER | JOHNSON&JOHNSON | STERRAD 100NX | SOORAJ | CSSD | UNOBLE TO CONDITION LOAD ALARM | RESTARTED THE SYSTEM AND RAN CYCLE. FOUND ISSUE WITH CONTROLLER BOARD CAUSING INTERMITTENT FAILURE. AS OF NOW WORKING BY RESTARTING SYSTEM AFTER EACH CYCLE. | NO | NO | 2023-11-07 | November 2023 | 4:30 | 2023-11-07 00:55 | 2023-11-07 04:30 | 3:35 | COMPLETED | SEMI-CRITICAL
672 | 2023-11-07 | Nov 2023 | 1:20 | 1:25 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SOORAJ | WARD | UNOBLE TO DO PREOP CHECK. | PRE OP CHEECK DONE USING TEST PLUGS. DONE USING PADDLES AND FOUND WORKING | NO | NO | 2023-11-07 | November 2023 | 1:30 | 2023-11-07 01:20 | 2023-11-07 01:30 | 0:10 | COMPLETED | CRITICAL
673 | 2023-11-07 | Nov 2023 | 15:10 | 15:15 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | HAFIZ | ICU | O2 CELL NEEDS TO BE REPLACED | REPLACED A NEW O2 CELL AND WORKING GOOD, | NO | NO | 2023-11-07 | November 2023 | 15:30 | 2023-11-07 15:10 | 2023-11-07 15:30 | 0:20 | COMPLETED | CRITICAL
674 | 2023-11-07 | Nov 2023 | 18:50 | 18:53 | 0.0020833 | BIPAP MACHINE | PHILIPS | A40 | SANTHOSH | ICU | VENTILATOR INOPERATIVE | CHECKED TEST RUN ,NO ISSUE FOUND | NO | NO | 2023-11-07 | November 2023 | 19:00 | 2023-11-07 18:50 | 2023-11-07 19:00 | 0:10 | COMPLETED | SEMI-CRITICAL
675 | 2023-11-08 | Nov 2023 | 9:30 | 9:45 | 0.010417 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | JERLIN | ICU | NOT WORKING | FIXED THE TEMP SENSOR PROPERLY AND ISSUE RECTIFIED. | NO | NO | 2023-11-08 | November 2023 | 9:50 | 2023-11-08 09:30 | 2023-11-08 09:50 | 0:20 | COMPLETED | SEMI-CRITICAL
676 | 2023-11-08 | Nov 2023 | 16:00 | 16:10 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | ICU | NOT WORKING | POWER CORD WAS DISCONNECTED FROM ITS SOCKET. FIXED IT AND ISSUE RECTIFIED | NO | NO | 2023-11-08 | November 2023 | 16:20 | 2023-11-08 16:00 | 2023-11-08 16:20 | 0:20 | COMPLETED | NON-CRITICAL
677 | 2023-11-08 | Nov 2023 | 10:00 | 10:05 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | ICU | NOT WORKING | REMOTE WAS LOCKED. UNLOCKED THE REMOTE AND RECTIFIED THE ISSUE | NO | NO | 2023-11-08 | November 2023 | 10:15 | 2023-11-08 10:00 | 2023-11-08 10:15 | 0:15 | COMPLETED | NON-CRITICAL
678 | 2023-11-08 | Nov 2023 | 16:15 | 16:20 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | ICU | PRINTER NOT WORKING | CHECKED AND ITS WORKING GOOD | NO | NO | 2023-11-08 | November 2023 | 16:30 | 2023-11-08 16:15 | 2023-11-08 16:30 | 0:15 | COMPLETED | CRITICAL
679 | 2023-11-08 | Nov 2023 | 15:20 | 15:30 | 0.0069444 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | RAMKIE | ICU | O2 SESNOR FAILURE | REPLACED O2 SESNOR | O2 SENSOR SERVO AIR | NO | 2023-11-08 | November 2023 | 15:40 | 2023-11-08 15:20 | 2023-11-08 15:40 | 0:20 | COMPLETED | CRITICAL
680 | 2023-11-08 | Nov 2023 | 16:10 | 16:30 | 0.013889 | OPTICAL COHERENCE TOMOGRAPHY, 3 DIMENSIONOL | TOPCON | 3D OCT-1 | RAMKIE | OPD | DISTURBANCE IN DISPLAY | CHECKED AND FOUND THE PROBLEM WITH DISPLAY MONITOR, SO NEED TO REPLACE IT | NO | DISPLAY | 2023-11-08 | November 2023 | 17:00 | 2023-11-08 16:10 | 2023-11-08 17:00 | 0:50 | PENDING | SEMI-CRITICAL
681 | 2023-11-09 | Nov 2023 | 10:15 | 10:20 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | THANISH | WARD | DEFIBRILLATOR NOT WORKING | PADDLES ISSUE RECTIFIED AND DONE THE OP CHECK.WORKING FINE. | NO | NO | 2023-11-09 | November 2023 | 10:30 | 2023-11-09 10:15 | 2023-11-09 10:30 | 0:15 | COMPLETED | CRITICAL
682 | 2023-11-09 | Nov 2023 | 12:05 | 12:10 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | THANISH | OPD | ECG NOT WORKING | REPLACED THE 10-LEAD ECG CABLE.WORKING FINE | 10-LEADS ECG CABLE | NO | 2023-11-09 | November 2023 | 12:30 | 2023-11-09 12:05 | 2023-11-09 12:30 | 0:25 | COMPLETED | SEMI-CRITICAL
683 | 2023-11-09 | Nov 2023 | 14:05 | 14:10 | 0.0034722 | LIGHT SOURCE UNIT,XENON | KARL STORZ | 20133720 | THANISH | OT | XENON LAMP FAILED | REPLACED THE XENON LAMP.WORKING FINE. | XENON LAMP | NO | 2023-11-09 | November 2023 | 14:30 | 2023-11-09 14:05 | 2023-11-09 14:30 | 0:25 | COMPLETED | SEMI-CRITICAL
684 | 2023-11-09 | Nov 2023 | 17:00 | 17:10 | 0.0069444 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | JERLIN | OPD | NOT WORKING | CHEANGED THE CABLE AND ISSUE RECTIFIED | ECG CABLE | NO | 2023-11-09 | November 2023 | 17:20 | 2023-11-09 17:00 | 2023-11-09 17:20 | 0:20 | COMPLETED | SEMI-CRITICAL
685 | 2023-11-09 | Nov 2023 | 22:25 | 22:30 | 0.0034722 | VENTILATOR, PEDIATRIC | MINDRAY MEDICAL INDIA | SV 300 | SOORAJ | ICU | AIRWAY OBSTRUCTION ALARM | FOUND ISSUE WITH PATIENT CIRCUIT. PATIENT CIRCUIT CHANGED AND ISSUE RECTIFIED. | NO | NO | 2023-11-09 | November 2023 | 22:45 | 2023-11-09 22:25 | 2023-11-09 22:45 | 0:20 | COMPLETED | CRITICAL
686 | 2023-11-09 | Nov 2023 | 13:10 | 13:15 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | HAFIZ | OT | NOT WORKING | SERVICED AND RETURNED TO ot. NOW WORKING GOOD | NO | NO | 2023-11-09 | November 2023 | 13:40 | 2023-11-09 13:10 | 2023-11-09 13:40 | 0:30 | COMPLETED | SEMI-CRITICAL
687 | 2023-11-09 | Nov 2023 | 10:10 | 10:30 | 0.013889 | BIOSAFETY CABINET | MICROFLOW DEVICES INDIA PVT LTD | MFD BS/B2 1200 | RAMKIE | LABORATORY | Not switching ON | CHECKED AND FOUND THE PROBLEM WITH SOCKET, THEN RECTIFIED | NO | NO | 2023-11-09 | November 2023 | 10:40 | 2023-11-09 10:10 | 2023-11-09 10:40 | 0:30 | COMPLETED | SEMI-CRITICAL
688 | 2023-11-09 | Nov 2023 | 14:50 | 14:55 | 0.0034722 | VENTILATOR, ADULT | MAQUET | SERVO-I | SANTHOSH | ICU | INTERNOL LEAK FAILED | CHECKED WITH ANOTHER CASSETTE PASS PRE-USE CHECK | NO | NO | 2023-11-09 | November 2023 | 15:00 | 2023-11-09 14:50 | 2023-11-09 15:00 | 0:10 | COMPLETED | CRITICAL
689 | 2023-11-10 | Nov 2023 | 10:05 | 10:10 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | THANISH | OT | DISPLAY NOT WORKING | RECTIFIED THE DISPLAY ISSUE.WORKING FINE. | NO | NO | 2023-11-10 | November 2023 | 10:25 | 2023-11-10 10:05 | 2023-11-10 10:25 | 0:20 | COMPLETED | SEMI-CRITICAL
690 | 2023-11-10 | Nov 2023 | 16:00 | 16:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | THANISH | OT | NOT SWITCHING ON | RESTARTED THE MACHINE. WORKING FINE. | NO | NO | 2023-11-10 | November 2023 | 16:20 | 2023-11-10 16:00 | 2023-11-10 16:20 | 0:20 | COMPLETED | CRITICAL
691 | 2023-11-10 | Nov 2023 | 10:40 | 11:00 | 0.013889 | WATERBATH,LABORATORY | VANI SCIENTIFIC & EQUIPMENTS | NA | RAMKIE | LABORATORY | NOT REACHING THE TARGET TEMP | CHANGED THE SET TEMPERATURE | NO | NO | 2023-11-10 | November 2023 | 11:20 | 2023-11-10 10:40 | 2023-11-10 11:20 | 0:40 | COMPLETED | SEMI-CRITICAL
692 | 2023-11-10 | Nov 2023 | 11:00 | 11:05 | 0.0034722 | ORTHOPANTOMOGRAPHY | SKANRAY DENTAL | MYRAY HYPERION X5 | SANTHOSH | OPD | COMMUNICATION ISSUE | CHECKED LAN CABLE AND RESTART, ISSUE RECTIFIED | NO | NO | 2023-11-10 | November 2023 | 11:10 | 2023-11-10 11:00 | 2023-11-10 11:10 | 0:10 | COMPLETED | SEMI-CRITICAL
693 | 2023-11-11 | Nov 2023 | 10:05 | 10:10 | 0.0034722 | DEFIBRILLATOR MONITOR WITH AED | PHILIPS | EFFICIA DFM 100 | THANISH | ICU | PRINTER NOT WORKING | CLEANED THE PRINTER HEAD/GIVEN TEST PRINT.WORKING FINE. | NO | NO | 2023-11-11 | November 2023 | 10:20 | 2023-11-11 10:05 | 2023-11-11 10:20 | 0:15 | COMPLETED | CRITICAL
694 | 2023-11-11 | Nov 2023 | 4:20 | 4:22 | 0.0013889 | BIPAP MACHINE | PHILIPS | A40 | SOORAJ | WARD | DETACHABLE BATTERY ALARM | FOUND BATTERY DEPLETED COMPLETELY. PUT ON MAINS AND ISSUE RECTIFIED. | NO | NO | 2023-11-11 | November 2023 | 4:40 | 2023-11-11 04:20 | 2023-11-11 04:40 | 0:20 | COMPLETED | SEMI-CRITICAL
695 | 2023-11-11 | Nov 2023 | 21:00 | 21:08 | 0.0055556 | PULSE OXIMETER,PORTABLE | COVIDIEN | NELLCOR | HAFIZ | WARD | NOT WORKING | REMOVED THE EXTENSION CABLE, CLEANED AND REFIXED. CHECKED AND NOW WORKING GOOD. | NO | NO | 2023-11-11 | November 2023 | 21:15 | 2023-11-11 21:00 | 2023-11-11 21:15 | 0:15 | COMPLETED | SEMI-CRITICAL
696 | 2023-11-11 | Nov 2023 | 22:15 | 22:18 | 0.0020833 | PULSE OXIMETER,PORTABLE | COVIDIEN | NELLCOR | HAFIZ | WARD | EEE 010 ERROR | BATTERY FULLY DEPLETED, PLUGGED TO AC SOURCE. REMOVED THE EXTENSION CABLE, CLEANED AND REFIXED. CHECKED AND WORKING GOOD. | NO | NO | 2023-11-11 | November 2023 | 22:30 | 2023-11-11 22:15 | 2023-11-11 22:30 | 0:15 | COMPLETED | SEMI-CRITICAL
697 | 2023-11-11 | Nov 2023 | 22:30 | 22:32 | 0.0013889 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | HAFIZ | WARD | BATTERY DEPLETED | GIVEN ONE POWERCHORD AND CONNECTED TO AC SOURCE. NOW WORKING GOOD. | NO | NO | 2023-11-11 | November 2023 | 22:40 | 2023-11-11 22:30 | 2023-11-11 22:40 | 0:10 | COMPLETED | SEMI-CRITICAL
698 | 2023-11-11 | Nov 2023 | 22:50 | 22:55 | 0.0034722 | BIPAP MACHINE | PHILIPS | A40 | HAFIZ | WARD | SEETINGS NEEDS TO RECTIFY | SETTINGS CHANGED AS PER THE REQUIREMENTS. NOW WORKING GOOD. | NO | NO | 2023-11-11 | November 2023 | 23:05 | 2023-11-11 22:50 | 2023-11-11 23:05 | 0:15 | COMPLETED | SEMI-CRITICAL
699 | 2023-11-12 | Nov 2023 | 9:30 | 9:40 | 0.0069444 | PULSE OXIMETER,PORTABLE | COVIDIEN | NELLCOR | JERLIN | WARD | SPO2 PROBE NOT WORKING | CHANGED THE SENSOR AND NOW ITS WORKING GOOD | SPO2 PROBE | NO | 2023-11-12 | November 2023 | 9:55 | 2023-11-12 09:30 | 2023-11-12 09:55 | 0:25 | COMPLETED | SEMI-CRITICAL
700 | 2023-11-12 | Nov 2023 | 15:00 | 15:10 | 0.0069444 | STEAM STERILIZER | GETINGE | GSS67H102E | JERLIN | CSSD | BOWIE DICK TEST FAILED | RESTARTED THE MACHINE AND ISSUE RECTIFIED | NO | NO | 2023-11-12 | November 2023 | 15:30 | 2023-11-12 15:00 | 2023-11-12 15:30 | 0:30 | COMPLETED | SEMI-CRITICAL
701 | 2023-11-12 | Nov 2023 | 15:00 | 15:10 | 0.0069444 | STEAM STERILIZER | GETINGE | GSS67H102E | JERLIN | CSSD | WATER LEVEL ERROR | REFILLED WATER UPTO THE LEVEL AND ISSUE RECTIFIED | NO | NO | 2023-11-12 | November 2023 | 15:20 | 2023-11-12 15:00 | 2023-11-12 15:20 | 0:20 | COMPLETED | SEMI-CRITICAL
702 | 2023-11-12 | Nov 2023 | 14:30 | 14:35 | 0.0034722 | ENDO INSUFFLATOR,50L | KARL STORZ | UI500 | JERLIN | OT | FLOW LEVEL NOT REACHED | CHECKED AND ITS WORKING GOOD | NO | NO | 2023-11-12 | November 2023 | 14:45 | 2023-11-12 14:30 | 2023-11-12 14:45 | 0:15 | COMPLETED | SEMI-CRITICAL
703 | 2023-11-12 | Nov 2023 | 4:30 | 4:35 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 500 | HAFIZ | EMERGENCY | A LINE CABLE NOT WORKING | REMOVED THE CABLE, REFIXED. NOW WORKING GOOD. | NO | NO | 2023-11-12 | November 2023 | 4:45 | 2023-11-12 04:30 | 2023-11-12 04:45 | 0:15 | COMPLETED | SEMI-CRITICAL
704 | 2023-11-12 | Nov 2023 | 4:45 | 4:46 | 0.00069444 | DEFIBRILLATOR MONITOR WITH AED | PHILIPS | EFFICIA DFM 100 | HAFIZ | EMERGENCY | LEADS ECG TEST FAILED | REMOVED THE ECG CABLE, CLEANED AND REFIXED. RERUNNED THE OPERATIONAL CHECK AND PASSED. NOW WORKING GOOD. | NO | NO | 2023-11-12 | November 2023 | 4:55 | 2023-11-12 04:45 | 2023-11-12 04:55 | 0:10 | COMPLETED | CRITICAL
705 | 2023-11-12 | Nov 2023 | 4:55 | 4:56 | 0.00069444 | DEFIBRILLATOR MONITOR WITH AED | PHILIPS | EFFICIA DFM 100 | HAFIZ | EMERGENCY | DATE AND TIME HAS TO CHANGE | RESTORED THE CURRENT DATE AND TIME. AND PASSED THE OPERATIONAL CHECK. | NO | NO | 2023-11-12 | November 2023 | 5:05 | 2023-11-12 04:55 | 2023-11-12 05:05 | 0:10 | COMPLETED | CRITICAL
706 | 2023-11-12 | Nov 2023 | 5:35 | 5:40 | 0.0034722 | DEFIBRILLATOR MONITOR WITH AED | PHILIPS | EFFICIA DFM 100 | HAFIZ | ICU | PRINTER ISSUE | CLEANED THE PRINTER HEAD AND PASSED THE OPERATIONAL CHECK. NOW WORKIING GOOD. | NO | NO | 2023-11-12 | November 2023 | 5:50 | 2023-11-12 05:35 | 2023-11-12 05:50 | 0:15 | COMPLETED | CRITICAL
707 | 2023-11-12 | Nov 2023 | 21:00 | 21:04 | 0.0027778 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | ICU | NOT WORKING | RESET DONE BUT STILL SAME RESULT. UNPLUGGED FROM THE MAIN SOURCEAND KEPT IN BATTERY MODE. AFTER AN HOUR, CHECKED AND WORKING GOOD. | NO | NO | 2023-11-12 | November 2023 | 22:00 | 2023-11-12 21:00 | 2023-11-12 22:00 | 1:00 | COMPLETED | NON-CRITICAL
708 | 2023-11-13 | Nov 2023 | 11:00 | 11:05 | 0.0034722 | VIDEO PROCESSING SYSTEM,ENDOSCOPES | FUJIFILM CORPORATION | VP-3500HD | THANISH | EMERGENCY | DISPLAY NOT WORKING | RECONNECTED THE DISPLAY CABLE.WORKING FINE. | NO | NO | 2023-11-13 | November 2023 | 11:20 | 2023-11-13 11:00 | 2023-11-13 11:20 | 0:20 | COMPLETED | SEMI-CRITICAL
709 | 2023-11-13 | Nov 2023 | 6:40 | 6:45 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | ICU | SIDE REEL NOT WORKING | SPRAYED WD-40 ON SIDE REEL, OBSERVED AND NOW WORKING GOOD. | NO | NO | 2023-11-13 | November 2023 | 6:50 | 2023-11-13 06:40 | 2023-11-13 06:50 | 0:10 | COMPLETED | NON-CRITICAL
710 | 2023-11-13 | Nov 2023 | 6:50 | 6:50 | 0 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | ICU | SIDE REEL NOT WORKING | SPRAYED WD-40 ON SIDE REEL, KEPT FOR OBSERVATION AND NOW WORKING GOOD. | NO | NO | 2023-11-13 | November 2023 | 7:00 | 2023-11-13 06:50 | 2023-11-13 07:00 | 0:10 | COMPLETED | NON-CRITICAL
711 | 2023-11-13 | Nov 2023 | 7:00 | 7:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | RESTARTED THE MACHINE, RERUNED THE TEST AND PASSED. NOW READY TO USE. | NO | NO | 2023-11-13 | November 2023 | 7:20 | 2023-11-13 07:00 | 2023-11-13 07:20 | 0:20 | COMPLETED | CRITICAL
712 | 2023-11-13 | Nov 2023 | 21:20 | 21:25 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | CVP TESTS PENDING | CVP TESTS DONE AND MACHINE NOW READY TO USE. | CARTRIDGE | NO | 2023-11-13 | November 2023 | 21:45 | 2023-11-13 21:20 | 2023-11-13 21:45 | 0:25 | COMPLETED | SEMI-CRITICAL
713 | 2023-11-14 | Nov 2023 | 10:00 | 10:05 | 0.0034722 | OXYGEN THERAPY UNIT | INSPIRED | O2FLO | THANISH | ICU | HFNC NOT WORKING | ADJUSTED THE SETTINGS.WORKING FINE. | NO | NO | 2023-11-14 | November 2023 | 10:20 | 2023-11-14 10:00 | 2023-11-14 10:20 | 0:20 | COMPLETED | SEMI-CRITICAL
714 | 2023-11-14 | Nov 2023 | 11:00 | 11:05 | 0.0034722 | BMI MACHINE | AKAS MEDICAL | QVS-100 | THANISH | OPD | NIBP NOT WORKING | NIBP HOSE LEAKAGE/REPLACED THE NIBP HOSE.WORKING FINE. | NIBP HOSE | NO | 2023-11-14 | November 2023 | 11:30 | 2023-11-14 11:00 | 2023-11-14 11:30 | 0:30 | COMPLETED | NON-CRITICAL
715 | 2023-11-14 | Nov 2023 | 7:30 | 7:40 | 0.0069444 | ACT MACHINE | HELENO LABORATORIES | MINI II | JERLIN | ICU | S1 ERROR | CHANGED THE CONTAINER AND RERAN THE SAMPLE ND FOUND THAT IT'S WORKING GOOD. | NO | NO | 2023-11-14 | November 2023 | 7:50 | 2023-11-14 07:30 | 2023-11-14 07:50 | 0:20 | COMPLETED | SEMI-CRITICAL
716 | 2023-11-14 | Nov 2023 | 16:30 | 16:40 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | ICU | NOT WORKING | RESET DONE. ISSUE RECTIFIED | NO | NO | 2023-11-14 | November 2023 | 16:55 | 2023-11-14 16:30 | 2023-11-14 16:55 | 0:25 | COMPLETED | NON-CRITICAL
717 | 2023-11-14 | Nov 2023 | 0:05 | 0:10 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | HAFIZ | WARD | NOT WORKING | FOUND THAT BUTTONS ARE LOCKED. UNLOCKED THE BUTTONS, NOW WORKING GOOD. | NO | NO | 2023-11-14 | November 2023 | 0:15 | 2023-11-14 00:05 | 2023-11-14 00:15 | 0:10 | COMPLETED | NON-CRITICAL
718 | 2023-11-14 | Nov 2023 | 6:50 | 6:55 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | WARD | PRINTER ISSUE | FOUND ISSUE IN THE PRINTER PAPER AND REPLACED. | PAPER ROLL | NO | 2023-11-14 | November 2023 | 7:10 | 2023-11-14 06:50 | 2023-11-14 07:10 | 0:20 | COMPLETED | CRITICAL
719 | 2023-11-14 | Nov 2023 | 12:10 | 12:30 | 0.013889 | TRACTION WITH TREATMENT TABLE | CHATTANOOGA | TRITON DTS | RAMKIE | PHYSIOTHERAPY | NOT SWITCHING ON | CHECKED AND RECTIFIED, UNDER OBSERVATION | NO | NO | 2023-11-14 | November 2023 | 12:40 | 2023-11-14 12:10 | 2023-11-14 12:40 | 0:30 | COMPLETED | SEMI-CRITICAL
720 | 2023-11-14 | Nov 2023 | 14:20 | 14:25 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 9000 | SANTHOSH | ICU | NOT WORKING | CHECKED AND POWERCORD RESET | NO | NO | 2023-11-14 | November 2023 | 14:30 | 2023-11-14 14:20 | 2023-11-14 14:30 | 0:10 | COMPLETED | NON-CRITICAL
721 | 2023-11-15 | Nov 2023 | 10:15 | 10:20 | 0.0034722 | VIDEO PROCESSING SYSTEM,ENDOSCOPES | KARL STORZ | TH102 | THANISH | OPD | DISPALY ISSUE | RECTIFIED THE VIDEO PROCESSING DISPAY CABLE.WORKING FINE. | NO | NO | 2023-11-15 | November 2023 | 10:45 | 2023-11-15 10:15 | 2023-11-15 10:45 | 0:30 | COMPLETED | SEMI-CRITICAL
722 | 2023-11-15 | Nov 2023 | 14:00 | 14:05 | 0.0034722 | MONITOR HIGH DEFINITION | KARL STORZ | TM220 | JERLIN | OT | ERROR MESSAGE DISPLAYED | RESTARTED THE MACHINE AND ISSUE RECTIFIED | NO | NO | 2023-11-15 | November 2023 | 14:10 | 2023-11-15 14:00 | 2023-11-15 14:10 | 0:10 | COMPLETED | SEMI-CRITICAL
723 | 2023-11-15 | Nov 2023 | 7:05 | 7:05 | 0 | FLUID WARMER SYSTEM | MATRIX | MN05 | JERLIN | OT | NOT WORKING | CHECKED AND ITS WORKING GOOD | NO | NO | 2023-11-15 | November 2023 | 7:10 | 2023-11-15 07:05 | 2023-11-15 07:10 | 0:05 | COMPLETED | SEMI-CRITICAL
724 | 2023-11-15 | Nov 2023 | 14:30 | 14:40 | 0.0069444 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | SOORAJ | OT | PAEDIATRIC TEE PROBE NOT WORKING | PROBE CONNECTOR ON DISPLAY ISSUE. CHANGED NEW CONNECTOR AND CHECKED, FOUND WORKING FINE. | PROBE CONNECTOR | NO | 2023-11-16 | November 2023 | 20:30 | 2023-11-15 14:30 | 2023-11-16 20:30 | 30:00 | COMPLETED | SEMI-CRITICAL
725 | 2023-11-15 | Nov 2023 | 23:00 | 23:05 | 0.0034722 | TRANSPORT VENTILATOR | RESMED | ASTRAL 100 | HAFIZ | WARD | UNOBLE TO CHANGE THE SETTINGS | FOUND THAT MACHINE SCREEN IS LOCKED. UNLOCKED THE SCREEN, SETTINGS CHANGED AS PER THE REQUIREMENTS. NOW WORKING GOOD. | NO | NO | 2023-11-15 | November 2023 | 23:15 | 2023-11-15 23:00 | 2023-11-15 23:15 | 0:15 | COMPLETED | SEMI-CRITICAL
726 | 2023-11-15 | Nov 2023 | 22:00 | 22:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | ICU | PRINTER ISSUE | CHECKED AND FOUND THAT PRINTER PAPER ALMOST FINISHED. PROPERLY FIXED THE PRINTER PAPER. OPERATIONAL CHECK DONE AND NOW PRINT TAKEN. WORKING GOOD. | NO | NO | 2023-11-15 | November 2023 | 22:30 | 2023-11-15 22:00 | 2023-11-15 22:30 | 0:30 | COMPLETED | CRITICAL
727 | 2023-11-15 | Nov 2023 | 12:40 | 12:50 | 0.0069444 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | RAMKIE | OT | NOT WORKING | REPLACED TEMP HOSE AND RECTIFIED | WARMER HOSE | NO | 2023-11-15 | November 2023 | 13:00 | 2023-11-15 12:40 | 2023-11-15 13:00 | 0:20 | COMPLETED | SEMI-CRITICAL
728 | 2023-11-15 | Nov 2023 | 13:40 | 13:45 | 0.0034722 | BMI MACHINE | AKAS MEDICAL | QVS-100 | SANTHOSH | OPD | BP READING ERROR | REPLACED NEW BP CUFF | BP CUFF | NO | 2023-11-15 | November 2023 | 13:50 | 2023-11-15 13:40 | 2023-11-15 13:50 | 0:10 | COMPLETED | NON-CRITICAL
729 | 2023-11-16 | Nov 2023 | 11:10 | 11:15 | 0.0034722 | PULMONORY FUNCTION TESTING MACHINE | NDD | EASY ONE PRO | THANISH | PHYSIOTHERAPY | SPIROMETRY NOT WORKING | RECTIFIED THE SPIROMETRY GUN.WORKING FINE. | NO | NO | 2023-11-16 | November 2023 | 11:30 | 2023-11-16 11:10 | 2023-11-16 11:30 | 0:20 | COMPLETED | SEMI-CRITICAL
730 | 2023-11-16 | Nov 2023 | 7:30 | 7:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | ICU | NOT WORKING | POWER CORD WAS DISCONNECTED FROM ITS SOCKET. FIXED IT AND ISSUE RECTIFIED | NO | NO | 2023-11-16 | November 2023 | 7:40 | 2023-11-16 07:30 | 2023-11-16 07:40 | 0:10 | COMPLETED | NON-CRITICAL
731 | 2023-11-16 | Nov 2023 | 8:00 | 8:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | OT | PRINTER NOT WORKING | FIXED THE DOOR PROPERLY AND ISSUE RECTIFIED | NO | NO | 2023-11-16 | November 2023 | 8:20 | 2023-11-16 08:00 | 2023-11-16 08:20 | 0:20 | COMPLETED | CRITICAL
732 | 2023-11-16 | Nov 2023 | 8:35 | 8:40 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | DIALYSIS | PRINTER NOT WORKING | CHECKED AND ITS WORKING GOOD | NO | NO | 2023-11-16 | November 2023 | 8:45 | 2023-11-16 08:35 | 2023-11-16 08:45 | 0:10 | COMPLETED | CRITICAL
733 | 2023-11-16 | Nov 2023 | 10:15 | 10:20 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | OT | DATE AND TIME CHANGED | CHANGED THE SETTINGS AND ISSUE RECTIFIED | NO | NO | 2023-11-16 | November 2023 | 10:30 | 2023-11-16 10:15 | 2023-11-16 10:30 | 0:15 | COMPLETED | CRITICAL
734 | 2023-11-16 | Nov 2023 | 14:00 | 14:10 | 0.0069444 | REFRIGERATOR, LOW TEMPERATURE | REMI | BR-50 | JERLIN | OT | TEMP LEVEL EXCEEDED | RESTARTED THE MACHINE AND ISSUE RECTIFIED | NO | NO | 2023-11-16 | November 2023 | 14:20 | 2023-11-16 14:00 | 2023-11-16 14:20 | 0:20 | COMPLETED | SEMI-CRITICAL
735 | 2023-11-16 | Nov 2023 | 16:00 | 16:10 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 550 | JERLIN | OT | GAS MODULE MALFUNCTION | CHECKED AND ITS WORKING GOOD | NO | NO | 2023-11-16 | November 2023 | 16:15 | 2023-11-16 16:00 | 2023-11-16 16:15 | 0:15 | COMPLETED | SEMI-CRITICAL
736 | 2023-11-17 | Nov 2023 | 9:00 | 9:05 | 0.0034722 | MONITOR HIGH DEFINITION | KARL STORZ | TM220 | JERLIN | OT | DISPLAY NOT WORKING | FIXED THE CAMERA CABLE PROPERLY AND ISSUE RECTIFIED | NO | NO | 2023-11-17 | November 2023 | 9:10 | 2023-11-17 09:00 | 2023-11-17 09:10 | 0:10 | COMPLETED | SEMI-CRITICAL
737 | 2023-11-17 | Nov 2023 | 15:00 | 15:00 | 0 | OT TABLE | MAQUET | ALPHA CLASSIC PRO | JERLIN | OT | REMOTE NOT WORKING | CHECKED AND ITS WORKING GOOD | NO | NO | 2023-11-17 | November 2023 | 15:05 | 2023-11-17 15:00 | 2023-11-17 15:05 | 0:05 | COMPLETED | NON-CRITICAL
738 | 2023-11-17 | Nov 2023 | 14:30 | 14:32 | 0.0013889 | MONITOR HIGH DEFINITION | KARL STORZ | TM220 | JERLIN | OT | MACHINE SWITCHING OFF AUTOMATICALLY | CONNNECTED THE POWERCORD PROPERLY AND RECTIFIED THE ISSUE | NO | NO | 2023-11-17 | November 2023 | 14:35 | 2023-11-17 14:30 | 2023-11-17 14:35 | 0:05 | COMPLETED | SEMI-CRITICAL
739 | 2023-11-17 | Nov 2023 | 6:55 | 7:00 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | RERUNNED THE CHECKOUT, ISSUE RECTIFIED. | NO | NO | 2023-11-17 | November 2023 | 7:20 | 2023-11-17 06:55 | 2023-11-17 07:20 | 0:25 | COMPLETED | CRITICAL
740 | 2023-11-17 | Nov 2023 | 21:45 | 21:55 | 0.0069444 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | WARD | NOT WORKING | IV LINE NOT PROPERLY FIXED. FXED IT PROPERLY, CHECKED AND NOW WORKING GOOD. | NO | NO | 2023-11-17 | November 2023 | 22:10 | 2023-11-17 21:45 | 2023-11-17 22:10 | 0:25 | COMPLETED | NON-CRITICAL
741 | 2023-11-17 | Nov 2023 | 20:45 | 20:50 | 0.0034722 | VENTILATOR, PEDIATRIC | MINDRAY MEDICAL INDIA | SV 300 | SANTHOSH | ICU | LIMB AIRWAY OBSTRUCT | REMOVED OBSTRUCTS OPENED MASK VALVE | NO | NO | 2023-11-17 | November 2023 | 21:00 | 2023-11-17 20:45 | 2023-11-17 21:00 | 0:15 | COMPLETED | CRITICAL
742 | 2023-11-17 | Nov 2023 | 21:00 | 21:10 | 0.0069444 | IABP MACHINE | MAQUET | CS300 | SANTHOSH | ICU | DDNT GET ECG WAVE | REFIXED ECG CABLE AND CLEANED ITS CONNECTOR END | NO | NO | 2023-11-17 | November 2023 | 21:10 | 2023-11-17 21:00 | 2023-11-17 21:10 | 0:10 | COMPLETED | CRITICAL
743 | 2023-11-18 | Nov 2023 | 12:05 | 12:10 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | THANISH | ICU | ABG CATRIDGE EXPIRED | REPLACED THE CATRIDGE AND DONE CVP.WORKING FINE. | NO | NO | 2023-11-18 | November 2023 | 12:45 | 2023-11-18 12:05 | 2023-11-18 12:45 | 0:40 | COMPLETED | SEMI-CRITICAL
744 | 2023-11-18 | Nov 2023 | 13:00 | 13:05 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | THANISH | ICU | ABG CATRIDGE EXPIRED | REPLACED THE CATRIDGE AND DONE CVP.WORKING FINE. | NO | NO | 2023-11-18 | November 2023 | 14:00 | 2023-11-18 13:00 | 2023-11-18 14:00 | 1:00 | COMPLETED | SEMI-CRITICAL
745 | 2023-11-18 | Nov 2023 | 10:30 | 10:35 | 0.0034722 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | JERLIN | OPD | ARTIFACTS IN ECG WAVES | CHANGED THE CABLE AND ISSUE RECTIFIED | 10-LEADS ECG CABLE | NO | 2023-11-18 | November 2023 | 10:50 | 2023-11-18 10:30 | 2023-11-18 10:50 | 0:20 | COMPLETED | NON-CRITICAL
746 | 2023-11-18 | Nov 2023 | 5:50 | 6:00 | 0.0069444 | VENTILATOR, ADULT | MAQUET | SERVO-I | HAFIZ | ICU | NOT TURNING ON | CHECKED AND FOUND ISSUE IN THE POWERSOCKET. CHANGED THE SOCKET AND ISSUE RECTIFIED. PASSED THE PRE-USE CHECK. MACHINE NOW READY TO USE. | NO | NO | 2023-11-18 | November 2023 | 6:20 | 2023-11-18 05:50 | 2023-11-18 06:20 | 0:30 | COMPLETED | CRITICAL
747 | 2023-11-18 | Nov 2023 | 13:20 | 13:30 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | OT | NOT WORKING | POWER CORD RE CONNECTED AND RECTIFIED | NO | NO | 2023-11-18 | November 2023 | 13:30 | 2023-11-18 13:20 | 2023-11-18 13:30 | 0:10 | COMPLETED | NON-CRITICAL
748 | 2023-11-18 | Nov 2023 | 13:10 | 13:15 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SANTHOSH | WARD | DEFIB PRINTER NOT WORKING | CHECKED WITH ANOTHER DEFIB PAPER WORKING GOOD TOLD TO REPLACE PAPER | NO | NO | 2023-11-18 | November 2023 | 13:25 | 2023-11-18 13:10 | 2023-11-18 13:25 | 0:15 | COMPLETED | CRITICAL
749 | 2023-11-18 | Nov 2023 | 19:10 | 19:15 | 0.0034722 | HOLTER,RECORDER | WIPRO GE HEALTH CARE | SEER 1000 | SANTHOSH | OPD | RED LIGHT INDICATES | RESET DONE ,CHECKED WITH PATIENT,WORKING GOOD , | NO | NO | 2023-11-18 | November 2023 | 19:20 | 2023-11-18 19:10 | 2023-11-18 19:20 | 0:10 | COMPLETED | SEMI-CRITICAL
750 | 2023-11-19 | Nov 2023 | 12:20 | 12:30 | 0.0069444 | ACT MACHINE | HELENO LABORATORIES | ACTALYKE | RAMKIE | OT | INTERMITTENTLY SHOWING ERROR | CHECKEC AND COMPANY PERSON CALIBRATED AND RETURNED TO USER, NOW IT WORKING | NO | NO | 2023-11-19 | November 2023 | 12:50 | 2023-11-19 12:20 | 2023-11-19 12:50 | 0:30 | COMPLETED | SEMI-CRITICAL
751 | 2023-11-19 | Nov 2023 | 11:30 | 11:35 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | CATRIDGE CVP FAILED | RERUN CVP1, PASSED ALL ,ISSUE RECTIFIED | NO | NO | 2023-11-19 | November 2023 | 11:40 | 2023-11-19 11:30 | 2023-11-19 11:40 | 0:10 | COMPLETED | SEMI-CRITICAL
752 | 2023-11-19 | Nov 2023 | 13:00 | 13:05 | 0.0034722 | IABP MACHINE | MAQUET | CS300 | SANTHOSH | OPD | GET UNWANTED NOISE | CHECKED AND OBSERVED ,ITS AN NORMAL NOISE | NO | NO | 2023-11-19 | November 2023 | 13:10 | 2023-11-19 13:00 | 2023-11-19 13:10 | 0:10 | COMPLETED | CRITICAL
753 | 2023-11-19 | Nov 2023 | 15:20 | 15:25 | 0.0034722 | NITRIC OXIDE DELIVERY SYSTEM | WIPRO GE HEALTH CARE | SLE | SANTHOSH | ICU | LIMIT ALARM ENOBLED | CHECKED AND CHANGED THE LIMIT ALARM,MACHINE WORKING GOOD | NO | NO | 2023-11-19 | November 2023 | 15:30 | 2023-11-19 15:20 | 2023-11-19 15:30 | 0:10 | COMPLETED | CRITICAL
754 | 2023-11-19 | Nov 2023 | 17:10 | 17:15 | 0.0034722 | HOLTER,RECORDER | WIPRO GE HEALTH CARE | SEER 1000 | SANTHOSH | OPD | UNOBLE TO RESET | CLEANED COMMUNICATION CABLE,REPLACED BATTERY,RESET DONE | NO | NO | 2023-11-19 | November 2023 | 17:20 | 2023-11-19 17:10 | 2023-11-19 17:20 | 0:10 | COMPLETED | SEMI-CRITICAL
755 | 2023-11-20 | Nov 2023 | 22:00 | 22:05 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | THANISH | WARD | ECG NOT WORKING | RESETTED THE SETTINGS.WORKING FINE. | NO | NO | 2023-11-20 | November 2023 | 22:30 | 2023-11-20 22:00 | 2023-11-20 22:30 | 0:30 | COMPLETED | SEMI-CRITICAL
756 | 2023-11-20 | Nov 2023 | 13:20 | 13:30 | 0.0069444 | ELISA MICROPLATE READER | TRANSASIA | LISA SCAN EM | RAMKIE | LABORATORY | NOT WORKING IN CERTAIN TEST | CHECKED AND CHANGED THE PROGRAM AND RECTIFIED | NO | NO | 2023-11-20 | November 2023 | 13:50 | 2023-11-20 13:20 | 2023-11-20 13:50 | 0:30 | COMPLETED | SEMI-CRITICAL
757 | 2023-11-20 | Nov 2023 | 8:30 | 8:35 | 0.0034722 | BMI MACHINE | AKAS MEDICAL | QVS-100 | SANTHOSH | OPD | CATRIDGE REAGENT GOT LEAKED | REPLACED ANOTHER CATRIDGE | CARTRIDGE | NO | 2023-11-20 | November 2023 | 8:35 | 2023-11-20 08:30 | 2023-11-20 08:35 | 0:05 | COMPLETED | NON-CRITICAL
758 | 2023-11-21 | Nov 2023 | 11:00 | 11:10 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | JERLIN | RADIOLOGY | CURSOR NOT WORKING PROPERLY | CLEANED THE TRACKER BALL AND ISSUE RECTIFIED | NO | NO | 2023-11-21 | November 2023 | 11:20 | 2023-11-21 11:00 | 2023-11-21 11:20 | 0:20 | COMPLETED | SEMI-CRITICAL
759 | 2023-11-21 | Nov 2023 | 10:00 | 10:10 | 0.0069444 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | JERLIN | OPD | IP CONFIGURATION ERROR | LAN CABLE WAS NOT PROPERLY FIXED . FIXED IT AND RECTIFIED THE ISUE | NO | NO | 2023-11-21 | November 2023 | 10:20 | 2023-11-21 10:00 | 2023-11-21 10:20 | 0:20 | COMPLETED | SEMI-CRITICAL
760 | 2023-11-21 | Nov 2023 | 8:50 | 8:53 | 0.0020833 | NITRIC OXIDE DELIVERY SYSTEM | WIPRO GE HEALTH CARE | SLE | SANTHOSH | ICU | PATIENT OUT PORT GOT BROKEN | FIXED O2 LOW FLOW ADAPTER AS PATIENT OUT CONNECTOR ISSUE RECTIFIED | NO | NO | 2023-11-21 | November 2023 | 9:00 | 2023-11-21 08:50 | 2023-11-21 09:00 | 0:10 | COMPLETED | CRITICAL
761 | 2023-11-21 | Nov 2023 | 11:10 | 11:15 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | SANTHOSH | ICU | O2 CELL FAIL IN PRE-USE CHECK | TEST RUN WITH TESTLUNG ,FOUND FIO2 WITH GOOD VALUES , MACHINE KEPT UNDER OBSERVATION | NO | NO | 2023-11-21 | November 2023 | 11:20 | 2023-11-21 11:10 | 2023-11-21 11:20 | 0:10 | COMPLETED | CRITICAL
762 | 2023-11-21 | Nov 2023 | 12:30 | 12:35 | 0.0034722 | OT TABLE | MAQUET | MEERA | SANTHOSH | OT | NOT WORKING WITH SERVICE MESSAGE | CLEANED AND REFIXED NOW WORKING ,ISSUE RECTIFIED | NO | NO | 2023-11-21 | November 2023 | 12:40 | 2023-11-21 12:30 | 2023-11-21 12:40 | 0:10 | COMPLETED | NON-CRITICAL
763 | 2023-11-22 | Nov 2023 | 13:00 | 13:05 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | OT | NOT WORKING | POWER CORD WAS DISCONNECTED FROM ITS SOCKET . FIXED IT AND ISSUE RECTIFIED | NO | NO | 2023-11-22 | November 2023 | 13:10 | 2023-11-22 13:00 | 2023-11-22 13:10 | 0:10 | COMPLETED | NON-CRITICAL
764 | 2023-11-22 | Nov 2023 | 15:00 | 15:05 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | JERLIN | OPD | NOT ABLE TO TRANSFER PATIENT DATA | PATIENT DATA TRANSFERRED AND PROBLEM RECTIFIED | NO | NO | 2023-11-22 | November 2023 | 15:15 | 2023-11-22 15:00 | 2023-11-22 15:15 | 0:15 | COMPLETED | SEMI-CRITICAL
765 | 2023-11-22 | Nov 2023 | 9:00 | 9:05 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | HAFIZ | ICU | NEED TO REPLACE A CARTRIDGE | REMOVED THE OLD CARTRIDGE (S.NO:301684863) AND REPLACED A NEW CARTRIDGE (S.NO:301709764). WARMING UP DONE AND READY TO USE. | CARTRIDGE | NO | 2023-11-22 | November 2023 | 9:45 | 2023-11-22 09:00 | 2023-11-22 09:45 | 0:45 | COMPLETED | SEMI-CRITICAL
766 | 2023-11-22 | Nov 2023 | 11:20 | 11:25 | 0.0034722 | HEATER COOLER | GETINGE | HCU40 | SANTHOSH | OT | BLANKET GOT WATER LEAK | GLUED PROPERLY FOUND NO WATER LEAK,ISSUE RECTIFIED | NO | NO | 2023-11-22 | November 2023 | 11:30 | 2023-11-22 11:20 | 2023-11-22 11:30 | 0:10 | COMPLETED | SEMI-CRITICAL
767 | 2023-11-22 | Nov 2023 | 17:30 | 17:35 | 0.0034722 | PULMONORY FUNCTION TESTING MACHINE | NDD | EASY ONE PRO | SANTHOSH | PHYSIOTHERAPY | FOUND PHYSICAL DAMAGE IN BLOWING PART CABLE | CABLE DISCONNECTED AND SOLDERED WELL NOW RUN A TEST , ISSUE RECTIFIED | NO | NO | 2023-11-22 | November 2023 | 17:45 | 2023-11-22 17:30 | 2023-11-22 17:45 | 0:15 | COMPLETED | SEMI-CRITICAL
768 | 2023-11-23 | Nov 2023 | 6:00 | 6:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | THANISH | OT | NIBP NOT WORKING | REPLACED THE NIBP CUFF CONNECTOR | BP CONNECTOR | NO | 2023-11-23 | November 2023 | 6:15 | 2023-11-23 06:00 | 2023-11-23 06:15 | 0:15 | COMPLETED | SEMI-CRITICAL
769 | 2023-11-23 | Nov 2023 | 10:20 | 10:30 | 0.0069444 | COBLATOR | ARTHO CARE CORPORATION | COBLATOR II | THANISH | OT | COBLATOR FOOT SWITCH NOT WORKING | REPAIRED THE COBLATOR FOOTSWITCH.WORKING FINE. | NO | NO | 2023-11-23 | November 2023 | 22:50 | 2023-11-23 10:20 | 2023-11-23 22:50 | 12:30 | COMPLETED | SEMI-CRITICAL
770 | 2023-11-23 | Nov 2023 | 11:45 | 11:50 | 0.0034722 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | JERLIN | OPD | ARTIFACTS IN ECG WAVES | FIXED THE CABLE AND POWER CORD PROPERLY AND ISSUE RECTIFIED | NO | NO | 2023-11-23 | November 2023 | 12:15 | 2023-11-23 11:45 | 2023-11-23 12:15 | 0:30 | COMPLETED | NON-CRITICAL
771 | 2023-11-23 | Nov 2023 | 11:20 | 11:30 | 0.0069444 | FREEZER,DEEP FREEZING -80 | THERMO FISHER SCIENTIFIC | FORMA 900 SERIES | RAMKIE | BLOODBANK | NOT ATTAINS THE TARGET TEMP | CHECKED AND SUGGESTED TO DEFROST AND RECTIFIED | NO | NO | 2023-11-23 | November 2023 | 11:50 | 2023-11-23 11:20 | 2023-11-23 11:50 | 0:30 | COMPLETED | SEMI-CRITICAL
772 | 2023-11-24 | Nov 2023 | 11:40 | 11:45 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | SOORAJ | ICU | SYSTEM VOLTAGE FAULT MESSAGE | COMPANY PERSON CHECKED AND FOUND ISSUE WITH POWER SUPPLY BOARD | NO | POWER SUPPLY BOARD | 2023-11-24 | November 2023 | 17:30 | 2023-11-24 11:40 | 2023-11-24 17:30 | 5:50 | PENDING | SEMI-CRITICAL
773 | 2023-11-24 | Nov 2023 | 17:30 | 17:35 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SOORAJ | ICU | ECG TEST FAILED | CLEANED CONNECTORS AND REINSERTED. RAN PREOP CHECK AND TEST PASSED. ISSUE RECTIFIED | NO | NO | 2023-11-24 | November 2023 | 17:45 | 2023-11-24 17:30 | 2023-11-24 17:45 | 0:15 | COMPLETED | CRITICAL
774 | 2023-11-24 | Nov 2023 | 15:20 | 15:30 | 0.0069444 | DIALYSIS MACHINE | FRESENIUS | 4008S NG | RAMKIE | DIALYSIS | BLD ALARM | BLD CALIBRATED AND TEST PASSED | NO | NO | 2023-11-24 | November 2023 | 15:40 | 2023-11-24 15:20 | 2023-11-24 15:40 | 0:20 | COMPLETED | CRITICAL
775 | 2023-11-25 | Nov 2023 | 15:00 | 15:05 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | THANISH | OPD | SOFTWARE ERROR | CLEANED THE KEYBOARD AND RESSTED THE CABLES.WORKING FINE. | NO | NO | 2023-11-25 | November 2023 | 15:40 | 2023-11-25 15:00 | 2023-11-25 15:40 | 0:40 | COMPLETED | SEMI-CRITICAL
776 | 2023-11-25 | Nov 2023 | 16:00 | 16:05 | 0.0034722 | CT MACHINE | PHILIPS | INGENUITY | THANISH | RADIOLOGY | TABLE MOVEMENT NOT WORKING | REPLACED THE RELATIVE ENCODER/PERFORMED THE TEST SCAN.WORKING FINE. | ENCODER | NO | 2023-11-25 | November 2023 | 18:00 | 2023-11-25 16:00 | 2023-11-25 18:00 | 2:00 | COMPLETED | CRITICAL
777 | 2023-11-25 | Nov 2023 | 9:30 | 9:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 9000 | JERLIN | WARD | NOT WORKING | REMOTE WAS LOCKED. UNLOCKED THE REMOTE AND RECTIFIED THE ISSUE | NO | NO | 2023-11-25 | November 2023 | 9:40 | 2023-11-25 09:30 | 2023-11-25 09:40 | 0:10 | COMPLETED | NON-CRITICAL
778 | 2023-11-25 | Nov 2023 | 7:10 | 7:15 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | JERLIN | OT | CIRCUIT LEAK TEST FAILED | CHECKOUT DONE AND ISSUE RECTIFIED | NO | NO | 2023-11-25 | November 2023 | 7:25 | 2023-11-25 07:10 | 2023-11-25 07:25 | 0:15 | COMPLETED | CRITICAL
779 | 2023-11-25 | Nov 2023 | 2:10 | 2:15 | 0.0034722 | FREEZER,DEEP FREEZING -80 | THERMO FISHER SCIENTIFIC | FORMA 900 SERIES | SANTHOSH | BLOODBANK | GOT HIGH TEMPERATURE ALARM | RESET DONE ,WORKING GOOD,KEPT UNDER OBSERVATION | NO | NO | 2023-11-25 | November 2023 | 2:20 | 2023-11-25 02:10 | 2023-11-25 02:20 | 0:10 | COMPLETED | SEMI-CRITICAL
780 | 2023-11-25 | Nov 2023 | 2:30 | 2:35 | 0.0034722 | VENTILATOR, PEDIATRIC | MINDRAY MEDICAL INDIA | SV 300 | SANTHOSH | ICU | O2 CELL FAIL IN PRE-USE CHECK | O2 CAL FAIL ,SO CURRENTLY TURN OFF THE O2 CELL | NO | O2 CELL | 2023-11-25 | November 2023 | 2:35 | 2023-11-25 02:30 | 2023-11-25 02:35 | 0:05 | PENDING | CRITICAL
781 | 2023-11-25 | Nov 2023 | 5:10 | 5:15 | 0.0034722 | VENTILATOR, PEDIATRIC | MINDRAY MEDICAL INDIA | SV 300 | SANTHOSH | EMERGENCY | LOW MINUTE VOLUME ALRM | FOUND ALARM LIMITS ARE IN PEDIATRIC VALUES,THEN CHANGED TO ADULT VALUES | NO | NO | 2023-11-25 | November 2023 | 5:20 | 2023-11-25 05:10 | 2023-11-25 05:20 | 0:10 | COMPLETED | CRITICAL
782 | 2023-11-25 | Nov 2023 | 5:50 | 5:55 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SANTHOSH | ICU | COT NOT WORKING | FOUND POWER CORD DISCONNECTED FROM CONTROL BOX,FIXED PROPERLY | NO | NO | 2023-11-25 | November 2023 | 6:00 | 2023-11-25 05:50 | 2023-11-25 06:00 | 0:10 | COMPLETED | NON-CRITICAL
783 | 2023-11-26 | Nov 2023 | 9:20 | 9:22 | 0.0013889 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | OT | THERAPY DISABLED, RUN OP CHECK | REFIXED THE KNOB, RERUNNED THE OP CHECK AND PASSED. NOW MACHINE READY TO USE. | NO | NO | 2023-11-26 | November 2023 | 9:40 | 2023-11-26 09:20 | 2023-11-26 09:40 | 0:20 | COMPLETED | CRITICAL
784 | 2023-11-26 | Nov 2023 | 11:50 | 11:55 | 0.0034722 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | SANTHOSH | OPD | GETTING ARTIFACTS IN ECG WAVE | CABLES CHECKED AND FIXED TIGHTLY,MACHINE KEPT OBSERVATION | NO | NO | 2023-11-26 | November 2023 | 12:00 | 2023-11-26 11:50 | 2023-11-26 12:00 | 0:10 | COMPLETED | NON-CRITICAL
785 | 2023-11-26 | Nov 2023 | 13:10 | 13:15 | 0.0034722 | CAMERA SYSTEM | KARL STORZ | TC200 | SANTHOSH | OPD | DDNT SAVE PATIENT DATA | CHECKED CABLES AND RESTARTED SYSTEM NOW RECTIFIED | NO | NO | 2023-11-26 | November 2023 | 13:20 | 2023-11-26 13:10 | 2023-11-26 13:20 | 0:10 | COMPLETED | SEMI-CRITICAL
786 | 2023-11-27 | Nov 2023 | 10:00 | 10:05 | 0.0034722 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | THANISH | ICU | SYRINGE PUMP NOT WORKING | PLUNGER KNOB RECTIFIED.WORKING FINE. | NO | NO | 2023-11-27 | November 2023 | 10:30 | 2023-11-27 10:00 | 2023-11-27 10:30 | 0:30 | COMPLETED | NON-CRITICAL
787 | 2023-11-27 | Nov 2023 | 10:30 | 10:32 | 0.0013889 | STEAM STERILIZER | GETINGE | GSS67H102E | SOORAJ | CSSD | GENERATOR CONTACTOR ALARM | FOUND ONE OF THE GENERATOR CONTACTOR IN CONTROL PANEL DEFECTIVE. SWAPPED THE CONTACTOR TEMPORARILY AND RAN CYCLES. MACHINE OBSERVED WORKING FINE. | NO | NO | 2023-11-28 | November 2023 | 17:00 | 2023-11-27 10:30 | 2023-11-28 17:00 | 30:30 | COMPLETED | SEMI-CRITICAL
788 | 2023-11-27 | Nov 2023 | 10:40 | 10:40 | 0 | TABLE TOP STEAM STERILIZER | GETINGE | HS 22K7 | SOORAJ | CSSD | PROGRAM NOT RUINNING | FOUND LOW WATER LEVEL. CHANGED WATER AND FILLED UPTO MAX LEVEL. RAN PROGRAM AND FOUND WORKING. | NO | NO | 2023-11-27 | November 2023 | 10:50 | 2023-11-27 10:40 | 2023-11-27 10:50 | 0:10 | COMPLETED | SEMI-CRITICAL
789 | 2023-11-27 | Nov 2023 | 12:20 | 12:25 | 0.0034722 | OT TABLE | MAQUET | ALPHA CLASSIC PRO | SOORAJ | OT | 0 LEVEL FUNCTION OT WORKING. | CHECKED CONNECTORS AND PORTS, FOUND NO ISSUE WITH REMOTE AND PORTS. SENSOR CALIBRATION REQUIRED. OTHER FUNCTIONS WORKING WELL. AWAITING ENGINEERS VISIT FOR SENSOR CALIBRATION. | NO | NO | 2023-11-28 | November 2023 | 12:40 | 2023-11-27 12:20 | 2023-11-28 12:40 | 24:20 | COMPLETED | NON-CRITICAL
790 | 2023-11-27 | Nov 2023 | 12:45 | 12:45 | 0 | PATIENT MONITOR | PHILIPS | MX 550 | SOORAJ | OT | BATTERY NOT CHARGING | FOUND BATTERY TO BE DEFECIVE AND WORKS ON 25% BACK UP ONLY. BATTERY NEEDS REPLACEMENT | NO | BATTERY | 2023-11-27 | November 2023 | 12:55 | 2023-11-27 12:45 | 2023-11-27 12:55 | 0:10 | PENDING | SEMI-CRITICAL
791 | 2023-11-27 | Nov 2023 | 10:00 | 10:05 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | WARD | NOT WORKING | CHECKED THE MACHINE AND FOUND THE POWERCHORD WAS DISCONNECTED FROM THE CONTROL BOX. RECONNECTED AND ISSUE RECTIFIED. NOW WORKING GOOD. | NO | NO | 2023-11-27 | November 2023 | 10:15 | 2023-11-27 10:00 | 2023-11-27 10:15 | 0:15 | COMPLETED | NON-CRITICAL
792 | 2023-11-27 | Nov 2023 | 12:30 | 12:35 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | EMERGENCY | LEADS ECG TEST FAILED | REMOVED THE ECG CABLE, CLEANED AND REFIXED. RERUNNED THE OPERATIONAL CHECK AND PASSED. NOW WORKING GOOD. | NO | NO | 2023-11-27 | November 2023 | 12:50 | 2023-11-27 12:30 | 2023-11-27 12:50 | 0:20 | COMPLETED | CRITICAL
793 | 2023-11-27 | Nov 2023 | 14:20 | 14:25 | 0.0034722 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | RAMKIE | EMERGENCY | Not switching ON | CHECKED AND FODUND THE PROBLM WITH POWER CORD AND REPLACED AND RECTIFIED | NO | NO | 2023-11-27 | November 2023 | 14:35 | 2023-11-27 14:20 | 2023-11-27 14:35 | 0:15 | COMPLETED | SEMI-CRITICAL
794 | 2023-11-28 | Nov 2023 | 22:30 | 22:35 | 0.0034722 | BIPAP MACHINE | PHILIPS | A40 | JERLIN | WARD | LOW CIRCUIT LEAK ALARM | ADJUSTED THE EXPIRATORY LEAK PORT AND ISSUE RECTIFIED | NO | NO | 2023-11-28 | November 2023 | 22:45 | 2023-11-28 22:30 | 2023-11-28 22:45 | 0:15 | COMPLETED | SEMI-CRITICAL
795 | 2023-11-28 | Nov 2023 | 9:45 | 9:50 | 0.0034722 | ECHO MACHINE | GE HEALTHCARE | VIVID IQ | SOORAJ | ICU | NOT ABLE TO SAVE PATIENT DATA | CONFIGURATION CHANGED SAVE TO LOCAL DRIVE AND FOUND ABLE TO SAVE PATIENT DATA. ISSUE RECTIFIED. | NO | NO | 2023-11-28 | November 2023 | 10:00 | 2023-11-28 09:45 | 2023-11-28 10:00 | 0:15 | COMPLETED | SEMI-CRITICAL
796 | 2023-11-28 | Nov 2023 | 9:00 | 9:05 | 0.0034722 | CENTRIFUGE TABLE TOP | REMI | NEYA 4 | HAFIZ | LABORATORY | CAN'T ABLE TO CHANGE THE TIME | OBSERVED AND CHANGED THE SETTINGS AS PER REQUIREMENTS. NOW, WORKING GOOD. | NO | NO | 2023-11-28 | November 2023 | 9:15 | 2023-11-28 09:00 | 2023-11-28 09:15 | 0:15 | COMPLETED | SEMI-CRITICAL
797 | 2023-11-28 | Nov 2023 | 11:45 | 11:50 | 0.0034722 | VIDEO LARYNGOSCOPE | AMBU KING VISION | KVLKIT 3 | HAFIZ | EMERGENCY | NOT WORKING | REMOVED THE BATTERY, CLEANED THE INTERFACE. REFIXED THE BATTERY AND RESTARTED. CHECKED AND NOW WORKING GOOD. | NO | NO | 2023-11-28 | November 2023 | 12:00 | 2023-11-28 11:45 | 2023-11-28 12:00 | 0:15 | COMPLETED | SEMI-CRITICAL
798 | 2023-11-28 | Nov 2023 | 12:20 | 12:25 | 0.0034722 | PATIENT MONITOR,VITAL SIGN | WELCH ALLYN | 75CT | HAFIZ | OPD | NOT WORKING AND BP VALUE DIFFERS | CHECKED THE MACHINE AND FOUND NO ISSUES WITH THE BP CUFF AND HOSE. TAKEN BP VALUES FROM AKAS AND CROSS-VERIFIED WITH THIS MACHINE. NO ISSUES, SAME VALUES FOUND. NOW, READY TO USE. | NO | NO | 2023-11-28 | November 2023 | 12:40 | 2023-11-28 12:20 | 2023-11-28 12:40 | 0:20 | COMPLETED | SEMI-CRITICAL
799 | 2023-11-28 | Nov 2023 | 13:40 | 13:45 | 0.0034722 | COBAS CORE UNIT | COBAS 6000 CORE | ROCHE | HAFIZ | LABORATORY | NOT WORKING | REALIGNED THE PROBE, ADJUSTED THE MODULE AND GRABBER AS PER THE COMPANY PERSON'S GUIDELINES. RESTARTED THE MACHINE AND NOW WORKING GOOD. | NO | NO | 2023-11-28 | November 2023 | 14:00 | 2023-11-28 13:40 | 2023-11-28 14:00 | 0:20 | COMPLETED | SEMI-CRITICAL
800 | 2023-11-28 | Nov 2023 | 16:20 | 16:30 | 0.0069444 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | RAMKIE | RADIOLOGY | WIFI BOX NOT SWITCHING ON | CHECKED AND POWER CORD GOT DISCONNECTED, RECTIFIED | NO | NO | 2023-11-28 | November 2023 | 16:40 | 2023-11-28 16:20 | 2023-11-28 16:40 | 0:20 | COMPLETED | SEMI-CRITICAL
801 | 2023-11-29 | Nov 2023 | 0:20 | 0:30 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | ICU | ECG EQUIPMENT MALFUNCTION | CLEANED THE ECG CABLE AND RERAN TEST. ISSUE RECTIFIED | NO | NO | 2023-11-29 | November 2023 | 0:40 | 2023-11-29 00:20 | 2023-11-29 00:40 | 0:20 | COMPLETED | CRITICAL
802 | 2023-11-29 | Nov 2023 | 7:00 | 7:10 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | CARTRIDGE EXPIRED | REPLACED THE OLD CARTRIDGE WITH A NEW ONE | CARTRIDGE | NO | 2023-11-29 | November 2023 | 7:20 | 2023-11-29 07:00 | 2023-11-29 07:20 | 0:20 | COMPLETED | SEMI-CRITICAL
803 | 2023-11-29 | Nov 2023 | 7:30 | 7:31 | 0.00069444 | OT TABLE | MAQUET | ALPHA CLASSIC PRO | SOORAJ | OT | NOT ABLE TO FIT CLAMPS ON SIDE RAIL. | FOUND SIDE RAILS LOOSENED. SECURED TIGHTELY AND ISSUE RECTIFIED. | NO | NO | 2023-11-29 | November 2023 | 7:40 | 2023-11-29 07:30 | 2023-11-29 07:40 | 0:10 | COMPLETED | NON-CRITICAL
804 | 2023-11-29 | Nov 2023 | 8:20 | 8:25 | 0.0034722 | MRI MACHINE | PHILIPS | INGENIA ELITON S | SOORAJ | RADIOLOGY | SYSTEM NOT SWITCHING ON | FOUND NO POWER ON UPS PORTS. INFORMED TO MAINTENANCE. UPS ISSUE GOT RECTIFIED AND ONE PATIENT SCAN WAS DONE, WORING FINE. | NO | NO | 2023-11-29 | November 2023 | 9:10 | 2023-11-29 08:20 | 2023-11-29 09:10 | 0:50 | COMPLETED | CRITICAL
805 | 2023-11-29 | Nov 2023 | 8:20 | 8:25 | 0.0034722 | RADIOGRAPHIC UNIT | CARESTREAM | DRF ASCEND ANOLOG HF | SOORAJ | RADIOLOGY | SYSTEM NOT SWITCHING ON | FOUND NO POWER ON UPS PORTS. INFORMED TO MAINTENANCE. UPS ISSUE GOT RECTIFIED AND ONE PATIENT SCAN WAS DONE, WORING FINE. | NO | NO | 2023-11-29 | November 2023 | 9:10 | 2023-11-29 08:20 | 2023-11-29 09:10 | 0:50 | COMPLETED | SEMI-CRITICAL
806 | 2023-11-29 | Nov 2023 | 8:30 | 8:35 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SOORAJ | ICU | CHANGE CARTRIDGE MESSAGE | NEW CARTRIDGE REPLACED AND CVP DONE. | CARTRIDGE | NO | 2023-11-29 | November 2023 | 9:27 | 2023-11-29 08:30 | 2023-11-29 09:27 | 0:57 | COMPLETED | SEMI-CRITICAL
807 | 2023-11-29 | Nov 2023 | 10:10 | 10:15 | 0.0034722 | IVUS MACHINE | PHILIPS | CORE SERIES S5 | SOORAJ | CATHLAB | CATHETER NOT FOUND | COMPANY PERSON CHECKED AND FOUND ROTATING PIM MODULE GONE DEFECTIVE AND NEEDS REPLACEMENT | NO | ROTATING PIM MODULE | 2023-11-29 | November 2023 | 10:45 | 2023-11-29 10:10 | 2023-11-29 10:45 | 0:35 | PENDING | SEMI-CRITICAL
808 | 2023-11-29 | Nov 2023 | 16:15 | 16:20 | 0.0034722 | ELECTRICAL SUCTION | YUWELL | 7A-23B | SOORAJ | OT | SUCTION NOT WORKING | CHECKED AND OBSERVED MOTOR NOT RUNNING, SUSPECT ISSUE WITH MOTOR. NEED TO SEND FOR SERVICE | NO | MOTOR | 2023-11-29 | November 2023 | 20:00 | 2023-11-29 16:15 | 2023-11-29 20:00 | 3:45 | PENDING | SEMI-CRITICAL
809 | 2023-11-29 | Nov 2023 | 16:30 | 16:32 | 0.0013889 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX 8 | SOORAJ | OT | BIPOLAR FORCEPS NOT WORKING DURING CASE. | CHECKED THE FORCEPS AND MACHINE AFTER CASE AND FOUND WORKING FINE. NO ISSUES FOUND WITH MACHINE AND WORKING SHOWN TO NURSING STAFF SUBASHINI | NO | NO | 2023-11-29 | November 2023 | 16:45 | 2023-11-29 16:30 | 2023-11-29 16:45 | 0:15 | COMPLETED | SEMI-CRITICAL
810 | 2023-11-29 | Nov 2023 | 16:45 | 16:50 | 0.0034722 | HARMONIC GENERATOR | LOTUS | UBERSONIC TM | SOORAJ | OT | HARMONIC NOT WORKING DURING CASE | CHECKED HANDPIECE AND CLEANED TIP WITH SPIRIT. CONNECTED AND OBSERVED THE MACHINE & HANDPIECE TO BE WORKING FINE. SHOWN TO NURSING STAFF SUBASHINI | NO | NO | 2023-11-29 | November 2023 | 17:00 | 2023-11-29 16:45 | 2023-11-29 17:00 | 0:15 | COMPLETED | SEMI-CRITICAL
811 | 2023-11-29 | Nov 2023 | 17:05 | 17:10 | 0.0034722 | LASER MACHINE,HOLMIUM | LUMEINUS | P-100H | SOORAJ | OT | 500NM PROBE NOT WORKING. | CROSS CHECKED WITH PROBE INSPECTION TOOL AND FOUND PROBE TO BE DEFECTIVE. INFORMED NURSING STAFF TO INDENT NEW PROBE FROM STORES. MACHINE CHECKED AND WORKING FINE. | NO | NO | 2023-11-29 | November 2023 | 17:15 | 2023-11-29 17:05 | 2023-11-29 17:15 | 0:10 | COMPLETED | SEMI-CRITICAL
812 | 2023-11-29 | Nov 2023 | 17:20 | 17:30 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | RAMKIE | ICU | ECG WAVE IS NOT GOOD | CHECKED AND ASKED TO CHANGE THE ELECTRODES AND RECTIFIED | ECG ELECTRODES | NO | 2023-11-29 | November 2023 | 17:50 | 2023-11-29 17:20 | 2023-11-29 17:50 | 0:30 | COMPLETED | CRITICAL
813 | 2023-11-30 | Nov 2023 | 1:00 | 1:10 | 0.0069444 | CEREBRAL OXIMETRY SYSTEM | COVIDIEN | SOMANETICS | JERLIN | OT | NOT WORKING | FOUND ISSUE WITH THE SENSOR, CHECKED WITH ANOTHER SENSOR AND CHECKED ITS WORKING GOOD | NO | NO | 2023-11-30 | November 2023 | 1:25 | 2023-11-30 01:00 | 2023-11-30 01:25 | 0:25 | COMPLETED | SEMI-CRITICAL
814 | 2023-11-30 | Nov 2023 | 1:00 | 1:10 | 0.0069444 | CEREBRAL OXIMETRY SYSTEM | COVIDIEN | SOMANETICS | JERLIN | OT | NOT WORKING | FOUND ISSUE WITH THE SENSOR, CHECKED WITH ANOTHER SENSOR AND CHECKED ITS WORKING GOOD | NO | NO | 2023-11-30 | November 2023 | 1:30 | 2023-11-30 01:00 | 2023-11-30 01:30 | 0:30 | COMPLETED | SEMI-CRITICAL
815 | 2023-11-30 | Nov 2023 | 23:00 | 23:05 | 0.0034722 | MULTIFUNCTIONAL BED | PARAMOUNT | PA93485V | JERLIN | WARD | REMOTE NOT WORKING | POWER SUPPLY ISSUE RECTIFIED | NO | NO | 2023-11-30 | November 2023 | 23:15 | 2023-11-30 23:00 | 2023-11-30 23:15 | 0:15 | COMPLETED | NON-CRITICAL
816 | 2023-11-30 | Nov 2023 | 21:15 | 21:20 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | JERLIN | OT | DISPLAY NOT WORKING | RESTARTED THE MACHINE AND IT WAS WORKING GOOD | NO | NO | 2023-11-30 | November 2023 | 21:30 | 2023-11-30 21:15 | 2023-11-30 21:30 | 0:15 | COMPLETED | SEMI-CRITICAL
817 | 2023-11-30 | Nov 2023 | 22:30 | 22:35 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | JERLIN | WARD | NIBP NOT WORKING | CHANGED THE CUFF'S CONNECTOR AND RECTIFIED THE ISSUE | NO | NO | 2023-11-30 | November 2023 | 22:45 | 2023-11-30 22:30 | 2023-11-30 22:45 | 0:15 | COMPLETED | SEMI-CRITICAL
818 | 2023-11-30 | Nov 2023 | 7:30 | 7:35 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SOORAJ | OT | VENTILATE MANUALLY ALARM | COMPANY PERSON CHECECKED AND FOUND ISSUE WITH SIGNAL INTERFACE BOARD. BOARD NEEDS REPLACEMENT | NO | SIGNAL INTERFACE BOARD | 2023-11-30 | November 2023 | 7:50 | 2023-11-30 07:30 | 2023-11-30 07:50 | 0:20 | PENDING | CRITICAL
819 | 2023-11-30 | Nov 2023 | 12:10 | 12:15 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | SOORAJ | OT | LUBDUB TEE PROBE ROTATION KNOB NOT WORKING BEFORE CASE | CHECKED AND FOUND MECHANICAL FAILURE OBSERVED ON FRONT & REVERSE KNOB, REST FUNCTION WORKING FINE ON TEE PROBE. MACHINE ALSO FOUND WORKING FINE. | NO | TEE PROBE | 2023-11-30 | November 2023 | 12:25 | 2023-11-30 12:10 | 2023-11-30 12:25 | 0:15 | PENDING | SEMI-CRITICAL
820 | 2023-11-30 | Nov 2023 | 15:30 | 15:40 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 550 | HAFIZ | OT | SWITCHING OFF AUTOMATICALLY ISSUE | SWAPPED THE INTERFACE BOARD FROM MONITOR (MGMNM/BME/00169) INTO MONITOR (MGMNM/BME/00170). CHECKED AND NOW WORKING GOOD. | NO | NO | 2023-11-30 | November 2023 | 16:30 | 2023-11-30 15:30 | 2023-11-30 16:30 | 1:00 | COMPLETED | SEMI-CRITICAL
821 | 2023-11-30 | Nov 2023 | 16:30 | 16:45 | 0.010417 | RADIOGRAPHIC UNIT | CARESTREAM | DRF ASCEND ANOLOG HF | HAFIZ | OPD | MONITOR TOUCH NOT WORKING | CLEANED THE TOUCH PAD, CHECKED AND NOW WORKING GOOD. | NO | NO | 2023-11-30 | November 2023 | 17:15 | 2023-11-30 16:30 | 2023-11-30 17:15 | 0:45 | COMPLETED | SEMI-CRITICAL
822 | 2023-12-01 | Dec 2023 | 10:40 | 10:42 | 0.0013889 | CAVITRON ULTRASONIC SURGICAL ASPIRATOR | INTEGRA | CUSA EXCEL + | SOORAJ | OT | CUSA PROBE NOT WORKING | CHECKED PROBE WASHER AND TIP USAGE, CHANGED DISTILLED WATER AND RAN TEST FOUND WORKING FINE. ISSUE RECTIFIED. | DISTILLED WATER | NO | 2023-12-01 | December 2023 | 10:55 | 2023-12-01 10:40 | 2023-12-01 10:55 | 0:15 | COMPLETED | CRITICAL
823 | 2023-12-01 | Dec 2023 | 17:00 | 17:20 | 0.013889 | HEAD LIGHT SOURCE, XENON | SUNOPTIC | TITAN X450 | HAFIZ | OT | LAMP HAS EXCEEDED USABLE LIFE | REMOVED THE OLD LAMP AND INSTALLED A NEW LAMP(S.NO: 4343009J, REF: SSX0450). NOW,WORKING GOOD. | XENON LAMP | NO | 2023-12-01 | December 2023 | 18:30 | 2023-12-01 17:00 | 2023-12-01 18:30 | 1:30 | COMPLETED | SEMI-CRITICAL
824 | 2023-12-01 | Dec 2023 | 12:30 | 12:35 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | WARD | ALARM INDICATION | IV LINE NOT PROPERLY FIXED, PROPERLY FIXED AND NOW WORKING GOOD. | NO | NO | 2023-12-01 | December 2023 | 12:40 | 2023-12-01 12:30 | 2023-12-01 12:40 | 0:10 | COMPLETED | NON-CRITICAL
825 | 2023-12-01 | Dec 2023 | 12:40 | 12:42 | 0.0013889 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | HAFIZ | WARD | NOT SWITCHING ON | FOUND THAT BATTERY FULLY DEPLETED. PLUUGED TO AC SOURCE, OBSERVED AND NOW WORKING GOOD. | NO | NO | 2023-12-01 | December 2023 | 12:50 | 2023-12-01 12:40 | 2023-12-01 12:50 | 0:10 | COMPLETED | SEMI-CRITICAL
826 | 2023-12-01 | Dec 2023 | 12:50 | 12:52 | 0.0013889 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | HAFIZ | WARD | PRINTER NOT WORKING | CHECKED THE MACHINE AND FOUND NO ISSUES WITH PRINTER PAPER. CLEANED THE PRINTER DOOR AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2023-12-01 | December 2023 | 13:10 | 2023-12-01 12:50 | 2023-12-01 13:10 | 0:20 | COMPLETED | SEMI-CRITICAL
827 | 2023-12-01 | Dec 2023 | 13:10 | 13:15 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | HAFIZ | WARD | NOT SWITCHING ON | CHECKED THE MACHINE AND THE CABLE GOT DETACHED AT THE BACK. FIXED IT AND WORKING GOOD. | NO | NO | 2023-12-01 | December 2023 | 13:20 | 2023-12-01 13:10 | 2023-12-01 13:20 | 0:10 | COMPLETED | SEMI-CRITICAL
828 | 2023-12-01 | Dec 2023 | 16:20 | 16:25 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | CATHLAB | BP CUFF NOT WORKING | FOUND THAT CONNECTOR PIN WAS DIFFERENT. REPLACED IT. NOW WORKING GOOD. | BP CONNECTOR | NO | 2023-12-01 | December 2023 | 16:40 | 2023-12-01 16:20 | 2023-12-01 16:40 | 0:20 | COMPLETED | SEMI-CRITICAL
829 | 2023-12-01 | Dec 2023 | 11:20 | 11:30 | 0.0069444 | RT POLYMERASE CHAIN REACTION | BIOMOLECULAR SYSTEM | MIC 9 | RAMKIE | LABORATORY | Error - Driver supply voltage | REPLACED THE POWER ADAPTER | POWER ADAPTOR | NO | 2023-12-01 | December 2023 | 11:40 | 2023-12-01 11:20 | 2023-12-01 11:40 | 0:20 | COMPLETED | SEMI-CRITICAL
830 | 2023-12-01 | Dec 2023 | 19:10 | 19:15 | 0.0034722 | MRI MACHINE | PHILIPS | INGENIA ELITON S | THANISH | RADIOLOGY | PATIENT DISCOMFORT ISSUE IN SPINE PROTOCOL | AIR FILTER BLOCKAGE IN MAGNETIC COOLING AREA, REPLACED IT | AIR FILTER | NO | 2023-12-01 | December 2023 | 20:30 | 2023-12-01 19:10 | 2023-12-01 20:30 | 1:20 | COMPLETED | CRITICAL
831 | 2023-12-01 | Dec 2023 | 20:45 | 20:55 | 0.0069444 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | THANISH | ICU | WARMER NOT WORKING | RECTIFIED THE WARMER HOSE HEATING CABLE ISSUE | NO | NO | 2023-12-01 | December 2023 | 21:05 | 2023-12-01 20:45 | 2023-12-01 21:05 | 0:20 | COMPLETED | SEMI-CRITICAL
832 | 2023-12-02 | Dec 2023 | 7:45 | 7:50 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | SOORAJ | OT | DISPLAY NOT SHOWING. | FOUND MACHINE WAS IN STANDBY MODE. DISABLED STANDBY MODE AND ISSUE RECTIFIED. | NO | NO | 2023-12-02 | December 2023 | 7:55 | 2023-12-02 07:45 | 2023-12-02 07:55 | 0:10 | COMPLETED | CRITICAL
833 | 2023-12-02 | Dec 2023 | 9:30 | 9:40 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 550 | SOORAJ | OT | ECG WAVEFORM NOT SHOWING. | FOUND ISSUE WITH ECG LEADS, REPLACED LEADS AND ISSUE RECTIFIED. | 5-LEADS ECG CABLE | NO | 2023-12-02 | December 2023 | 9:50 | 2023-12-02 09:30 | 2023-12-02 09:50 | 0:20 | COMPLETED | SEMI-CRITICAL
834 | 2023-12-02 | Dec 2023 | 10:00 | 10:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | HAFIZ | ICU | SPO2 PROBE NOT WORKING | PROPERLY FIXED THE PROBE. CHECKED AND NOW WORKING GOOD. | NO | NO | 2023-12-02 | December 2023 | 10:15 | 2023-12-02 10:00 | 2023-12-02 10:15 | 0:15 | COMPLETED | SEMI-CRITICAL
835 | 2023-12-02 | Dec 2023 | 12:00 | 12:10 | 0.0069444 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | HAFIZ | EMERGENCY | ARTIFACTS ISSUE | PROPERLY FIXED THE LEADS, CHANGED THE SETTINGS. NOW WORKING GOOD. | NO | NO | 2023-12-02 | December 2023 | 12:30 | 2023-12-02 12:00 | 2023-12-02 12:30 | 0:30 | COMPLETED | SEMI-CRITICAL
836 | 2023-12-02 | Dec 2023 | 15:00 | 15:05 | 0.0034722 | OXYGEN THERAPY UNIT | INSPIRED | O2FLO | HAFIZ | ICU | OXYGEN DELIVERY FAILED | CLEANED THE AIR FILTER, RESTARTED THE MACHINE. WORKING GOOD. | NO | NO | 2023-12-02 | December 2023 | 15:15 | 2023-12-02 15:00 | 2023-12-02 15:15 | 0:15 | COMPLETED | SEMI-CRITICAL
837 | 2023-12-02 | Dec 2023 | 15:30 | 15:40 | 0.0069444 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | RAMKIE | CATHLAB | NOT WORKING | CHECKED, NO ISSUES FOUND, IT WORKING FINE | NO | NO | 2023-12-02 | December 2023 | 16:00 | 2023-12-02 15:30 | 2023-12-02 16:00 | 0:30 | COMPLETED | NON-CRITICAL
838 | 2023-12-02 | Dec 2023 | 16:00 | 16:05 | 0.0034722 | ACT MACHINE | HELENO LABORATORIES | ACTALYKE | RAMKIE | CATHLAB | ERROR | COMPANY ENGG CHECKED AND CALIBRATED, IT WORKING FINE | NO | NO | 2023-12-02 | December 2023 | 16:25 | 2023-12-02 16:00 | 2023-12-02 16:25 | 0:25 | COMPLETED | SEMI-CRITICAL
839 | 2023-12-02 | Dec 2023 | 17:10 | 17:15 | 0.0034722 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | THANISH | RADIOLOGY | MOBILE X-RAY NOT SWITCHING ON | RECTIFIED THE PLUG TOP ISSUE | NO | NO | 2023-12-02 | December 2023 | 17:30 | 2023-12-02 17:10 | 2023-12-02 17:30 | 0:20 | COMPLETED | SEMI-CRITICAL
840 | 2023-12-03 | Dec 2023 | 15:00 | 15:15 | 0.010417 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | CVP FAILED | CHANGED THE HEATER BLOCK AND ISSUE RECTIFIED | HEATER BLOCK | NO | 2023-12-03 | December 2023 | 15:25 | 2023-12-03 15:00 | 2023-12-03 15:25 | 0:25 | COMPLETED | SEMI-CRITICAL
841 | 2023-12-03 | Dec 2023 | 17:30 | 17:40 | 0.0069444 | LENS INTEGRATED SYSTEM,ARTHROSCOPY | SMITH & NEPHEW | 72204354 | JERLIN | OT | PROBLEM IN THE DISPLAY | FOUND A SMALL PATCH IN THE SCOPE , CLEANED THE SCOPE AND CHECKED. FOUND NO ISSUES | NO | NO | 2023-12-03 | December 2023 | 17:50 | 2023-12-03 17:30 | 2023-12-03 17:50 | 0:20 | COMPLETED | SEMI-CRITICAL
842 | 2023-12-03 | Dec 2023 | 21:00 | 21:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | ICU | PRINTER NOT WORKIG | FOUND THAT PAPER WAS NOT FIXED PROPERLY. FIXED IT AND RECTIFIED THE ISSUE. | NO | NO | 2023-12-03 | December 2023 | 21:25 | 2023-12-03 21:00 | 2023-12-03 21:25 | 0:25 | COMPLETED | CRITICAL
843 | 2023-12-03 | Dec 2023 | 23:00 | 23:10 | 0.0069444 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | JERLIN | WARD | CLAMP NOT FIXING PROPERLY | FIXED IT AND RECTIFIED THE ISSUE | NO | NO | 2023-12-03 | December 2023 | 23:30 | 2023-12-03 23:00 | 2023-12-03 23:30 | 0:30 | COMPLETED | NON-CRITICAL
844 | 2023-12-04 | Dec 2023 | 5:30 | 5:35 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | JERLIN | WARD | DATE AND TIME CHANGED | CHANGED IT ISSUE RECTIFIED. | NO | NO | 2023-12-04 | December 2023 | 5:40 | 2023-12-04 05:30 | 2023-12-04 05:40 | 0:10 | COMPLETED | SEMI-CRITICAL
845 | 2023-12-04 | Dec 2023 | 21:00 | 21:10 | 0.0069444 | BIPAP MACHINE | PHILIPS | A40 | HAFIZ | WARD | ALARM INDICATION | SETTINGS CHANGED AS PER REQUIREMENTS. NOW WORKING GOOD. | NO | NO | 2023-12-04 | December 2023 | 21:20 | 2023-12-04 21:00 | 2023-12-04 21:20 | 0:20 | COMPLETED | SEMI-CRITICAL
846 | 2023-12-04 | Dec 2023 | 16:00 | 16:05 | 0.0034722 | ACT MACHINE | HELENO LABORATORIES | ACTALYKE | RAMKIE | CATHLAB | ABNORMAL NOISE AND WRONG OUTPUT | COMPANY ENGG CHECKED AND CALIBRATED, UNDER OBSERVATION | NO | NO | 2023-12-04 | December 2023 | 16:30 | 2023-12-04 16:00 | 2023-12-04 16:30 | 0:30 | COMPLETED | SEMI-CRITICAL
847 | 2023-12-05 | Dec 2023 | 6:30 | 6:35 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | ICU | NOT WORKING | REFIXED THE IV LINE, CHECKED AND WORKING GOOD. | NO | NO | 2023-12-05 | December 2023 | 6:50 | 2023-12-05 06:30 | 2023-12-05 06:50 | 0:20 | COMPLETED | NON-CRITICAL
848 | 2023-12-05 | Dec 2023 | 15:20 | 15:30 | 0.0069444 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | RAMKIE | ICU | PRINTER NOT WORKING | CHECKED AND CHANGED THE PAPER DIRECTION, NOW IT WORKING GOOD | NO | NO | 2023-12-05 | December 2023 | 15:40 | 2023-12-05 15:20 | 2023-12-05 15:40 | 0:20 | COMPLETED | SEMI-CRITICAL
849 | 2023-12-06 | Dec 2023 | 13:15 | 13:20 | 0.0034722 | IVUS MACHINE | PHILIPS | CORE SERIES S5 | SOORAJ | CATHLAB | PIM MODULE ALARM | FOUND RPIM MODULE DEFECTIVE AND REPLACED | PIM MODULE | NO | 2023-12-06 | December 2023 | 14:00 | 2023-12-06 13:15 | 2023-12-06 14:00 | 0:45 | COMPLETED | SEMI-CRITICAL
850 | 2023-12-06 | Dec 2023 | 7:00 | 7:02 | 0.0013889 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SOORAJ | OT | ECG TEST MALFUNCTION | CLEANED ECG CABLE AND RAN PREOP CHECK, ISSUE RECTIFIED. | NO | NO | 2023-12-06 | December 2023 | 7:05 | 2023-12-06 07:00 | 2023-12-06 07:05 | 0:05 | COMPLETED | CRITICAL
851 | 2023-12-06 | Dec 2023 | 21:10 | 21:15 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SOORAJ | WARD | COT NOT WORKING | FOUND ISSUE WITH POWER SOCKET. CONNECTED TO ANOTHER POWER SOCKET AND FOUND WORKING FINE. | NO | NO | 2023-12-06 | December 2023 | 21:25 | 2023-12-06 21:10 | 2023-12-06 21:25 | 0:15 | COMPLETED | NON-CRITICAL
852 | 2023-12-06 | Dec 2023 | 14:00 | 14:10 | 0.0069444 | CAMERA SYSTEM | KARL STORZ | TC200 | JERLIN | OPD | DISPLAY NOT WORKING | FOUND THAT LINK CABLE NOT FIXED PROPERLY. FIXEDD IT AND ISSUE RECTIFIED. | NO | NO | 2023-12-06 | December 2023 | 14:25 | 2023-12-06 14:00 | 2023-12-06 14:25 | 0:25 | COMPLETED | SEMI-CRITICAL
853 | 2023-12-06 | Dec 2023 | 11:50 | 11:55 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | ICU | OP CHECK FAILED | REFIXED THE THERAPHY KNOB, RERUNNED THE OPERATIONAL CHECK AND WORKING GOOD. | NO | NO | 2023-12-06 | December 2023 | 12:10 | 2023-12-06 11:50 | 2023-12-06 12:10 | 0:20 | COMPLETED | CRITICAL
854 | 2023-12-06 | Dec 2023 | 15:00 | 15:20 | 0.013889 | IVUS MACHINE | PHILIPS | CORE SERIES S5 | HAFIZ | CATHLAB | DEFECTIVE PIMr MODULE | REPLACED A NEW PIMr MODULE (PN:300003311863, S.NO:110061). NOW WORKING GOOD. | PIM MODULE | NO | 2023-12-06 | December 2023 | 15:40 | 2023-12-06 15:00 | 2023-12-06 15:40 | 0:40 | COMPLETED | SEMI-CRITICAL
855 | 2023-12-06 | Dec 2023 | 16:10 | 16:20 | 0.0069444 | BINOCULAR MICROSCOPE | MAGNUS | CH20iLED | RAMKIE | LABORATORY | VISION NOT CLEAR | CLEANED SESNOR AND ALSO NEED REPLACEMENT | NO | NO | 2023-12-06 | December 2023 | 16:40 | 2023-12-06 16:10 | 2023-12-06 16:40 | 0:30 | COMPLETED | SEMI-CRITICAL
856 | 2023-12-07 | Dec 2023 | 0:45 | 0:50 | 0.0034722 | BIPAP MACHINE | PHILIPS | A40 | SOORAJ | WARD | LOW CIRCUIT ALARM | FOUND LEAK LESS THAN 10ML, ADJUSTED PATIENT CIRCUIT AND ISSUE RECTIFIED. | NO | NO | 2023-12-07 | December 2023 | 0:55 | 2023-12-07 00:45 | 2023-12-07 00:55 | 0:10 | COMPLETED | SEMI-CRITICAL
857 | 2023-12-07 | Dec 2023 | 1:00 | 1:00 | 0 | PATIENT MONITOR | PHILIPS | MX 450 | SOORAJ | WARD | SPO2 READING NOT SHOWING. | FOUND SPO2 PROBE DEFECTIVE. CHANGED NEW SPO2 PROBE AND ISSUE RECTIFIED. | SPO2 PROBE | NO | 2023-12-07 | December 2023 | 1:15 | 2023-12-07 01:00 | 2023-12-07 01:15 | 0:15 | COMPLETED | SEMI-CRITICAL
858 | 2023-12-07 | Dec 2023 | 19:35 | 19:40 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | SOORAJ | OT | VIEW MONITOR NOT SHOWING | FOUND ISSUE WITH POWER SOCKET. CONNECTED TO ANOTHER POWER SOCKET AND FOUND WORKING FINE. | NO | NO | 2023-12-07 | December 2023 | 19:50 | 2023-12-07 19:35 | 2023-12-07 19:50 | 0:15 | COMPLETED | SEMI-CRITICAL
859 | 2023-12-07 | Dec 2023 | 21:20 | 21:25 | 0.0034722 | HEART LUNG MACHINE | MAQUET | HL20-4PUMPS | SOORAJ | OT | BURNING SMELL | SUSPECTED TEMPERORY ISSUE DUE TO HIGH HUMIDITY INSIDE OT. KEPT THE MACHINE OUTSIDE OT FOR TWO HOURS AND AGAIN CHECKED, FOUND NOT ISSUES. CROSS CHECKED WITH COMPANY ENGINEER NEXT DAY AND FOUND MACHINE WOORKING FINE. | NO | NO | 2023-12-07 | December 2023 | 23:45 | 2023-12-07 21:20 | 2023-12-07 23:45 | 2:25 | COMPLETED | CRITICAL
860 | 2023-12-07 | Dec 2023 | 22:50 | 22:52 | 0.0013889 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | SOORAJ | OT | PATIENT DATA NOT SAVING | FOUND ISSUE WITH CONFIGURATION. CHANGED CONFIGURATION AND ISSUE RECTIFIED. | NO | NO | 2023-12-07 | December 2023 | 23:00 | 2023-12-07 22:50 | 2023-12-07 23:00 | 0:10 | COMPLETED | SEMI-CRITICAL
861 | 2023-12-07 | Dec 2023 | 7:30 | 7:34 | 0.0027778 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | HAFIZ | OT | HOLDING CLAMP ISSUE | REMOVED THE CURRENTLY FIXED CLAMP, REPLACED WITH ANOTHER DEFECTIVE PUMP. NOW ISSUE RECTIFIED. | CLAMP | NO | 2023-12-07 | December 2023 | 7:50 | 2023-12-07 07:30 | 2023-12-07 07:50 | 0:20 | COMPLETED | NON-CRITICAL
862 | 2023-12-07 | Dec 2023 | 7:50 | 7:52 | 0.0013889 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | OT | KNOB BROKEN | KNOB REPLACED FROM DEFECTIVE PUMP AND NOW ISSUE RECTIFIED. | NO | NO | 2023-12-07 | December 2023 | 8:10 | 2023-12-07 07:50 | 2023-12-07 08:10 | 0:20 | COMPLETED | NON-CRITICAL
863 | 2023-12-07 | Dec 2023 | 14:10 | 14:30 | 0.013889 | WEIGHING MACHINE,5KG | ATMA | ATM5/05 | RAMKIE | BLOODBANK | NOT WORKING | PROBLEM WITH LOAD CELL | LOADCELL | NO | 2023-12-07 | December 2023 | 14:50 | 2023-12-07 14:10 | 2023-12-07 14:50 | 0:40 | COMPLETED | NON-CRITICAL
864 | 2023-12-08 | Dec 2023 | 23:30 | 23:32 | 0.0013889 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | SOORAJ | WARD | MONITOR NOT SWITCHING ON | BATTERY DEPLETED, CONNECTED TO POWER SUPPLY AND CHECKED WORKING FINE. ISSUE RECTIFIED. | NO | NO | 2023-12-08 | December 2023 | 23:40 | 2023-12-08 23:30 | 2023-12-08 23:40 | 0:10 | COMPLETED | SEMI-CRITICAL
865 | 2023-12-08 | Dec 2023 | 15:00 | 15:15 | 0.010417 | BIPAP MACHINE | PHILIPS | A40 | JERLIN | ICU | LOW LEAK ALARM | ADJUSTED THE LEAK PORT IN THE CONNECTOR AND RECTIFIED THE ISSUE | NO | NO | 2023-12-08 | December 2023 | 15:25 | 2023-12-08 15:00 | 2023-12-08 15:25 | 0:25 | COMPLETED | SEMI-CRITICAL
866 | 2023-12-08 | Dec 2023 | 16:30 | 16:35 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | JERLIN | OT | PEDIATRIC TEE PROBE NOT WORKING | REFIXED THE PROBE AND FOUND THAT IT'S IN GOOD WORKING CONDITION. | NO | NO | 2023-12-08 | December 2023 | 16:35 | 2023-12-08 16:30 | 2023-12-08 16:35 | 0:05 | COMPLETED | SEMI-CRITICAL
867 | 2023-12-08 | Dec 2023 | 8:45 | 8:50 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | OT | POWER EQUIPMENT MALFUNCTION | REMOVED THE BATTERY, REFIXED PROPERLY. RERUNNED THE OP-CHECK IN SERVICE MODE. PASSED THE CHECK AND NOW READY TO USE. | NO | NO | 2023-12-08 | December 2023 | 9:10 | 2023-12-08 08:45 | 2023-12-08 09:10 | 0:25 | COMPLETED | CRITICAL
868 | 2023-12-08 | Dec 2023 | 11:35 | 11:40 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | HAFIZ | OT | APL VALVE KNOB LOOSED | CLOSING VALVE IS LOOSENED. TIGHNED THE SCREWS AT BOTTOM, AND CHECKED. NOW READY TO USE. | NO | NO | 2023-12-08 | December 2023 | 12:15 | 2023-12-08 11:35 | 2023-12-08 12:15 | 0:40 | COMPLETED | CRITICAL
869 | 2023-12-08 | Dec 2023 | 13:00 | 13:10 | 0.0069444 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | HAFIZ | ICU | POWER SUPPLY BOARD ISSUE | REPLACED THE POWER SUPPLY BOARD WITH COMPANY PERSON. NOW WORKING GOOD. | POWER SUPPLY BOARD | NO | 2023-12-08 | December 2023 | 13:30 | 2023-12-08 13:00 | 2023-12-08 13:30 | 0:30 | COMPLETED | SEMI-CRITICAL
870 | 2023-12-08 | Dec 2023 | 14:00 | 14:10 | 0.0069444 | HEART LUNG MACHINE | MAQUET | HL20-4PUMPS | HAFIZ | OT | BURNING SMELL WHEN TURNED ON | CHECKED WITH THE COMPANY PERSON, NO BURNING SMELL. READY TO USE. | NO | NO | 2023-12-08 | December 2023 | 14:30 | 2023-12-08 14:00 | 2023-12-08 14:30 | 0:30 | COMPLETED | CRITICAL
871 | 2023-12-08 | Dec 2023 | 13:30 | 13:40 | 0.0069444 | AUTOMATED MICROBIAL ANALYSER | BIOMERIUX | VITEK 2 COMPACT 30 | RAMKIE | LABORATORY | PRINTER PROBLEM | AFTER REPLACING PRINTER PROBLEM, IT WORKING | PRINTER HEAD | NO | 2023-12-08 | December 2023 | 13:50 | 2023-12-08 13:30 | 2023-12-08 13:50 | 0:20 | COMPLETED | SEMI-CRITICAL
872 | 2023-12-08 | Dec 2023 | 14:00 | 14:10 | 0.0069444 | HAEMATOLOGY ANALYSER | XN 1000 | TRANSASIA BIOMEDICALS | RAMKIE | LABORATORY | DETECTION ERROR | LUBRICATED AND CLEANED, NOT IT WORKING FINE | NO | NO | 2023-12-08 | December 2023 | 14:30 | 2023-12-08 14:00 | 2023-12-08 14:30 | 0:30 | COMPLETED | SEMI-CRITICAL
873 | 2023-12-08 | Dec 2023 | 11:20 | 11:30 | 0.0069444 | AUTOMATED MICROBIAL ANALYSER | BIOMERIUX | mini VIDAS | RAMKIE | LABORATORY | ERROR HARDWARE FAILURE | AFTER RESTARTING PROBLEM RECTIFIED | NO | NO | 2023-12-08 | December 2023 | 11:50 | 2023-12-08 11:20 | 2023-12-08 11:50 | 0:30 | COMPLETED | SEMI-CRITICAL
874 | 2023-12-08 | Dec 2023 | 14:05 | 14:10 | 0.0034722 | SCOPE CABINET | ITPL | EC-GF-2018F | THANISH | ENDOSCOPY | SCOPE CABINET LOCK NOT WORKING | REPLACED THE LOCK.WORKING FINE | CABINET LOCK | NO | 2023-12-08 | December 2023 | 15:15 | 2023-12-08 14:05 | 2023-12-08 15:15 | 1:10 | COMPLETED | SEMI-CRITICAL
875 | 2023-12-09 | Dec 2023 | 21:30 | 21:35 | 0.0034722 | IVUS MACHINE | PHILIPS | CORE SERIES S5 | SOORAJ | CATHLAB | CARD CORDINOTOR ERROR | REMOVED GRAPHICS CARD AND REINSERTED, ISSUE RECTIFIED. | NO | NO | 2023-12-09 | December 2023 | 22:00 | 2023-12-09 21:30 | 2023-12-09 22:00 | 0:30 | COMPLETED | SEMI-CRITICAL
876 | 2023-12-09 | Dec 2023 | 21:45 | 21:50 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SOORAJ | ICU | REPLACE CARTRIDGE MESSAGE | CARTRIDGE EMPTY, NEW CARTRIDGE REPLACED AND CVP DONE | CARTRIDGE | NO | 2023-12-09 | December 2023 | 23:00 | 2023-12-09 21:45 | 2023-12-09 23:00 | 1:15 | COMPLETED | SEMI-CRITICAL
877 | 2023-12-09 | Dec 2023 | 9:45 | 9:50 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | JERLIN | ICU | MACHINE NOT SWITICHING ON | MACHINE WAS DETACHED FROM IT'S CHARGING PORT. FIXED IT AND RECTIFIED THE ISSUE. | NO | NO | 2023-12-09 | December 2023 | 10:00 | 2023-12-09 09:45 | 2023-12-09 10:00 | 0:15 | COMPLETED | SEMI-CRITICAL
878 | 2023-12-09 | Dec 2023 | 10:05 | 10:10 | 0.0034722 | MAMMOGRAPH UNIT | HOLOGIC | SELENIA DIMENSIONS | THANISH | RADIOLOGY | MAMMO NOT WORKING | RECTIFIED THE CONNECTIVITY ISSUE BETWEEN GENERATOR AND SYSTEM | NO | NO | 2023-12-09 | December 2023 | 10:35 | 2023-12-09 10:05 | 2023-12-09 10:35 | 0:30 | COMPLETED | SEMI-CRITICAL
879 | 2023-12-09 | Dec 2023 | 10:50 | 10:55 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | THANISH | EMERGENCY | ECG NOT WORKING | RECTIFIED THE ECG CABLE ISSUE.WORKING FINE. | NO | NO | 2023-12-09 | December 2023 | 11:10 | 2023-12-09 10:50 | 2023-12-09 11:10 | 0:20 | COMPLETED | SEMI-CRITICAL
880 | 2023-12-09 | Dec 2023 | 9:30 | 9:35 | 0.0034722 | C-ARM MACHINE | WIPRO GE HEALTH CARE | OEC ONE | THANISH | OT | C-ARM NOT WORKING | CAMERA CAPTURE ASSEMBLY FAILED | NO | CAMERA CAPTURE ASSEMBLY | 2023-12-09 | December 2023 | 10:30 | 2023-12-09 09:30 | 2023-12-09 10:30 | 1:00 | PENDING | CRITICAL
881 | 2023-12-10 | Dec 2023 | 22:15 | 22:20 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | SOORAJ | WARD | NOT ABLE TO TAKE BP FOR PATIENT. | OBSERVED CUFF SIZE SMALL FOR PATIENT. CHANGED CUFF SIZE AND ISSUE RECTIFIED. | BP CUFF | NO | 2023-12-10 | December 2023 | 22:25 | 2023-12-10 22:15 | 2023-12-10 22:25 | 0:10 | COMPLETED | SEMI-CRITICAL
882 | 2023-12-10 | Dec 2023 | 12:30 | 12:35 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | ENDOSCOPY | ECG EQUIPMENT MALFUNCTION | RERAN OP CHECK AND RECTIFIED THE ISSUE | NO | NO | 2023-12-10 | December 2023 | 12:45 | 2023-12-10 12:30 | 2023-12-10 12:45 | 0:15 | COMPLETED | CRITICAL
883 | 2023-12-10 | Dec 2023 | 14:00 | 14:15 | 0.010417 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | CVP DUE | CVP DONE | NO | NO | 2023-12-10 | December 2023 | 14:30 | 2023-12-10 14:00 | 2023-12-10 14:30 | 0:30 | COMPLETED | SEMI-CRITICAL
884 | 2023-12-11 | Dec 2023 | 0:30 | 0:35 | 0.0034722 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | SOORAJ | RADIOLOGY | DETECTOR PLATE NOT DETECTING | REMOVED BATTERY, CLEANED CONNECTORS AND SHAKED A LITTLE BEFORE SETTING UP. ISSUE RECTIFIED. | NO | NO | 2023-12-11 | December 2023 | 0:50 | 2023-12-11 00:30 | 2023-12-11 00:50 | 0:20 | COMPLETED | SEMI-CRITICAL
885 | 2023-12-11 | Dec 2023 | 0:50 | 0:50 | 0 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | SOORAJ | RADIOLOGY | DETECTOR PLATE NOT DETECTING | REMOVED BATTERY, CLEANED CONNECTORS AND SHAKED A LITTLE BEFORE SETTING UP. ISSUE RECTIFIED. | NO | NO | 2023-12-11 | December 2023 | 1:00 | 2023-12-11 00:50 | 2023-12-11 01:00 | 0:10 | COMPLETED | SEMI-CRITICAL
886 | 2023-12-11 | Dec 2023 | 9:30 | 9:35 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | WARD | MACHINE NOT CHARGING | CHECKED AND FOUND NO SUPPLY IN THE POWER SOCKET. ISSUE RECTIFIED | NO | NO | 2023-12-11 | December 2023 | 9:45 | 2023-12-11 09:30 | 2023-12-11 09:45 | 0:15 | COMPLETED | CRITICAL
887 | 2023-12-11 | Dec 2023 | 19:00 | 19:10 | 0.0069444 | BMI MACHINE | WELCH ALLYN | TANITA | HAFIZ | OPD | NOT WORKING | INTERFACE CABLE GOT DETACHED. REFIXED ALONG WITH RAMKIE, JERLIN. NOW, WORKING GOOD. | NO | NO | 2023-12-11 | December 2023 | 20:00 | 2023-12-11 19:00 | 2023-12-11 20:00 | 1:00 | COMPLETED | NON-CRITICAL
888 | 2023-12-11 | Dec 2023 | 21:20 | 21:25 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | HAFIZ | WARD | NOT WORKING | PLUGGED TO AC SOURCE, CHECKED AND WORKING GOOD. | NO | NO | 2023-12-11 | December 2023 | 21:30 | 2023-12-11 21:20 | 2023-12-11 21:30 | 0:10 | COMPLETED | NON-CRITICAL
889 | 2023-12-11 | Dec 2023 | 9:10 | 9:15 | 0.0034722 | ORTHOPANTOMOGRAPHY | SKANRAY DENTAL | MYRAY HYPERION X5 | THANISH | OPD | SYSTEM HANGING ISSUE | DELETED THE OLD DATA.WORKING FINE. | NO | NO | 2023-12-11 | December 2023 | 9:25 | 2023-12-11 09:10 | 2023-12-11 09:25 | 0:15 | COMPLETED | SEMI-CRITICAL
890 | 2023-12-11 | Dec 2023 | 11:10 | 11:15 | 0.0034722 | VIDEO PROCESSING SYSTEM,ENDOSCOPES | KARL STORZ | TH102 | THANISH | OPD | CAMERA CONSOLE NOT WORKING | CAMERA CONSOLE CABLE ISSUE RECTIFIED | NO | NO | 2023-12-11 | December 2023 | 11:25 | 2023-12-11 11:10 | 2023-12-11 11:25 | 0:15 | COMPLETED | SEMI-CRITICAL
891 | 2023-12-11 | Dec 2023 | 14:10 | 14:15 | 0.0034722 | DEFIBRILLATOR MONITOR WITH AED | PHILIPS | EFFICIA DFM 100 | THANISH | OT | INTERNOL PADDLES NOT WORKING | INTERNOL PADDLES CONNECTOR BEING DAMAGED BY USER.NOT WORKING AND NEED TO BUY NEW ONE. | NO | PADDLES | 2023-12-11 | December 2023 | 14:30 | 2023-12-11 14:10 | 2023-12-11 14:30 | 0:20 | PENDING | CRITICAL
892 | 2023-12-12 | Dec 2023 | 15:30 | 15:35 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | JERLIN | OT | NOT SWITCHING ON | RESTARTED THE MACHINE AND RECTIFIED THE ISSUE | NO | NO | 2023-12-12 | December 2023 | 15:55 | 2023-12-12 15:30 | 2023-12-12 15:55 | 0:25 | COMPLETED | CRITICAL
893 | 2023-12-12 | Dec 2023 | 16:30 | 16:35 | 0.0034722 | VIDEO PROCESSING SYSTEM,ENDOSCOPES | OLYMPUS | CV-190 | JERLIN | ENDOSCOPY | VISUAL NOT CAPTURED IN THE SOFTWARE | FOUND THAT OUTPUT VIDEO CABLE WAS NOT CONNECTED IN THE SYSTEM. CONNECTED AND RECTIFIED THE ISSUE | NO | NO | 2023-12-12 | December 2023 | 16:45 | 2023-12-12 16:30 | 2023-12-12 16:45 | 0:15 | COMPLETED | SEMI-CRITICAL
894 | 2023-12-12 | Dec 2023 | 22:10 | 22:15 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | HAFIZ | OT | AC POWER FAIL ISSUE | BATTERY DRAINED. PLUGGED TO AC SOURCE. NOW WORKING GOOD. | NO | BATTERY | 2023-12-12 | December 2023 | 22:30 | 2023-12-12 22:10 | 2023-12-12 22:30 | 0:20 | PENDING | CRITICAL
895 | 2023-12-12 | Dec 2023 | 22:30 | 22:30 | 0 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | HAFIZ | OT | CALIBRATION DUE | WEEKLY CALIBRATION DONE. NOW WORKING GOOD. | NO | NO | 2023-12-12 | December 2023 | 22:50 | 2023-12-12 22:30 | 2023-12-12 22:50 | 0:20 | COMPLETED | CRITICAL
896 | 2023-12-12 | Dec 2023 | 11:00 | 11:30 | 0.020833 | LAMINAR AIR FLOW | MICROFLOW DEVICES INDIA PVT LTD | MFD V 1200 | RAMKIE | LABORATORY | PLATE GOT DISCONNCETED | ATTACHED TAPES AND FIXED | NO | NO | 2023-12-12 | December 2023 | 12:50 | 2023-12-12 11:00 | 2023-12-12 12:50 | 1:50 | COMPLETED | NON-CRITICAL
897 | 2023-12-12 | Dec 2023 | 11:50 | 12:00 | 0.0069444 | FREEZER,DEEP FREEZING -40 | REMI | RQVD-200 PLUS | RAMKIE | LABORATORY | LOCKER PROBLEM | ADJUSTED LOCK AND RECTIFIED | NO | NO | 2023-12-12 | December 2023 | 12:10 | 2023-12-12 11:50 | 2023-12-12 12:10 | 0:20 | COMPLETED | SEMI-CRITICAL
898 | 2023-12-12 | Dec 2023 | 10:05 | 10:10 | 0.0034722 | VIDEO PROCESSING SYSTEM,ENDOSCOPES | KARL STORZ | TH102 | THANISH | OT | CAMERA CONSOLE NOT WORKING | CAMERA CONSOLE CONNECTOR PORT DAMAGED.PARTIALLY WORKING/NEED TO REPLACE WITH NEW ONE. | NO | CAMERA CONSOLE CONNECTOR CABLE | 2023-12-12 | December 2023 | 10:30 | 2023-12-12 10:05 | 2023-12-12 10:30 | 0:25 | PENDING | SEMI-CRITICAL
899 | 2023-12-13 | Dec 2023 | 15:30 | 15:35 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | ICU | NOT WORKING | CHANGED THE PRESSURE SENSOR AND RECTIFIED THE ISSUE. | NO | NO | 2023-12-13 | December 2023 | 15:55 | 2023-12-13 15:30 | 2023-12-13 15:55 | 0:25 | COMPLETED | NON-CRITICAL
900 | 2023-12-13 | Dec 2023 | 22:00 | 22:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | WARD | OP CHECK FAILED | REFIXED THE PADDLES AND RERUNNED THE OPERATIONAL CHECK AND WORKING GOOD. | NO | NO | 2023-12-13 | December 2023 | 22:30 | 2023-12-13 22:00 | 2023-12-13 22:30 | 0:30 | COMPLETED | CRITICAL
901 | 2023-12-13 | Dec 2023 | 12:30 | 12:40 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | RAMKIE | RADIOLOGY | WORKLIST AND PACS PROBLEM | CHANGED THE FLOW NAME AND RECTIFIED | NO | NO | 2023-12-13 | December 2023 | 12:50 | 2023-12-13 12:30 | 2023-12-13 12:50 | 0:20 | COMPLETED | SEMI-CRITICAL
902 | 2023-12-13 | Dec 2023 | 9:25 | 9:30 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | THANISH | OT | DISPLAY COLOUR CHANGING ISSUE | CLEANED THE DISPLAY PANEL AND CONNECTOR.WORKING FINE | NO | NO | 2023-12-13 | December 2023 | 9:50 | 2023-12-13 09:25 | 2023-12-13 09:50 | 0:25 | COMPLETED | SEMI-CRITICAL
903 | 2023-12-14 | Dec 2023 | 4:00 | 4:10 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | SODA LIME QUANTITY REDUCED AND CHANGED PATIENT CIRCUIT, ISSUE RECTIFIED. | NO | NO | 2023-12-14 | December 2023 | 4:50 | 2023-12-14 04:00 | 2023-12-14 04:50 | 0:50 | COMPLETED | CRITICAL
904 | 2023-12-14 | Dec 2023 | 12:30 | 12:40 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | RAMKIE | RADIOLOGY | PACS AND SYSTEM HANGING PROBLEM | AFTER CLEARING THE SPOOLER, PROBLEM RECTIFIED | NO | NO | 2023-12-14 | December 2023 | 12:50 | 2023-12-14 12:30 | 2023-12-14 12:50 | 0:20 | COMPLETED | SEMI-CRITICAL
905 | 2023-12-14 | Dec 2023 | 13:30 | 13:40 | 0.0069444 | ANKLE MOTION DEVICE | ANKLE MOTION PRIVATE LIMITED | AM-03 | RAMKIE | PHYSIOTHERAPY | CHARGER PROBLEM | SOLDERED AND RECTIFIED AND MACHINE SIDE CONNECTOR | NO | NO | 2023-12-14 | December 2023 | 14:10 | 2023-12-14 13:30 | 2023-12-14 14:10 | 0:40 | COMPLETED | SEMI-CRITICAL
906 | 2023-12-14 | Dec 2023 | 10:10 | 10:11 | 0.00069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SANTHOSH | OT | CIRCUIT LEAK TEST FAILED | ISSUE RECTIFIED PASSED FULL TEST | NO | NO | 2023-12-14 | December 2023 | 10:20 | 2023-12-14 10:10 | 2023-12-14 10:20 | 0:10 | COMPLETED | CRITICAL
907 | 2023-12-15 | Dec 2023 | 1:45 | 1:50 | 0.0034722 | BIPAP MACHINE | RESMED | STELLAR 150 | HAFIZ | WARD | ALARM INDICATION | SETTINGS NEED TO BE CHANGED, DONE. NOW WORKING GOOD. | NO | NO | 2023-12-15 | December 2023 | 2:10 | 2023-12-15 01:45 | 2023-12-15 02:10 | 0:25 | COMPLETED | SEMI-CRITICAL
908 | 2023-12-15 | Dec 2023 | 2:10 | 2:12 | 0.0013889 | BIPAP MACHINE | RESMED | STELLAR 150 | HAFIZ | WARD | NOT SWITCHING ON | BATTERY DRAINED. PLUGGED TO AC SOURCE. NOW WORKING GOOD. | NO | NO | 2023-12-15 | December 2023 | 2:30 | 2023-12-15 02:10 | 2023-12-15 02:30 | 0:20 | COMPLETED | SEMI-CRITICAL
909 | 2023-12-15 | Dec 2023 | 13:30 | 13:40 | 0.0069444 | CAMERA SYSTEM | KARL STORZ | TC200 | RAMKIE | OT | ERROR SOFTWARE UPDATE 3.0 NOT SUPPORTED | PROBLEM RECTIFIED AFTER RECONNECTING CABLES, NOW IT WORKING GOOD | NO | NO | 2023-12-15 | December 2023 | 14:10 | 2023-12-15 13:30 | 2023-12-15 14:10 | 0:40 | COMPLETED | SEMI-CRITICAL
910 | 2023-12-15 | Dec 2023 | 10:15 | 10:19 | 0.0027778 | PATIENT MONITOR | PHILIPS | MX 550 | SANTHOSH | OT | AUTOMATICALLY SCREEN TURN OFF,THEN DDNT WORK TOUCH SCREEN ALSO | CHECKED AND WORKING FINE MACHINE UNDER OBSERVATION | NO | NO | 2023-12-15 | December 2023 | 10:25 | 2023-12-15 10:15 | 2023-12-15 10:25 | 0:10 | COMPLETED | SEMI-CRITICAL
911 | 2023-12-15 | Dec 2023 | 15:10 | 15:15 | 0.0034722 | MONITOR HIGH DEFINITION | KARL STORZ | TM220 | SANTHOSH | OT | INCOMPATIBLE DEVICE FOUND | JUST AN NOTIFICATION CHECKED KARL STORZ AND HUGEMED PROCESSOR | NO | NO | 2023-12-15 | December 2023 | 15:18 | 2023-12-15 15:10 | 2023-12-15 15:18 | 0:08 | COMPLETED | SEMI-CRITICAL
912 | 2023-12-15 | Dec 2023 | 15:20 | 15:25 | 0.0034722 | MONITOR HIGH DEFINITION | KARL STORZ | TM220 | SANTHOSH | OT | MONITOR SCREEN SHOWS UPSIDE DOWN | CHANGED SETTINGS ISSUE RECTIFIED | NO | NO | 2023-12-15 | December 2023 | 15:30 | 2023-12-15 15:20 | 2023-12-15 15:30 | 0:10 | COMPLETED | SEMI-CRITICAL
913 | 2023-12-16 | Dec 2023 | 9:10 | 9:15 | 0.0034722 | HEART LUNG MACHINE | GETINGE | HL20-4PUMPS | SOORAJ | OT | WHILE SWITCHING BETWEEN AC & BATTERY, MAINS INDICATION IS NOT SHOWING. | FOUND ISSUE WITH CONTROL CONSOLE MODULE. REPLACED CONTROL CONSOLE MODULE AND ISSUE RECTIFIED. | CONTROL CONSOLE MODULE | NO | 2023-12-16 | December 2023 | 12:00 | 2023-12-16 09:10 | 2023-12-16 12:00 | 2:50 | COMPLETED | CRITICAL
914 | 2023-12-16 | Dec 2023 | 14:30 | 14:35 | 0.0034722 | IVUS MACHINE | PHILIPS | CORE SERIES S5 | SOORAJ | CATHLAB | CATHETER ERROR MESSAGE | FOUND ISSUE WITH CATHETER BEING USED. CHANGED CATHTER AND ISSUE RECTIFIED. | CATHETER | NO | 2023-12-16 | December 2023 | 14:50 | 2023-12-16 14:30 | 2023-12-16 14:50 | 0:20 | COMPLETED | SEMI-CRITICAL
915 | 2023-12-16 | Dec 2023 | 15:30 | 15:35 | 0.0034722 | HEART LUNG MACHINE | GETINGE | HL20-4PUMPS | JERLIN | OT | MACHINE NOT WORKING ON MAIN POWER SUPPLY | CHANGED THE CCO MODULE AND RECTIFIED THE ISSUE. | NO | NO | 2023-12-16 | December 2023 | 15:45 | 2023-12-16 15:30 | 2023-12-16 15:45 | 0:15 | COMPLETED | CRITICAL
916 | 2023-12-16 | Dec 2023 | 22:45 | 22:50 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | WARD | SPO2 PROBE NOT WORKING | PROPERLY FIXED THE PROBE. CHECKED AND NOW WORKING GOOD. | NO | NO | 2023-12-16 | December 2023 | 23:00 | 2023-12-16 22:45 | 2023-12-16 23:00 | 0:15 | COMPLETED | SEMI-CRITICAL
917 | 2023-12-16 | Dec 2023 | 17:40 | 17:50 | 0.0069444 | CARDIO PULMONORY EXERCISE TESTER | COSMED | FITMATE MED | RAMKIE | OPD | PRINTER PAPER REPLACEMENT | REPLACED PRINTER PAPER AND CHECKED IT WORKING | NO | NO | 2023-12-16 | December 2023 | 18:10 | 2023-12-16 17:40 | 2023-12-16 18:10 | 0:30 | COMPLETED | SEMI-CRITICAL
918 | 2023-12-16 | Dec 2023 | 10:10 | 10:15 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | THANISH | WARD | SHOCK NOT DELIVERING | RECTIFIED THE PADDLES ISSUE.WORKING FINE. | NO | NO | 2023-12-16 | December 2023 | 10:25 | 2023-12-16 10:10 | 2023-12-16 10:25 | 0:15 | COMPLETED | CRITICAL
919 | 2023-12-17 | Dec 2023 | 9:30 | 9:35 | 0.0034722 | STEAM STERILIZER | GETINGE | GSS67H102E | SOORAJ | CSSD | WATER LEAK | FOUND HEATING COIL AND GASKET DEFECTIVE. NEED REPLACMENT | NO | HEATING COIL | 2023-12-17 | December 2023 | 12:00 | 2023-12-17 09:30 | 2023-12-17 12:00 | 2:30 | PENDING | SEMI-CRITICAL
920 | 2023-12-17 | Dec 2023 | 1:30 | 1:35 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | HAFIZ | ICU | PREUSE CHECK FAILED | RERUNNED THE PRE-USE CHECK. PASSED AND NOW READY TO USE. | NO | NO | 2023-12-17 | December 2023 | 1:50 | 2023-12-17 01:30 | 2023-12-17 01:50 | 0:20 | COMPLETED | CRITICAL
921 | 2023-12-17 | Dec 2023 | 2:30 | 2:35 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | HAFIZ | ICU | NOT WORKING | FOUND THAT THERMISTOR RECEPTACLE WAS DETACHED FROM THE BACK. REFIXED AND NOW WORKING GOOD. | NO | NO | 2023-12-17 | December 2023 | 2:50 | 2023-12-17 02:30 | 2023-12-17 02:50 | 0:20 | COMPLETED | SEMI-CRITICAL
922 | 2023-12-18 | Dec 2023 | 14:45 | 14:50 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | JERLIN | ENDOSCOPY | ETCO2 NOT WORKING | ACTIVATED THE GAS MODULE SETTINGS IN THE MONITOR AND RECTIFIED THE ISSUE. | NO | NO | 2023-12-18 | December 2023 | 15:00 | 2023-12-18 14:45 | 2023-12-18 15:00 | 0:15 | COMPLETED | SEMI-CRITICAL
923 | 2023-12-18 | Dec 2023 | 16:10 | 16:20 | 0.0069444 | PRINTER,COMPUTED RADIOGRAPHY | CARESTREAM | DRY VIEW 6950 | RAMKIE | RADIOLOGY | PRINTER NOT WORKING | CLEANED AND COMP PERSON REMOVED THE JAMMED SHEET, ISSUE RECTIFIED | NO | NO | 2023-12-18 | December 2023 | 17:00 | 2023-12-18 16:10 | 2023-12-18 17:00 | 0:50 | COMPLETED | SEMI-CRITICAL
924 | 2023-12-19 | Dec 2023 | 11:50 | 11:55 | 0.0034722 | INCUBATOR,TRANSPORT | PHOENIX | TINC 101 | SOORAJ | ICU | LOW TEMPERATURE ALARM | INCUBATOR NOT REACHED SET TEMPERATURE. CHECKED ON SELF AND FOUND ALARM SILENCED AFTER REACHING SET TEMPERATURE, NO ISSUES FOUND NO ISSUES | NO | NO | 2023-12-19 | December 2023 | 12:10 | 2023-12-19 11:50 | 2023-12-19 12:10 | 0:20 | COMPLETED | SEMI-CRITICAL
925 | 2023-12-19 | Dec 2023 | 9:15 | 9:20 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | JERLIN | ENDOSCOPY | NOT SWITCHING ON | FOUND THAT POWER SUPPLY SWITCH WAS OFF. RECTIFIED THE ISSUE | NO | NO | 2023-12-19 | December 2023 | 9:30 | 2023-12-19 09:15 | 2023-12-19 09:30 | 0:15 | COMPLETED | CRITICAL
926 | 2023-12-19 | Dec 2023 | 10:00 | 10:15 | 0.010417 | PATIENT MONITOR | PHILIPS | MX 550 | JERLIN | OT | ETCO2 NOT DETECTED PROPERLY | CHANGED THE WATER TRAP AND ISSUE RECTIFIED. | WATER TRAP | NO | 2023-12-19 | December 2023 | 10:25 | 2023-12-19 10:00 | 2023-12-19 10:25 | 0:25 | COMPLETED | SEMI-CRITICAL
927 | 2023-12-19 | Dec 2023 | 11:10 | 11:15 | 0.0034722 | CAMERA SYSTEM | KARL STORZ | TC200 | JERLIN | OT | CONNECTOR END DAMAGED AND NOT WORKING. | FIXED THE CONNECTOR PROPERLY AND FOUND IN GOOD WORKING CONDITION | NO | NO | 2023-12-19 | December 2023 | 11:30 | 2023-12-19 11:10 | 2023-12-19 11:30 | 0:20 | COMPLETED | SEMI-CRITICAL
928 | 2023-12-19 | Dec 2023 | 15:10 | 15:20 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | RAMKIE | WARD | NOT WORKING | REMOTE PROBLEM, REPLACED IT AND RECTIFIED | NO | NO | 2023-12-19 | December 2023 | 15:40 | 2023-12-19 15:10 | 2023-12-19 15:40 | 0:30 | COMPLETED | NON-CRITICAL
929 | 2023-12-19 | Dec 2023 | 21:10 | 21:15 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | THANISH | ICU | MONITOR NOT SWITCHING ON | POWER SUPPLY BOARD FAILED. | NO | PSB | 2023-12-19 | December 2023 | 21:50 | 2023-12-19 21:10 | 2023-12-19 21:50 | 0:40 | PENDING | SEMI-CRITICAL
930 | 2023-12-19 | Dec 2023 | 22:10 | 22:15 | 0.0034722 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | THANISH | OPD | ECG WAVE DISTURBANCE | CLEANED THE ACQUISTION MODULE BOARD.WORKING FINE. | NO | NO | 2023-12-19 | December 2023 | 22:40 | 2023-12-19 22:10 | 2023-12-19 22:40 | 0:30 | COMPLETED | NON-CRITICAL
931 | 2023-12-20 | Dec 2023 | 11:32 | 11:35 | 0.0020833 | BMI MACHINE | AKAS MEDICAL | QVS-100 | SOORAJ | OPD | BP VALE VARIATION. | OBSERVED BP CUFF AND FOUND NO LEAK ON BP CUFF. CHECEKD ON SELF AND TWO OTHER PATIENTS WHILE CROSS CHECKING WITH MANUAL BP, VARIATION WAS WITHING ACCEPTABLE RANGE. NO ISSUES FOUND | NO | NO | 2023-12-20 | December 2023 | 12:20 | 2023-12-20 11:32 | 2023-12-20 12:20 | 0:48 | COMPLETED | NON-CRITICAL
932 | 2023-12-20 | Dec 2023 | 15:20 | 15:25 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | SOORAJ | ICU | ECG LEAD NOT DETECTING | CHECKED AND FOUND V3, V4 LEADS NOT CLEANED PROPERLY. CLEANED THE LEADS AND CHECKED ON SELF, WORKING FINE. ISSUE RECTIFIED. | NO | NO | 2023-12-20 | December 2023 | 15:45 | 2023-12-20 15:20 | 2023-12-20 15:45 | 0:25 | COMPLETED | SEMI-CRITICAL
933 | 2023-12-20 | Dec 2023 | 11:40 | 11:50 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 550 | JERLIN | OT | MACHINE RESTARTING AUTOMATICALLY | CHECKED AND FOUND NO ISSUES IN THE MONITOR | NO | NO | 2023-12-20 | December 2023 | 12:00 | 2023-12-20 11:40 | 2023-12-20 12:00 | 0:20 | COMPLETED | SEMI-CRITICAL
934 | 2023-12-20 | Dec 2023 | 15:05 | 15:10 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | HAFIZ | ICU | DATE AND TIME CHANGED | CHANGED THE SETTINGS AS PER REQUIREMENTS. WORKING GOOD. | NO | NO | 2023-12-20 | December 2023 | 15:20 | 2023-12-20 15:05 | 2023-12-20 15:20 | 0:15 | COMPLETED | SEMI-CRITICAL
935 | 2023-12-20 | Dec 2023 | 16:00 | 16:10 | 0.0069444 | WEIGHING MACHINE,5KG | ATMA | ATM5/05 | HAFIZ | BLOODBANK | NOT WORKING | NEW LOADCELL (SENSOR) REPLACED BY THE COMPANY PERSON | LOADCELL | NO | 2023-12-20 | December 2023 | 16:30 | 2023-12-20 16:00 | 2023-12-20 16:30 | 0:30 | COMPLETED | NON-CRITICAL
936 | 2023-12-20 | Dec 2023 | 19:50 | 19:55 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | HAFIZ | ICU | NEEDS TO REPLACE A CARTRIDGE | REMOVED THE OLD CARTRIDGE AND INSTALLED A NEW CARTRIDGE. MACHINE, WARMING UP. READY TO USE. | CARTRIDGE | NO | 2023-12-20 | December 2023 | 20:00 | 2023-12-20 19:50 | 2023-12-20 20:00 | 0:10 | COMPLETED | SEMI-CRITICAL
937 | 2023-12-20 | Dec 2023 | 20:00 | 20:05 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | HAFIZ | ICU | NEEDS TO REPLACE A CARTRIDGE | REMOVED THE OLD CARTRIDGE AND INSTALLED A NEW CARTRIDGE. MACHINE, WARMING UP. READY TO USE. | CARTRIDGE | NO | 2023-12-20 | December 2023 | 20:10 | 2023-12-20 20:00 | 2023-12-20 20:10 | 0:10 | COMPLETED | SEMI-CRITICAL
938 | 2023-12-20 | Dec 2023 | 14:10 | 14:20 | 0.0069444 | TRACTION WITH TREATMENT TABLE | CHATTANOOGA | TRITON DTS | RAMKIE | PHYSIOTHERAPY | CENTRE HIP ADDJUSTMENT NOT WORKING | ADJUSTED BOLT, FIXED AND RECTIFIED | NO | NO | 2023-12-20 | December 2023 | 14:40 | 2023-12-20 14:10 | 2023-12-20 14:40 | 0:30 | COMPLETED | SEMI-CRITICAL
939 | 2023-12-20 | Dec 2023 | 21:05 | 21:10 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | THANISH | ICU | NEED TO INSTALL ABG CATRIDGE | INSTALLED THE ABG CATRIDGE AND CVP PERFORMED.WORKING FINE | CARTRIDGE | NO | 2023-12-20 | December 2023 | 21:55 | 2023-12-20 21:05 | 2023-12-20 21:55 | 0:50 | COMPLETED | SEMI-CRITICAL
940 | 2023-12-20 | Dec 2023 | 22:05 | 22:10 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | THANISH | ICU | NEED TO INSTALL ABG CATRIDGE | INSTALLED THE ABG CATRIDGE AND CVP PERFORMED.WORKING FINE | CARTRIDGE | NO | 2023-12-20 | December 2023 | 23:10 | 2023-12-20 22:05 | 2023-12-20 23:10 | 1:05 | COMPLETED | SEMI-CRITICAL
941 | 2023-12-21 | Dec 2023 | 8:00 | 8:05 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | SOORAJ | RADIOLOGY | WHITE LINE SHOWING WHILE USING C16-D PROBE | FOUND PROBE DEFECTIVE, NEED TO REPLACE | NO | C1-6-D PROBE | 2023-12-21 | December 2023 | 13:00 | 2023-12-21 08:00 | 2023-12-21 13:00 | 5:00 | PENDING | SEMI-CRITICAL
942 | 2023-12-21 | Dec 2023 | 15:30 | 15:40 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 550 | JERLIN | CATHLAB | NOT ABLE TO CONNECT WITH SLAVE MONITOR | REFIXED THE CABLE AND RECTIFIED THE ISSUE | NO | NO | 2023-12-21 | December 2023 | 15:50 | 2023-12-21 15:30 | 2023-12-21 15:50 | 0:20 | COMPLETED | SEMI-CRITICAL
943 | 2023-12-21 | Dec 2023 | 8:00 | 8:10 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | JERLIN | OT | CIRCUIT LEAK TEST FAILED | REDUCED SODA LIME QUANTITY AND RERAN THE TEST AND RECTIFIED THE ISSUE | NO | NO | 2023-12-21 | December 2023 | 8:25 | 2023-12-21 08:00 | 2023-12-21 08:25 | 0:25 | COMPLETED | CRITICAL
944 | 2023-12-21 | Dec 2023 | 14:15 | 14:20 | 0.0034722 | VENTILATOR, NEONATAL | WIPRO GE HEALTH CARE | SLE 6000 | HAFIZ | ICU | LEAKING FRESH GAS - ALARM | FOUND ISSUE WITH TEST-LUNG. REPLACED AND WORKING GOOD | TEST-LUNG | NO | 2023-12-21 | December 2023 | 14:30 | 2023-12-21 14:15 | 2023-12-21 14:30 | 0:15 | COMPLETED | CRITICAL
945 | 2023-12-21 | Dec 2023 | 16:15 | 16:20 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | HAFIZ | ICU | BP CUFF NOT WORKING | CHECKED AND FOUND ISSUE WITH THE BP CUFF. REPLACED AND WORKING GOOD. | NO | NO | 2023-12-21 | December 2023 | 16:30 | 2023-12-21 16:15 | 2023-12-21 16:30 | 0:15 | COMPLETED | SEMI-CRITICAL
946 | 2023-12-21 | Dec 2023 | 15:00 | 15:10 | 0.0069444 | BLOOD CULTURE DETECTION SYSTEM | BACT ALERT 3D | BIOMERIUX | RAMKIE | LABORATORY | ABNORNOL NOISE | AFTER RESTART PROBLEM RECTIFIED | NO | NO | 2023-12-21 | December 2023 | 15:40 | 2023-12-21 15:00 | 2023-12-21 15:40 | 0:40 | COMPLETED | SEMI-CRITICAL
947 | 2023-12-21 | Dec 2023 | 20:10 | 20:15 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | THANISH | WARD | NIBP NOT WORKING | NIBP CUFF LEAKAGE ISSUE RECTIFIED. | NO | NO | 2023-12-21 | December 2023 | 20:30 | 2023-12-21 20:10 | 2023-12-21 20:30 | 0:20 | COMPLETED | SEMI-CRITICAL
948 | 2023-12-21 | Dec 2023 | 21:15 | 21:20 | 0.0034722 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | THANISH | WARD | NIBP NOT WORKING | NIBP CUFF LEAKAGE ISSUE RECTIFIED. | NO | NO | 2023-12-21 | December 2023 | 21:40 | 2023-12-21 21:15 | 2023-12-21 21:40 | 0:25 | COMPLETED | SEMI-CRITICAL
949 | 2023-12-22 | Dec 2023 | 9:05 | 9:10 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | JERLIN | OT | CALIBRATION FAILED | O2 SENSOR CLEANED AND REFIXED | NO | NO | 2023-12-22 | December 2023 | 9:20 | 2023-12-22 09:05 | 2023-12-22 09:20 | 0:15 | COMPLETED | CRITICAL
950 | 2023-12-22 | Dec 2023 | 16:30 | 16:40 | 0.0069444 | MICROSCOPE OPERATING, SURGICAL | CARL ZEISS | LUMERA 300 | JERLIN | OT | NOT ABLE TO CONNECT WITH SLAVE MONITOR | CONNECTED IT AND RECTIFIED THE ISSUE | NO | NO | 2023-12-22 | December 2023 | 16:55 | 2023-12-22 16:30 | 2023-12-22 16:55 | 0:25 | COMPLETED | SEMI-CRITICAL
951 | 2023-12-22 | Dec 2023 | 13:45 | 13:50 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | HAFIZ | ICU | NEEDS TO REPLACE A CARTRIDGE | REMOVED THE OLD CARTRIDGE AND INSTALLED A NEW CARTRIDGE(S.NO:301759073). MACHINE, WARMING UP. READY TO USE. | CARTRIDGE | NO | 2023-12-22 | December 2023 | 14:00 | 2023-12-22 13:45 | 2023-12-22 14:00 | 0:15 | COMPLETED | SEMI-CRITICAL
952 | 2023-12-22 | Dec 2023 | 11:00 | 11:20 | 0.013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | ICU | NOT WORKING | NEED TO REPLACE 2 NUMBERS OF HEIGHT ACTUATORS | NO | HEIGHT ACTUATORS | 2023-12-22 | December 2023 | 11:40 | 2023-12-22 11:00 | 2023-12-22 11:40 | 0:40 | PENDING | NON-CRITICAL
953 | 2023-12-22 | Dec 2023 | 22:10 | 22:15 | 0.0034722 | VENTILATOR, NEONATAL | WIPRO GE HEALTH CARE | SLE 5000 | THANISH | ICU | LOW MINUTE ALARM | ADJUSTED THE SETTINGS.WORKING FINE. | NO | NO | 2023-12-22 | December 2023 | 22:25 | 2023-12-22 22:10 | 2023-12-22 22:25 | 0:15 | COMPLETED | CRITICAL
954 | 2023-12-23 | Dec 2023 | 10:05 | 10:10 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | ICU | NOT WORKING | CHECKED AND FOUND THAT POWERCHORD WAS DETACHED FROM BOTTOM. RECONNECTED AND WORKING GOOD. | NO | NO | 2023-12-23 | December 2023 | 10:25 | 2023-12-23 10:05 | 2023-12-23 10:25 | 0:20 | COMPLETED | NON-CRITICAL
955 | 2023-12-23 | Dec 2023 | 11:30 | 11:35 | 0.0034722 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | HAFIZ | OPD | ARTIFACTS ISSUE | REMOVED THE INTERFACE CABLE, CLEANED AND REFIXED. CHECKED WITH A PATIENT. WORKING GOOD. | NO | NO | 2023-12-23 | December 2023 | 12:05 | 2023-12-23 11:30 | 2023-12-23 12:05 | 0:35 | COMPLETED | NON-CRITICAL
956 | 2023-12-23 | Dec 2023 | 11:30 | 11:35 | 0.0034722 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | HAFIZ | OPD | NO COMMUNICATION WITH ERGOMETER/TREADMILL | REMOVED THE CABLE & REFIXED. ENOBLED THE SETTINGS. CHECKED WITH A PATIENT. WORKING GOOD. | NO | NO | 2023-12-23 | December 2023 | 11:40 | 2023-12-23 11:30 | 2023-12-23 11:40 | 0:10 | COMPLETED | NON-CRITICAL
957 | 2023-12-23 | Dec 2023 | 11:00 | 11:20 | 0.013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | RAMKIE | WARD | NOT WORKING | PRAPLACED CONTROL BOX | CONTROL BOX | NO | 2023-12-23 | December 2023 | 11:40 | 2023-12-23 11:00 | 2023-12-23 11:40 | 0:40 | COMPLETED | NON-CRITICAL
958 | 2023-12-24 | Dec 2023 | 11:20 | 11:24 | 0.0027778 | MICROSCOPE OPERATING, SURGICAL | LEICA MICROSYSTEMS | LEICA OHX | HAFIZ | OT | PLUGTOP WIRE DISCONNECTED | NO DAMAGE TO PLUGTOP, RECONNECTED THE WIRES IN THE SAME PLUGTOP, CHECKED AND WORKING GOOD. | NO | NO | 2023-12-24 | December 2023 | 11:50 | 2023-12-24 11:20 | 2023-12-24 11:50 | 0:30 | COMPLETED | SEMI-CRITICAL
959 | 2023-12-24 | Dec 2023 | 11:50 | 11:55 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | WARD | SPO2 PROBE NOT WORKING | REPLACED SPO2 PROBE FROM WARDSIDE. NOW WORKING GOOD. | SPO2 PROBE | NO | 2023-12-24 | December 2023 | 12:10 | 2023-12-24 11:50 | 2023-12-24 12:10 | 0:20 | COMPLETED | SEMI-CRITICAL
960 | 2023-12-25 | Dec 2023 | 22:45 | 22:50 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | REFIXED BELLOWS AND REDUCED SODA LIME QTY. CLEANED MOISTURE ON SODA LIME JAR AND CHECKED, CHECKOUT PASSED. | NO | NO | 2023-12-25 | December 2023 | 23:40 | 2023-12-25 22:45 | 2023-12-25 23:40 | 0:55 | COMPLETED | CRITICAL
961 | 2023-12-26 | Dec 2023 | 6:10 | 6:15 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | SOORAJ | WARD | BP NOT SHWOING. | FOUND BP CUFF LEAK. REPLACED BP CUFF FROM ANOTHER MONITOR AND ISSUE RECTIFIED | BP CUFF | NO | 2023-12-26 | December 2023 | 6:30 | 2023-12-26 06:10 | 2023-12-26 06:30 | 0:20 | COMPLETED | SEMI-CRITICAL
962 | 2023-12-26 | Dec 2023 | 22:50 | 22:55 | 0.0034722 | VENTILATOR, NEONATAL | WIPRO GE HEALTH CARE | SLE 6000 | SOORAJ | ICU | LOW PRESSURE ALARM | FOUND LEAK ON AIR HOSE CONENCTOR. TIGHTENED THE CONNECTOR AND REFIXED PROPERLY, ISSUE RECTIFIED. | NO | NO | 2023-12-26 | December 2023 | 23:30 | 2023-12-26 22:50 | 2023-12-26 23:30 | 0:40 | COMPLETED | CRITICAL
963 | 2023-12-26 | Dec 2023 | 16:00 | 16:10 | 0.0069444 | INTRA ORAL XRAY | SKANRAY DENTAL | INTRA SKANDC | JERLIN | OPD | NOT WORKING | FIXED THE SENSOR AND RESTARTED THE MACHINE . ISSUE RECTIFIED | NO | NO | 2023-12-26 | December 2023 | 16:20 | 2023-12-26 16:00 | 2023-12-26 16:20 | 0:20 | COMPLETED | SEMI-CRITICAL
964 | 2023-12-26 | Dec 2023 | 15:20 | 15:30 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC V2 | JERLIN | EMERGENCY | NOT SWITCHING ON | FIXED THE POWER CORD PROPERLY AND ISSUE RECTIFIED | NO | NO | 2023-12-26 | December 2023 | 15:35 | 2023-12-26 15:20 | 2023-12-26 15:35 | 0:15 | COMPLETED | SEMI-CRITICAL
965 | 2023-12-26 | Dec 2023 | 8:15 | 8:20 | 0.0034722 | HAEMATOLOGY ANALYSER | XN 1000 | TRANSASIA BIOMEDICALS | HAFIZ | LABORATORY | MPA PRESSURE ERROR | REMOVED THE FILTER AS PER GUIDELINES. RUNNED A SAMPLE. NOW WORKING GOOD. | NO | NO | 2023-12-26 | December 2023 | 8:34 | 2023-12-26 08:15 | 2023-12-26 08:34 | 0:19 | COMPLETED | SEMI-CRITICAL
966 | 2023-12-26 | Dec 2023 | 8:30 | 8:37 | 0.0048611 | RADIOGRAPHIC UNIT | CARESTREAM | DRF ASCEND ANOLOG HF | HAFIZ | OPD | NOT SWITCHING ON | ISSUE WITH THE SOCKET. RECONNECTED AND NOW WORKING GOOD. | NO | NO | 2023-12-26 | December 2023 | 8:45 | 2023-12-26 08:30 | 2023-12-26 08:45 | 0:15 | COMPLETED | SEMI-CRITICAL
967 | 2023-12-26 | Dec 2023 | 8:45 | 8:45 | 0 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | HAFIZ | OPD | PACS ERROR (SERVICE IS UNOVAILABLE OR NETWORK IS OFFLINE) | FOUND THAT LAN CABLE DISCONNECTED. RECONNECTED AND ISSUE RECTIFIED. | NO | NO | 2023-12-26 | December 2023 | 8:55 | 2023-12-26 08:45 | 2023-12-26 08:55 | 0:10 | COMPLETED | SEMI-CRITICAL
968 | 2023-12-27 | Dec 2023 | 8:30 | 8:40 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S7 | JERLIN | RADIOLOGY | PROBE NOT DETECTED | REFIXED THE PROBES AND FOUND THAT IT'S WORKING GOOD | NO | NO | 2023-12-27 | December 2023 | 8:50 | 2023-12-27 08:30 | 2023-12-27 08:50 | 0:20 | COMPLETED | SEMI-CRITICAL
969 | 2023-12-27 | Dec 2023 | 7:45 | 7:50 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | REMOVED THE ABSORBER CANISTER, ADJUSTED. RUNNED THE FULL TEST AND PASSED. CIRCUIT LEAK IS IN NORMAL RANGE. NOW MACHINE READY TO USE. | NO | NO | 2023-12-27 | December 2023 | 8:10 | 2023-12-27 07:45 | 2023-12-27 08:10 | 0:25 | COMPLETED | CRITICAL
970 | 2023-12-28 | Dec 2023 | 0:00 | 0:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SOORAJ | OT | VAPORISER CHECK FAILED. | CHECKED AND FOUND VAPORISER NOT FIXED PROPERLY. REFIXED AND CHECCKED, ISSUE RECTIFIED. | NO | NO | 2023-12-28 | December 2023 | 0:30 | 2023-12-28 | 2023-12-28 00:30 | 0:30 | COMPLETED | CRITICAL
971 | 2023-12-28 | Dec 2023 | 2:30 | 2:35 | 0.0034722 | VENTILATOR, NEONATAL | MAQUET | SERVO-I | SOORAJ | ICU | BATTERY TEST SKIPPED DURING PREOP CHECK | FOUND BATTERIES DEFECTIVE, CONNECTED ON UPS AND WORKING FINE. BATTERY NEEDS REPLACEMENT | NO | BATTERY | 2023-12-28 | December 2023 | 2:45 | 2023-12-28 02:30 | 2023-12-28 02:45 | 0:15 | PENDING | CRITICAL
972 | 2023-12-28 | Dec 2023 | 22:00 | 22:05 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SOORAJ | WARD | ECG TEST FAILED | FOUND ECG CABLE SET DEFECTIVE. CHANGED CABLE SET AND ISSUER RECTIFIED. | ECG CABLE | NO | 2023-12-28 | December 2023 | 22:30 | 2023-12-28 22:00 | 2023-12-28 22:30 | 0:30 | COMPLETED | CRITICAL
973 | 2023-12-28 | Dec 2023 | 16:30 | 16:40 | 0.0069444 | NEBULIZER | JK MEDICALS | ULTIMA | JERLIN | WARD | KNOB BROKEN | FIXED IT WITH ANOTHER MACHINE'S UPPER CASE | NO | NO | 2023-12-28 | December 2023 | 16:55 | 2023-12-28 16:30 | 2023-12-28 16:55 | 0:25 | COMPLETED | SEMI-CRITICAL
974 | 2023-12-28 | Dec 2023 | 10:30 | 10:35 | 0.0034722 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | JERLIN | RADIOLOGY | MACHINE NOT WORKING | CHECKED AND FOUND THAT ITS WORKING GOOD. | NO | NO | 2023-12-28 | December 2023 | 10:45 | 2023-12-28 10:30 | 2023-12-28 10:45 | 0:15 | COMPLETED | SEMI-CRITICAL
975 | 2023-12-28 | Dec 2023 | 11:00 | 11:03 | 0.0020833 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX 8 | HAFIZ | OT | NOT WORKING | FOUND ISSUES WITH BIPOLAR CABLE. REPLACED IT, CHECKED AND NOW WORKING GOOD. | NO | NO | 2023-12-28 | December 2023 | 11:20 | 2023-12-28 11:00 | 2023-12-28 11:20 | 0:20 | COMPLETED | SEMI-CRITICAL
976 | 2023-12-29 | Dec 2023 | 3:00 | 3:05 | 0.0034722 | VENTILATOR, PEDIATRIC | MINDRAY MEDICAL INDIA | SV 300 | SOORAJ | EMERGENCY | BLOWER TEMPERATURE HIGH ALARM | FOUND FAN FILTER WITH DUST. CLEANED THE FILTER AND RECHECKED, CHECKEOUT PASSED AND ISSUE RECTIFIED. | NO | NO | 2023-12-29 | December 2023 | 4:00 | 2023-12-29 03:00 | 2023-12-29 04:00 | 1:00 | COMPLETED | CRITICAL
977 | 2023-12-29 | Dec 2023 | 6:10 | 6:15 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | SOORAJ | WARD | FOUND TUBE HOLDER BROKEN. | REPLACED NEW TUBE HOLDER FROM ANOTHER DEFECTIVE MACHINE AND CHECKED, ISSUE RECTIFIED. | TUBE HOLDER | NO | 2023-12-29 | December 2023 | 6:45 | 2023-12-29 06:10 | 2023-12-29 06:45 | 0:35 | COMPLETED | NON-CRITICAL
978 | 2023-12-29 | Dec 2023 | 5:30 | 5:35 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | SOORAJ | OT | INTERMITTENTLY SWITCHING ON/OFF | REFIXED ALL CONNECTORS AND OBSERVED FOR 1 HR, WORKING FINE AND ISSUE RECTIFIED. | NO | NO | 2023-12-29 | December 2023 | 6:50 | 2023-12-29 05:30 | 2023-12-29 06:50 | 1:20 | COMPLETED | SEMI-CRITICAL
979 | 2023-12-29 | Dec 2023 | 14:30 | 14:40 | 0.0069444 | CEREBRAL OXIMETRY SYSTEM | COVIDIEN | SOMANETICS | JERLIN | ICU | LEFT SENSOR NOT WORKING | CLEANED AND REFIXED THE CABLE AND FOUND IT'S WORKING GOOD | NO | NO | 2023-12-29 | December 2023 | 14:55 | 2023-12-29 14:30 | 2023-12-29 14:55 | 0:25 | COMPLETED | SEMI-CRITICAL
980 | 2023-12-29 | Dec 2023 | 13:20 | 13:30 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | ICU | THERAPY DELIVERY TEST DISABLED | RERAN OP CHECK AND RECTIFIED THE ISSUE | NO | NO | 2023-12-29 | December 2023 | 13:40 | 2023-12-29 13:20 | 2023-12-29 13:40 | 0:20 | COMPLETED | CRITICAL
981 | 2023-12-29 | Dec 2023 | 11:00 | 11:03 | 0.0020833 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX 8 | HAFIZ | OT | NOT WORKING | RESET DONE. CHECKED AND NOW WORKING GOOD. | NO | NO | 2023-12-29 | December 2023 | 11:20 | 2023-12-29 11:00 | 2023-12-29 11:20 | 0:20 | COMPLETED | SEMI-CRITICAL
982 | 2023-12-29 | Dec 2023 | 14:00 | 14:10 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | RAMKIE | WARD | SIDE RAIL PROBLEM | COMP PERSON, CHECKED AND RECTIFIED | NO | NO | 2023-12-29 | December 2023 | 14:30 | 2023-12-29 14:00 | 2023-12-29 14:30 | 0:30 | COMPLETED | NON-CRITICAL
983 | 2023-12-29 | Dec 2023 | 14:30 | 14:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | RAMKIE | WARD | NOT WORKING | REMOTE PROBLEM, REPLACED IT AND RECTIFIED | REMOTE | NO | 2023-12-29 | December 2023 | 14:45 | 2023-12-29 14:30 | 2023-12-29 14:45 | 0:15 | COMPLETED | NON-CRITICAL
984 | 2023-12-29 | Dec 2023 | 14:45 | 14:50 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 9000 | RAMKIE | OPD | Side rail problem | COMP PERSON, CHECKED AND RECTIFIED | NO | NO | 2023-12-29 | December 2023 | 15:00 | 2023-12-29 14:45 | 2023-12-29 15:00 | 0:15 | COMPLETED | NON-CRITICAL
985 | 2023-12-29 | Dec 2023 | 14:30 | 14:40 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | THANISH | ICU | NEED TO INSTALL ABG CATRIDGE | INSTALLED THE ABG CATRIDGE AND CVP PERFORMED.WORKING FINE | CARTRIDGE | NO | 2023-12-29 | December 2023 | 15:30 | 2023-12-29 14:30 | 2023-12-29 15:30 | 1:00 | COMPLETED | SEMI-CRITICAL
986 | 2023-12-29 | Dec 2023 | 16:10 | 16:15 | 0.0034722 | MRI MACHINE | PHILIPS | INGENIA ELITON S | THANISH | RADIOLOGY | IMAGE RE-CONSTRUCTION ERROR | RE-STARTED THE APPLICATION AND PROTOCOL.WORKING FINE | NO | NO | 2023-12-29 | December 2023 | 16:40 | 2023-12-29 16:10 | 2023-12-29 16:40 | 0:30 | COMPLETED | CRITICAL
987 | 2023-12-30 | Dec 2023 | 15:30 | 15:40 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 550 | JERLIN | OT | TOUCH NOT WORKING | CHECKED AND FOUND THAT ITS WORKING GOOD. | NO | NO | 2023-12-30 | December 2023 | 15:45 | 2023-12-30 15:30 | 2023-12-30 15:45 | 0:15 | COMPLETED | SEMI-CRITICAL
988 | 2023-12-30 | Dec 2023 | 7:35 | 7:40 | 0.0034722 | HEATER COOLER | GETINGE | HCU40 | JERLIN | OT | LEAK IN THE BLANKET | SEALED THE TEAR IN THE BLANKET AND RECTIFIED THE ISSUE | NO | NO | 2023-12-30 | December 2023 | 8:00 | 2023-12-30 07:35 | 2023-12-30 08:00 | 0:25 | COMPLETED | CRITICAL
989 | 2023-12-30 | Dec 2023 | 16:30 | 16:50 | 0.013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | RAMKIE | WARD | SIDE RAIL PROBLEM | SPRAYED WD-40 AND ISSUE RECTIFIED | NO | NO | 2023-12-30 | December 2023 | 17:00 | 2023-12-30 16:30 | 2023-12-30 17:00 | 0:30 | COMPLETED | NON-CRITICAL
990 | 2023-12-30 | Dec 2023 | 15:05 | 15:10 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | THANISH | ICU | NEED TO INSTALL ABG CATRIDGE | INSTALLED THE ABG CATRIDGE AND CVP PERFORMED.WORKING FINE | CARTRIDGE | NO | 2023-12-30 | December 2023 | 16:10 | 2023-12-30 15:05 | 2023-12-30 16:10 | 1:05 | COMPLETED | SEMI-CRITICAL
991 | 2024-01-01 | Jan 2024 | 21:40 | 21:45 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | HAFIZ | WARD | NOT WORKING | CHECKED AND FOUND THAT THE THERMISTOR RECEPTACLE WAS DETACHED FROM THE BACK. REFIXED AND CHECKED. NOW WORKING GOOD. | NO | NO | 2024-01-01 | January 2024 | 22:00 | 2024-01-01 21:40 | 2024-01-01 22:00 | 0:20 | COMPLETED | SEMI-CRITICAL
992 | 2024-01-01 | Jan 2024 | 22:00 | 22:03 | 0.0020833 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | WARD | SPO2 PROBE NOT WORKING | CHECKED AND FOUND NO ISSUES WITH THE PROBE. RECONNECTED THE PROBE PROPERLY AT BOTH ENDS. NOW WORKING GOOD. | NO | NO | 2024-01-01 | January 2024 | 22:10 | 2024-01-01 22:00 | 2024-01-01 22:10 | 0:10 | COMPLETED | SEMI-CRITICAL
993 | 2024-01-01 | Jan 2024 | 22:55 | 23:00 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | PRINTER NOT WORKING | CHECKED AND FOUND ISSUES WITH THE PRINTER PAPER. CROSS-CHECKED WTH ANOTHER PAPER ROLL. NOW WORKING GOOD. | PAPER ROLL | NO | 2024-01-01 | January 2024 | 23:15 | 2024-01-01 22:55 | 2024-01-01 23:15 | 0:20 | COMPLETED | SEMI-CRITICAL
994 | 2024-01-01 | Jan 2024 | 8:05 | 8:10 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | THANISH | ICU | REMOTE NOT WORKING | RESETTED THE COT REMOTE.WORKING FINE. | NO | NO | 2024-01-01 | January 2024 | 8:15 | 2024-01-01 08:05 | 2024-01-01 08:15 | 0:10 | COMPLETED | NON-CRITICAL
995 | 2024-01-01 | Jan 2024 | 9:10 | 9:15 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | THANISH | OT | TIDAL VOLUME ABNORMAL ISSUE | CLEANED THE FLOW SENSOR AND PERFORMED THE FLOW SENSOR CALIBRATION.WORKING FINE. | NO | NO | 2024-01-01 | January 2024 | 9:30 | 2024-01-01 09:10 | 2024-01-01 09:30 | 0:20 | COMPLETED | CRITICAL
996 | 2024-01-01 | Jan 2024 | 9:35 | 9:40 | 0.0034722 | PLASMA STERILIZER | JOHNSON&JOHNSON | STERRAD 100NX | THANISH | CSSD | H2O2 CYCLE ISSUE | RE-STARTED THE MACHINE.WORKING FINE. | NO | NO | 2024-01-01 | January 2024 | 9:55 | 2024-01-01 09:35 | 2024-01-01 09:55 | 0:20 | COMPLETED | SEMI-CRITICAL
997 | 2024-01-01 | Jan 2024 | 11:40 | 11:50 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | RAMKIE | RADIOLOGY | MACHINE HANGING AND UNOBLE TO PUSH PACS | CHECKED AND CLEARED PENDING PACS AND PROBLEM GOT RECTIFIED | NO | NO | 2024-01-01 | January 2024 | 12:10 | 2024-01-01 11:40 | 2024-01-01 12:10 | 0:30 | COMPLETED | SEMI-CRITICAL
998 | 2024-01-02 | Jan 2024 | 8:10 | 8:15 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | THANISH | ICU | PRINTER NOT WORKING FOR DEFIBRILLATOR | CLEANED THR PRINTER HEAD.WORKIGN FINE. | NO | NO | 2024-01-02 | January 2024 | 8:20 | 2024-01-02 08:10 | 2024-01-02 08:20 | 0:10 | COMPLETED | CRITICAL
999 | 2024-01-02 | Jan 2024 | 9:15 | 9:20 | 0.0034722 | PLASMA STERILIZER | JOHNSON&JOHNSON | STERRAD 100NX | THANISH | CSSD | H202 ADJUSTMENT FAILED. | REPLACED THE UV LAMP AND OPTICAL SENSOR.WORKING FINE. | UV LAMP & OPTICAL SENSOR | NO | 2024-01-02 | January 2024 | 10:50 | 2024-01-02 09:15 | 2024-01-02 10:50 | 1:35 | COMPLETED | SEMI-CRITICAL
1000 | 2024-01-02 | Jan 2024 | 11:15 | 11:20 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | THANISH | OT | SHOCK ABORTED ISSUE. | RECTIFIED THE PADDLES ISSUE.WORKING FINE. | NO | NO | 2024-01-02 | January 2024 | 11:30 | 2024-01-02 11:15 | 2024-01-02 11:30 | 0:15 | COMPLETED | CRITICAL
1001 | 2024-01-02 | Jan 2024 | 13:05 | 13:10 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID E95 | THANISH | OPD | SETTINGS ISSUE | SETTINGS ISSUE RECTIFIED. | NO | NO | 2024-01-02 | January 2024 | 13:20 | 2024-01-02 13:05 | 2024-01-02 13:20 | 0:15 | COMPLETED | SEMI-CRITICAL
1002 | 2024-01-02 | Jan 2024 | 14:30 | 14:40 | 0.0069444 | ANKLE MOTION DEVICE | ANKLE MOTION PRIVATE LIMITED | AM-03 | RAMKIE | PHYSIOTHERAPY | NOT WORKING | CHECKED AND FOUND THE BREAKAGE IN POWER CONNECTING, SO IT TAKEN BY COMPANY PERSON FOR INHOUSE SERVICE | NO | POWER SOCKET | 2024-01-02 | January 2024 | 15:10 | 2024-01-02 14:30 | 2024-01-02 15:10 | 0:40 | PENDING | SEMI-CRITICAL
1003 | 2024-01-02 | Jan 2024 | 15:10 | 15:15 | 0.0034722 | ANKLE MOTION DEVICE | ANKLE MOTION PRIVATE LIMITED | AM-03 | RAMKIE | PHYSIOTHERAPY | NOT WORKING | CHECKED AND FOUND THE PROBLEM WITH POWER ADAPTER, SO SERVICED THAT ADAPTER AND PROBLEM RECTIFIED | NO | NO | 2024-01-02 | January 2024 | 15:55 | 2024-01-02 15:10 | 2024-01-02 15:55 | 0:45 | COMPLETED | SEMI-CRITICAL
1004 | 2024-01-02 | Jan 2024 | 11:30 | 11:40 | 0.0069444 | FREEZER,DEEP FREEZING -80 | WHEECON INSTRUMENTS PVT LTD | HG-ULTF-ARC200-80 | JERLIN | LABORATORY | TEMPERATURE ALARM | RESTARTED THE COMPRESSOR AND IT'S WORKING GOOD. | NO | NO | 2024-01-02 | January 2024 | 11:50 | 2024-01-02 11:30 | 2024-01-02 11:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1005 | 2024-01-03 | Jan 2024 | 22:40 | 22:43 | 0.0020833 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | WARD | NOT WORKING | CHECKED AND FOUND BATTERY FULLY DEPLETED. PLUGGED TO AC SOURCE, OBSERVED AND WORKING GOOD. | NO | NO | 2024-01-03 | January 2024 | 23:00 | 2024-01-03 22:40 | 2024-01-03 23:00 | 0:20 | COMPLETED | NON-CRITICAL
1006 | 2024-01-03 | Jan 2024 | 23:00 | 23:05 | 0.0034722 | HOLTER,RECORDER | WIPRO GE HEALTH CARE | SEER 1000 | HAFIZ | OPD | NOT WORKING | CHECKED AND FOUND BATTERY NOT IN CORRECT POSITION. PROPERLY PLACED. NOW WORKING GOOD. | NO | NO | 2024-01-03 | January 2024 | 23:15 | 2024-01-03 23:00 | 2024-01-03 23:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1007 | 2024-01-03 | Jan 2024 | 9:10 | 9:15 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | THANISH | ICU | PADDLES ISSUE | CLEANED THE PADDLES CONENCTOR.WORKING FINE. | NO | NO | 2024-01-03 | January 2024 | 9:30 | 2024-01-03 09:10 | 2024-01-03 09:30 | 0:20 | COMPLETED | CRITICAL
1008 | 2024-01-03 | Jan 2024 | 11:10 | 11:15 | 0.0034722 | BIPAP MACHINE | PHILIPS | A40 | THANISH | WARD | SETTINGS ISSUE | RECTIFIED THE SETTINGS ISSUE.WORKING FINE. | NO | NO | 2024-01-03 | January 2024 | 11:30 | 2024-01-03 11:10 | 2024-01-03 11:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1009 | 2024-01-03 | Jan 2024 | 14:00 | 14:30 | 0.020833 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | RAMKIE | EMERGENCY | CLAMP BROKEN | REPLACED THE CLAMP AND RECTIFIED | CLAMP | NO | 2024-01-03 | January 2024 | 14:50 | 2024-01-03 14:00 | 2024-01-03 14:50 | 0:50 | COMPLETED | NON-CRITICAL
1010 | 2024-01-03 | Jan 2024 | 15:10 | 15:15 | 0.0034722 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | RAMKIE | EMERGENCY | CLAMP BROKEN | REPLACED BACK PANEL AND RECTIFIED | BACK PANEL | NO | 2024-01-03 | January 2024 | 15:50 | 2024-01-03 15:10 | 2024-01-03 15:50 | 0:40 | COMPLETED | NON-CRITICAL
1011 | 2024-01-03 | Jan 2024 | 14:00 | 14:10 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 9000 | JERLIN | ICU | HEIGHT ADJUSTMENT PROBLEM | CLEANED THE SENSORS AND RECTIFIED THE ISSUE. | NO | NO | 2024-01-03 | January 2024 | 14:20 | 2024-01-03 14:00 | 2024-01-03 14:20 | 0:20 | COMPLETED | NON-CRITICAL
1012 | 2024-01-03 | Jan 2024 | 15:00 | 15:05 | 0.0034722 | CATHLAB - BIPLANE | PHILIPS | AZURION 7B20 | JERLIN | CATHLAB | COLLISION SENSOR PROBLEM | CHECKED AND IT'S WORKING GOOD AFTER MAINTAINING TEMP. | NO | NO | 2024-01-03 | January 2024 | 15:20 | 2024-01-03 15:00 | 2024-01-03 15:20 | 0:20 | COMPLETED | CRITICAL
1013 | 2024-01-03 | Jan 2024 | 12:30 | 12:35 | 0.0034722 | MONITOR HIGH DEFINITION | OLYMPUS | OEV262H | JERLIN | OPD | NEED TO CHANGE PIP SETTINGS | CHANGED THE SETTINGS AND RECTIFIED THE ISSUE. | NO | NO | 2024-01-03 | January 2024 | 12:50 | 2024-01-03 12:30 | 2024-01-03 12:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1014 | 2024-01-03 | Jan 2024 | 10:30 | 10:40 | 0.0069444 | MOVEMENT THERAPHY UNIT, UPPER AND LOWER LIMBS | RECK | MOTOMED LETTO2 | JERLIN | PHYSIOTHERAPY | NOT WORKING | CHECKED AND FOUND IT'S WORKING GOOD | NO | NO | 2024-01-03 | January 2024 | 10:55 | 2024-01-03 10:30 | 2024-01-03 10:55 | 0:25 | COMPLETED | SEMI-CRITICAL
1015 | 2024-01-03 | Jan 2024 | 9:30 | 9:40 | 0.0069444 | BMI MACHINE | WELCH ALLYN | TANITA | JERLIN | OPD | NOT WORKING | CHECKED AND IT'S WORKING GOOD. | NO | NO | 2024-01-03 | January 2024 | 9:54 | 2024-01-03 09:30 | 2024-01-03 09:54 | 0:24 | COMPLETED | NON-CRITICAL
1016 | 2024-01-03 | Jan 2024 | 12:15 | 12:20 | 0.0034722 | MONITOR HIGH DEFINITION | OLYMPUS | OEV262H | JERLIN | OPD | NOT ABLE TO RECORD THE PROCEDURE | CONNECTED THE VIDEO OUT CABLE AND RECORDED THE PROCEDURE | NO | NO | 2024-01-03 | January 2024 | 12:30 | 2024-01-03 12:15 | 2024-01-03 12:30 | 0:15 | COMPLETED | SEMI-CRITICAL
1017 | 2024-01-04 | Jan 2024 | 10:05 | 10:10 | 0.0034722 | HEATER COOLER | GETINGE | HCU40 | THANISH | OT | HEMOTHERM BLANKET WATER LEAKAGE | LEAK ARRESTED IN HEMOTHERM BLANKET. | NO | NO | 2024-01-04 | January 2024 | 10:30 | 2024-01-04 10:05 | 2024-01-04 10:30 | 0:25 | COMPLETED | CRITICAL
1018 | 2024-01-04 | Jan 2024 | 14:10 | 14:15 | 0.0034722 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX 8 | THANISH | OT | BIPOLAR NOT WORKING | REPAIRED THE BIPOLAR FOOTSWITCH.WORKING FINE. | NO | NO | 2024-01-04 | January 2024 | 14:40 | 2024-01-04 14:10 | 2024-01-04 14:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1019 | 2024-01-04 | Jan 2024 | 15:20 | 15:30 | 0.0069444 | CT MACHINE | PHILIPS | INGENUITY | RAMKIE | RADIOLOGY | CT SYSTEM MONITOR NOT WORKING | CHECKED AND REPLACED THE DP TO HDMI CABLE | DP TO HDMI CABLE | NO | 2024-01-04 | January 2024 | 15:50 | 2024-01-04 15:20 | 2024-01-04 15:50 | 0:30 | COMPLETED | CRITICAL
1020 | 2024-01-05 | Jan 2024 | 10:05 | 10:10 | 0.0034722 | IABP MACHINE | MAQUET | CS300 | THANISH | OT | BATTERY ISSUE | RESETTED THE BATTERY CHANGE ALARM IN SERVICE MODE.WORKING FINE. | NO | NO | 2024-01-05 | January 2024 | 10:30 | 2024-01-05 10:05 | 2024-01-05 10:30 | 0:25 | COMPLETED | CRITICAL
1021 | 2024-01-05 | Jan 2024 | 13:05 | 13:10 | 0.0034722 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | THANISH | WARD | NIBP CUFF LEAKAGE | REPAIRED THE NIBP CUFF.WORKING FINE | NO | NO | 2024-01-05 | January 2024 | 13:30 | 2024-01-05 13:05 | 2024-01-05 13:30 | 0:25 | COMPLETED | SEMI-CRITICAL
1022 | 2024-01-06 | Jan 2024 | 23:30 | 23:35 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | HAFIZ | ICU | UNOBLE TO STORE DATA | FOUND SETTINGS HAVE BEEN CHANGED. ENOBLED THE LOCAL STORAGE, NOW WORKING GOOD. | NO | NO | 2024-01-06 | January 2024 | 23:45 | 2024-01-06 23:30 | 2024-01-06 23:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1023 | 2024-01-06 | Jan 2024 | 11:00 | 11:20 | 0.013889 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | RAMKIE | OT | NOT WORKING | SERVICED THE BOARD CONNECTOR AND RECTIFIED | NO | NO | 2024-01-06 | January 2024 | 11:25 | 2024-01-06 11:00 | 2024-01-06 11:25 | 0:25 | COMPLETED | SEMI-CRITICAL
1024 | 2024-01-06 | Jan 2024 | 12:30 | 12:35 | 0.0034722 | DIGITAL PLATFORM SCALE | SECA | 676 | JERLIN | DIALYSIS | POWER CORD NOT WORKING | CHANGED THE CHARGER AND IT'S WORKING GOOD | POWER ADAPTOR | NO | 2024-01-06 | January 2024 | 12:50 | 2024-01-06 12:30 | 2024-01-06 12:50 | 0:20 | COMPLETED | NON-CRITICAL
1025 | 2024-01-06 | Jan 2024 | 10:00 | 10:10 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 550 | JERLIN | CATHLAB | WATER TRAP ISSUE | CHANGED THE WATER TRAP AND RECTIFIED THE ISSUE | WATER TRAP | NO | 2024-01-06 | January 2024 | 10:15 | 2024-01-06 10:00 | 2024-01-06 10:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1026 | 2024-01-07 | Jan 2024 | 11:10 | 11:20 | 0.0069444 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | RAMKIE | OT | OCCLUSION ERROR | CHECKED AND FEED PURPOSE USED, SO AFTER DILUTION, PROBLEM CLEARED | NO | NO | 2024-01-07 | January 2024 | 11:25 | 2024-01-07 11:10 | 2024-01-07 11:25 | 0:15 | COMPLETED | NON-CRITICAL
1027 | 2024-01-08 | Jan 2024 | 23:00 | 23:05 | 0.0034722 | BABY WEIGHING SCALE | SECA | 354 | HAFIZ | ICU | IMPROPER READINGS | CHECKED AND FOUND THAT UNIT WAS CHANGED TO LBS FROM KGS. CHANGES DONE. NOW WORKING GOOD. | NO | NO | 2024-01-08 | January 2024 | 23:10 | 2024-01-08 23:00 | 2024-01-08 23:10 | 0:10 | COMPLETED | NON-CRITICAL
1028 | 2024-01-08 | Jan 2024 | 12:00 | 12:05 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | HAFIZ | ICU | NEEDS TO REPLACE A CARTRIDGE | REMOVED THE OLD CARTRIDGE (S.NO:301759073) AND INSTALLED A NEW CARTRIDGE(S.NO:301759071). MACHINE, WARMING UP. READY TO USE. | CARTRIDGE | NO | 2024-01-08 | January 2024 | 12:20 | 2024-01-08 12:00 | 2024-01-08 12:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1029 | 2024-01-08 | Jan 2024 | 10:10 | 10:20 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | RAMKIE | OT | CIRCUIT LEAK TEST FAILED | CHECKED AND FOUND SODA LIME CHAMBER FULLY FILLED, AFTER ADJUSTED PROBLEM CLEARED | NO | NO | 2024-01-08 | January 2024 | 10:30 | 2024-01-08 10:10 | 2024-01-08 10:30 | 0:20 | COMPLETED | CRITICAL
1030 | 2024-01-08 | Jan 2024 | 15:00 | 15:15 | 0.010417 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | WARD | SIDE RAIL NOT WORKING | FIXED IT AND RECTIFIED THE ISSUE | NO | NO | 2024-01-08 | January 2024 | 15:25 | 2024-01-08 15:00 | 2024-01-08 15:25 | 0:25 | COMPLETED | NON-CRITICAL
1031 | 2024-01-09 | Jan 2024 | 16:00 | 16:05 | 0.0034722 | RADIANT HEAT WARMER | WIPRO GE HEALTH CARE | GIRAFFE | HAFIZ | ICU | NOT WORKING | CHECKED AND FOUND CIRCUIT BREAKER TURNED OFF. SWITCHED ON AND NOW WORKING GOOD. | NO | NO | 2024-01-09 | January 2024 | 16:10 | 2024-01-09 16:00 | 2024-01-09 16:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1032 | 2024-01-09 | Jan 2024 | 17:20 | 17:25 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | HAFIZ | OT | O2 CELL ISSUE | FOUND MOISTURE IN FLOW SENSOR MODULE, CLEANED AND REINSTALLED O2 CELL. WEEKLY O2 CELL CALIBRATION DONE. NOW WORKING GOOD. | NO | NO | 2024-01-09 | January 2024 | 18:00 | 2024-01-09 17:20 | 2024-01-09 18:00 | 0:40 | COMPLETED | CRITICAL
1033 | 2024-01-09 | Jan 2024 | 18:00 | 18:03 | 0.0020833 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | HAFIZ | ICU | NOT WORKING | REPLACED THE CONTROL BOARD FROM ANOTHER PUMP AND NOW WORKING GOOD. | CONTROL BOARD | NO | 2024-01-09 | January 2024 | 18:30 | 2024-01-09 18:00 | 2024-01-09 18:30 | 0:30 | COMPLETED | NON-CRITICAL
1034 | 2024-01-09 | Jan 2024 | 7:50 | 7:55 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | SOORAJ | OT | SATURATION NOT SHOWING | CHECKED AND FOUND ISSUES WITH SPO2 PROBE. CHANGED SPO2 PROBE AND ISSUE RECTIFIED. | SPO2 PROBE | NO | 2024-01-09 | January 2024 | 8:00 | 2024-01-09 07:50 | 2024-01-09 08:00 | 0:10 | COMPLETED | SEMI-CRITICAL
1035 | 2024-01-09 | Jan 2024 | 13:55 | 14:00 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | O2 CELL ISSUE | CHECKED AND FOUND O2 CELL DEFECTIVE. O2 CELL FROM ANOTHER MACHINE BME/00149 REPLACED IN TO PROCEED WITH THE CASE | O2 CELL | NO | 2024-01-09 | January 2024 | 14:10 | 2024-01-09 13:55 | 2024-01-09 14:10 | 0:15 | COMPLETED | CRITICAL
1036 | 2024-01-09 | Jan 2024 | 10:10 | 10:20 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | OT | NOT WORKING | CHECKED AND FOUND THE POWER CORD DISCONNECTED, NOW PROBLEM RECTIFIED | NO | NO | 2024-01-09 | January 2024 | 10:30 | 2024-01-09 10:10 | 2024-01-09 10:30 | 0:20 | COMPLETED | NON-CRITICAL
1037 | 2024-01-10 | Jan 2024 | 12:30 | 12:35 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | CVP DUE | CVP TESTS DONE AND PASSED. NOW, MACHINE READY TO USE. | CARTRIDGE | NO | 2024-01-10 | January 2024 | 13:00 | 2024-01-10 12:30 | 2024-01-10 13:00 | 0:30 | COMPLETED | SEMI-CRITICAL
1038 | 2024-01-10 | Jan 2024 | 13:55 | 14:00 | 0.0034722 | IMMUNOASSY AUTO ANOLYZER | BIOMERIUX | MINIVIDAS | HAFIZ | LABORATORY | BARCODE ERROR | REMOVED THE BARCODE CABLE, CLEANED AND REFIXED. NOW, WORKING GOOD. | NO | NO | 2024-01-10 | January 2024 | 14:10 | 2024-01-10 13:55 | 2024-01-10 14:10 | 0:15 | COMPLETED | SEMI-CRITICAL
1039 | 2024-01-10 | Jan 2024 | 17:30 | 17:35 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID E95 | HAFIZ | CATHLAB | TEE PROBE NOT DETECTING | FOUND THAT IT WAS CONNECTED PROPERY. REFIXED AND ISSUE RECTIFIED. | NO | NO | 2024-01-10 | January 2024 | 17:46 | 2024-01-10 17:30 | 2024-01-10 17:46 | 0:16 | COMPLETED | SEMI-CRITICAL
1040 | 2024-01-10 | Jan 2024 | 11:05 | 11:10 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SOORAJ | ICU | PRINTER NOT WORKING | CLEANED PRINTER HEAD AND TEST PRINT TAKEN, ISSUE RECTIFIED. | NO | NO | 2024-01-10 | January 2024 | 11:20 | 2024-01-10 11:05 | 2024-01-10 11:20 | 0:15 | COMPLETED | SEMI-CRITICAL
1041 | 2024-01-10 | Jan 2024 | 13:40 | 13:45 | 0.0034722 | C-ARM MACHINE | WIPRO GE HEALTH CARE | OEC ONE | SOORAJ | OT | TUBE TOO HOT ALARM | CHECKED AND FOUND WATER SPILLAGE ON TUBE BOARDS. TUBE REPLACED AND ISSUE RECTIFIED. | TUBE KIT | NO | 2024-01-11 | January 2024 | 19:00 | 2024-01-10 13:40 | 2024-01-11 19:00 | 29:20 | COMPLETED | CRITICAL
1042 | 2024-01-10 | Jan 2024 | 14:15 | 14:20 | 0.0034722 | MICROSCOPE OPERATING, SURGICAL | LEICA MICROSYSTEMS | LEICA OHX | SOORAJ | OT | IMAGE NOT SHOWING ON EXTERNOL DISPLAY | FOUND ISSUE WITH SOFTWARE CRASH. REBOOTED THE RECORDER ON RECOVERY MODE AND ISSUE RECTIFIED. | NO | NO | 2024-01-10 | January 2024 | 14:30 | 2024-01-10 14:15 | 2024-01-10 14:30 | 0:15 | COMPLETED | SEMI-CRITICAL
1043 | 2024-01-10 | Jan 2024 | 15:05 | 15:10 | 0.0034722 | NAVIGATION SYSTEM | STRYKER | MAC3 | THANISH | OT | UPS FAILED | BYPASSED THE UPS SUPPLY FROM CONSOLE.GIVEN POWER SUPPLY DIRECTLY TO CONSOLE. | NO | NO | 2024-01-10 | January 2024 | 15:50 | 2024-01-10 15:05 | 2024-01-10 15:50 | 0:45 | COMPLETED | CRITICAL
1044 | 2024-01-10 | Jan 2024 | 9:10 | 9:20 | 0.0069444 | ACT MACHINE | HELENO LABORATORIES | ACTALYKE | RAMKIE | CATHLAB | ERROR PO1 | COMPANY PERSON CHECKED AND CALIBRATED | NO | NO | 2024-01-10 | January 2024 | 9:40 | 2024-01-10 09:10 | 2024-01-10 09:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1045 | 2024-01-10 | Jan 2024 | 9:40 | 9:45 | 0.0034722 | ACT MACHINE | HELENO LABORATORIES | ACTALYKE | RAMKIE | CATHLAB | ERROR TAKES MORE TIME | COMPANY PERSON CHECKED AND CALIBRATED | NO | NO | 2024-01-10 | January 2024 | 10:10 | 2024-01-10 09:40 | 2024-01-10 10:10 | 0:30 | COMPLETED | SEMI-CRITICAL
1046 | 2024-01-10 | Jan 2024 | 10:10 | 10:15 | 0.0034722 | ACT MACHINE | HELENO LABORATORIES | ACTALYKE | RAMKIE | OT | ERROR TAKES MORE TIME | COMPANY PERSON CHECKED AND CALIBRATED | NO | NO | 2024-01-10 | January 2024 | 10:30 | 2024-01-10 10:10 | 2024-01-10 10:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1047 | 2024-01-10 | Jan 2024 | 21:10 | 21:20 | 0.0069444 | VENTILATOR, NEONATAL | WIPRO GE HEALTH CARE | SLE 5000 | JERLIN | ICU | NOT SHOWING WAVEFORM | CHECKED THE SETTINGS AND IT'S WORKING FINE. | NO | NO | 2024-01-10 | January 2024 | 21:30 | 2024-01-10 21:10 | 2024-01-10 21:30 | 0:20 | COMPLETED | CRITICAL
1048 | 2024-01-10 | Jan 2024 | 13:30 | 13:40 | 0.0069444 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | JERLIN | WARD | BULB ELECTRODE TORN | CHANGED A BULB ELECTRODE AND RECTIFIED THE ISSUE. | BULB ELECTRODE | NO | 2024-01-10 | January 2024 | 13:50 | 2024-01-10 13:30 | 2024-01-10 13:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1049 | 2024-01-11 | Jan 2024 | 14:30 | 14:35 | 0.0034722 | RADIANT HEAT WARMER CUM INCUBATOR | WIPRO GE HEALTH CARE | GIRAFFE | HAFIZ | ICU | FILTER TO BE REPLACED | NEW FILTERS INSTALLED IN TWO GIRAFFE INCUBATORS (BME/), (BME/). NOW READY TO USE. | GIRAFFE FILTERS | NO | 2024-01-11 | January 2024 | 15:00 | 2024-01-11 14:30 | 2024-01-11 15:00 | 0:30 | COMPLETED | SEMI-CRITICAL
1050 | 2024-01-11 | Jan 2024 | 15:55 | 16:00 | 0.0034722 | BMI MACHINE | AKAS MEDICAL | QVS-100 | HAFIZ | OPD | NOT WORKING | FOUND ISSUE WITH THE POWER SOCKET. PLUGGED TO ANOTHER SOCKET AND WORKING GOOD. | NO | NO | 2024-01-11 | January 2024 | 16:05 | 2024-01-11 15:55 | 2024-01-11 16:05 | 0:10 | COMPLETED | NON-CRITICAL
1051 | 2024-01-11 | Jan 2024 | 7:40 | 7:50 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SOORAJ | ICU | SWITHCED OFF | FOUND UPS CONNEECTION CUT OFF. RESTARTED THE SYSTEM AND RAN IQM PROCESS, ISSUE RECTIFIED. | NO | NO | 2024-01-11 | January 2024 | 8:00 | 2024-01-11 07:40 | 2024-01-11 08:00 | 0:20 | COMPLETED | SEMI-CRITICAL
1052 | 2024-01-11 | Jan 2024 | 7:30 | 7:35 | 0.0034722 | NAVIGATION SYSTEM | STRYKER | MAC3 | SOORAJ | OT | EXTERNAL DISPLAY NOT SHOWING | FOUND ISSUE WITH UPS. SEPERATE POWER SOURCE CABLE CONNECTED TO THE MONITOR, ISSUE RECTIFIED. | NO | NO | 2024-01-11 | January 2024 | 16:00 | 2024-01-11 07:30 | 2024-01-11 16:00 | 8:30 | COMPLETED | CRITICAL
1053 | 2024-01-11 | Jan 2024 | 14:10 | 14:20 | 0.0069444 | C-ARM MACHINE | WIPRO GE HEALTH CARE | OEC ONE | THANISH | ENDOSCOPY | TUBE OVER HEATED | REPLACED THE TUBE AND PERFORMED THE CALIBRATION.WORKING FINE | TUBE KIT | NO | 2024-01-11 | January 2024 | 19:30 | 2024-01-11 14:10 | 2024-01-11 19:30 | 5:20 | COMPLETED | CRITICAL
1054 | 2024-01-12 | Jan 2024 | 13:00 | 13:04 | 0.0027778 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | WARD | NOT WORKING | MANUAL KEYS LOCKED, UNLOCKED AND NOW WORKING GOOD. | NO | NO | 2024-01-12 | January 2024 | 13:15 | 2024-01-12 13:00 | 2024-01-12 13:15 | 0:15 | COMPLETED | NON-CRITICAL
1055 | 2024-01-12 | Jan 2024 | 15:00 | 15:10 | 0.0069444 | C-ARM MACHINE | WIPRO GE HEALTH CARE | OEC ONE | HAFIZ | OT | CAMERA NOT WORKING | RECEIVED THE CAMERA MODULE FROM GE, CHANGED AND CALIBRATED BY THE COMPANY PERSON AND WORKING GOOD. | CAMERA | NO | 2024-01-12 | January 2024 | 20:00 | 2024-01-12 15:00 | 2024-01-12 20:00 | 5:00 | COMPLETED | CRITICAL
1056 | 2024-01-12 | Jan 2024 | 9:20 | 9:25 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | PPEAK NOT SHOWING | REFIXED BELLOWS AND CHECKOUT DONE, ISSUE RECTIFIED. | NO | NO | 2024-01-12 | January 2024 | 9:30 | 2024-01-12 09:20 | 2024-01-12 09:30 | 0:10 | COMPLETED | CRITICAL
1057 | 2024-01-12 | Jan 2024 | 16:15 | 16:20 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | O2 CELL ISSUE | NEW O2 CELL REPLACED AND ISSUE RECTIFIED. | O2 CELL | NO | 2024-01-12 | January 2024 | 16:25 | 2024-01-12 16:15 | 2024-01-12 16:25 | 0:10 | COMPLETED | CRITICAL
1058 | 2024-01-12 | Jan 2024 | 9:05 | 9:10 | 0.0034722 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | THANISH | OPD | TMT NOT WORKING | UNLOCKED THE EMERGENCY SWITCH.WORKING FINE. | NO | NO | 2024-01-12 | January 2024 | 9:15 | 2024-01-12 09:05 | 2024-01-12 09:15 | 0:10 | COMPLETED | NON-CRITICAL
1059 | 2024-01-12 | Jan 2024 | 10:20 | 10:25 | 0.0034722 | BMI MACHINE | AKAS MEDICAL | QVS-100 | THANISH | OPD | NOT SWITCHING ON | OLD POWER CORD FAILED/REPLCAED WITH NEW POWER CORD. | POWER CORD | NO | 2024-01-12 | January 2024 | 10:40 | 2024-01-12 10:20 | 2024-01-12 10:40 | 0:20 | COMPLETED | NON-CRITICAL
1060 | 2024-01-12 | Jan 2024 | 1:00 | 1:06 | 0.0041667 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | JERLIN | EMERGENCY | NOT ABLE TO TAKE PRINT | BATTERY WAS TOO LOW , CONNECTED TO AC SOURCE AND TEST DONE. | NO | NO | 2024-01-12 | January 2024 | 1:15 | 2024-01-12 01:00 | 2024-01-12 01:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1061 | 2024-01-13 | Jan 2024 | 13:10 | 13:15 | 0.0034722 | OT TABLE | MAQUET | ALPHA MAX | SOORAJ | OT | BURNING SMELL FROM POWER CHORD | FOUND SALINE SOLUTION SPILLAGE ON POWER CHORD, CHANGED POWER CHORD AND FOUND WORKING FINE. | NO | NO | 2024-01-13 | January 2024 | 13:30 | 2024-01-13 13:10 | 2024-01-13 13:30 | 0:20 | COMPLETED | NON-CRITICAL
1062 | 2024-01-13 | Jan 2024 | 13:55 | 14:00 | 0.0034722 | MICROSCOPE OPERATING, SURGICAL | CARL ZEISS | OPMI SENSARA | SOORAJ | OT | ENT MICROSCOPE ARM SUSPENDING AUTOMATICALY | TIGHTENED THE ARM SCREWS, ISSUE RECTIFIED. | NO | NO | 2024-01-13 | January 2024 | 14:10 | 2024-01-13 13:55 | 2024-01-13 14:10 | 0:15 | COMPLETED | SEMI-CRITICAL
1063 | 2024-01-13 | Jan 2024 | 11:10 | 11:30 | 0.013889 | EXAMINATION COUCH | MIDMARK JANAK | E122012 | RAMKIE | OPD | NOT WORKING | NEED TO REPLACE THE ACTUATOR | NO | ACTUATOR | 2024-01-13 | January 2024 | 11:50 | 2024-01-13 11:10 | 2024-01-13 11:50 | 0:40 | PENDING | NON-CRITICAL
1064 | 2024-01-13 | Jan 2024 | 11:50 | 11:55 | 0.0034722 | EXAMINATION COUCH | MIDMARK JANAK | E122012 | RAMKIE | OPD | NOT WORKING | NEED TO REPLACE THE GAS PISTON AND REMOTE | NO | GAS PISTON AND REMOTE | 2024-01-13 | January 2024 | 12:20 | 2024-01-13 11:50 | 2024-01-13 12:20 | 0:30 | PENDING | NON-CRITICAL
1065 | 2024-01-13 | Jan 2024 | 2:00 | 2:10 | 0.0069444 | VENTILATOR, NEONATAL | MAQUET | SERVO-I | JERLIN | ICU | NOT ABLE TO PERFORM PRE USE CHECK | PERFORMED PREUSE CHECK AND RECTIFIED THE ISSUE | NO | NO | 2024-01-13 | January 2024 | 2:20 | 2024-01-13 02:00 | 2024-01-13 02:20 | 0:20 | COMPLETED | CRITICAL
1066 | 2024-01-15 | Jan 2024 | 9:45 | 9:45 | 0 | PATIENT MONITOR | PHILIPS | CM12 | SOORAJ | WARD | NIBP NOT SHOWING | FOUND BP HOSE DEFECTIVE. ALTERNOTE HOSE GIVEN AND ISSUE RECTIFIED. | NIBP HOSE | NO | 2024-01-15 | January 2024 | 9:55 | 2024-01-15 09:45 | 2024-01-15 09:55 | 0:10 | COMPLETED | SEMI-CRITICAL
1067 | 2024-01-15 | Jan 2024 | 3:30 | 3:40 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | WARD | NOT WORKING | CHANGED THE POWER CORD AND RECTIFIED THE ISSUE | NO | NO | 2024-01-15 | January 2024 | 3:55 | 2024-01-15 03:30 | 2024-01-15 03:55 | 0:25 | COMPLETED | NON-CRITICAL
1068 | 2024-01-16 | Jan 2024 | 7:20 | 7:25 | 0.0034722 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX 8 | HAFIZ | OT | NOT WORKING | RESET DONE, AND NOW WORKING GOOD | NO | NO | 2024-01-16 | January 2024 | 7:30 | 2024-01-16 07:20 | 2024-01-16 07:30 | 0:10 | COMPLETED | SEMI-CRITICAL
1069 | 2024-01-16 | Jan 2024 | 7:30 | 7:31 | 0.00069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | HAFIZ | OT | CALIBRATION DUE | CALIBRATION DONE, NOW READY TO USE. | NO | NO | 2024-01-16 | January 2024 | 8:00 | 2024-01-16 07:30 | 2024-01-16 08:00 | 0:30 | COMPLETED | CRITICAL
1070 | 2024-01-16 | Jan 2024 | 8:30 | 8:35 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | HAFIZ | ENDOSCOPY | CIRCUIT LEAK TEST FAILED | ISSUES WITH THE BELLOW. REMOVED IT AND CLEANED. RERUNNED THE PRE-USE CHECK AND NOW READY TO USE. | NO | NO | 2024-01-16 | January 2024 | 9:00 | 2024-01-16 08:30 | 2024-01-16 09:00 | 0:30 | COMPLETED | CRITICAL
1071 | 2024-01-16 | Jan 2024 | 9:00 | 9:00 | 0 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | HAFIZ | ENDOSCOPY | CALIBRATION DUE | WEEKLY AND MONTHLY O2 CELL CALIBRATION DONE, NOW READY TO USE. | NO | NO | 2024-01-16 | January 2024 | 9:20 | 2024-01-16 09:00 | 2024-01-16 09:20 | 0:20 | COMPLETED | CRITICAL
1072 | 2024-01-16 | Jan 2024 | 10:30 | 10:35 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | ICU | DOOR BROKEN | SWAPPED THE DOOR FROM DEFECTIVE INFUSION PUMP, NOW WORKING GOOD. | NO | NO | 2024-01-16 | January 2024 | 11:00 | 2024-01-16 10:30 | 2024-01-16 11:00 | 0:30 | COMPLETED | NON-CRITICAL
1073 | 2024-01-16 | Jan 2024 | 13:15 | 13:18 | 0.0020833 | CORE CONSOLE WITH HANDPIECE | STRYKER | 5400-052-000 | HAFIZ | OT | NOT WORKING | RESET DONE, AND NOW WORKING GOOD | NO | NO | 2024-01-16 | January 2024 | 13:30 | 2024-01-16 13:15 | 2024-01-16 13:30 | 0:15 | COMPLETED | SEMI-CRITICAL
1074 | 2024-01-16 | Jan 2024 | 13:55 | 14:00 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | HAFIZ | EMERGENCY | NOT SWITCHING ON | FOUND THAT CIRCUIT BREAKER TURNED OFF. SWITCHED ON AND NOW WORKING GOOD. | NO | NO | 2024-01-16 | January 2024 | 14:05 | 2024-01-16 13:55 | 2024-01-16 14:05 | 0:10 | COMPLETED | CRITICAL
1075 | 2024-01-16 | Jan 2024 | 18:10 | 18:20 | 0.0069444 | PACEMAKER,SINGLE CHAMBER | MEDTRONIC | 53401 | RAMKIE | ICU | LOW BATTERY ALARM | CHECKED AND REPLACED BATTERY AND PROBLEM RECTIFIED | BATTERY | NO | 2024-01-16 | January 2024 | 18:30 | 2024-01-16 18:10 | 2024-01-16 18:30 | 0:20 | COMPLETED | CRITICAL
1076 | 2024-01-17 | Jan 2024 | 12:00 | 12:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | HAFIZ | OT | AUTOMATICALLY SWITCHED OFF | FOUND FUNGI FORMED IN IV2FLEX BOARD. REPLACED IT AND NOW WORKING GOOD. | INTERFACE BOARD | NO | 2024-01-17 | January 2024 | 15:00 | 2024-01-17 12:00 | 2024-01-17 15:00 | 3:00 | COMPLETED | SEMI-CRITICAL
1077 | 2024-01-17 | Jan 2024 | 19:10 | 19:15 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | THANISH | ICU | ECG NOT WORKING | REPAIRED THE 5-LEAD ECG CABLE.WORKING FINE. | NO | NO | 2024-01-17 | January 2024 | 19:35 | 2024-01-17 19:10 | 2024-01-17 19:35 | 0:25 | COMPLETED | SEMI-CRITICAL
1078 | 2024-01-17 | Jan 2024 | 10:10 | 10:20 | 0.0069444 | OPTICAL COHERENCE TOMOGRAPHY, 3 DIMENSIONOL | TOPCON | 3D OCT-1 | RAMKIE | OPD | BLUR IN CERTAIN PART | COMPANY PERSON CHECKED AND FOUND THAT THE OBJECT SCRATCHES, INTERNOL FIXATION TARGET IS WEAK AND LCD TO BE REPLACED. WORKING WITH CONTRAINTS | NO | TARGER FIXATION, OBJEST LEN AND LCD MONITOR | 2024-01-17 | January 2024 | 11:00 | 2024-01-17 10:10 | 2024-01-17 11:00 | 0:50 | PENDING | SEMI-CRITICAL
1079 | 2024-01-17 | Jan 2024 | 11:15 | 11:20 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | JERLIN | OT | NOT WORKING | CONNECTED THE TEMP SENSOR PROPERLY AND RECTIFIED THE ISSUE. | NO | NO | 2024-01-17 | January 2024 | 11:30 | 2024-01-17 11:15 | 2024-01-17 11:30 | 0:15 | COMPLETED | SEMI-CRITICAL
1080 | 2024-01-18 | Jan 2024 | 12:00 | 12:04 | 0.0027778 | CAVITRON ULTRASONIC SURGICAL ASPIRATOR | INTEGRA | CUSA EXCEL + | HAFIZ | OT | ALARM INDICATION | FOUND IMPURITIES IN DISTILLED WATER, REPLACED WITH FRESH WATER AND NOW ISSUE RECTIFIED | NO | NO | 2024-01-18 | January 2024 | 12:20 | 2024-01-18 12:00 | 2024-01-18 12:20 | 0:20 | COMPLETED | CRITICAL
1081 | 2024-01-18 | Jan 2024 | 11:25 | 11:30 | 0.0034722 | RADIOGRAPHIC UNIT | CARESTREAM | DRF ASCEND ANOLOG HF | SOORAJ | OPD | PATIENT DATA NOT SHOWING ON SOFTWARE. | RECONNETCED THE LAN CABLE AND RESTARTED THE SYSTEM. ISSUE RECTIFIED | NO | NO | 2024-01-18 | January 2024 | 11:45 | 2024-01-18 11:25 | 2024-01-18 11:45 | 0:20 | COMPLETED | SEMI-CRITICAL
1082 | 2024-01-18 | Jan 2024 | 12:00 | 12:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | SOORAJ | WARD | ALARM SOUND HIGH | SETTINGS CONFIGURED AND ISSUE RECTIFIED. | NO | NO | 2024-01-18 | January 2024 | 12:10 | 2024-01-18 12:00 | 2024-01-18 12:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1083 | 2024-01-18 | Jan 2024 | 13:00 | 13:00 | 0 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SOORAJ | WARD | PRINTER NOT WORKING. | REFIXED PAPER AND TEST PRINT PASSED, ISSUE RECTIFIED. | NO | NO | 2024-01-18 | January 2024 | 13:05 | 2024-01-18 13:00 | 2024-01-18 13:05 | 0:05 | COMPLETED | CRITICAL
1084 | 2024-01-18 | Jan 2024 | 13:10 | 13:10 | 0 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SOORAJ | ICU | PRINTER NOT WORKING. | REFIXED PAPER AND TEST PRINT PASSED, ISSUE RECTIFIED. | NO | NO | 2024-01-18 | January 2024 | 13:15 | 2024-01-18 13:10 | 2024-01-18 13:15 | 0:05 | COMPLETED | CRITICAL
1085 | 2024-01-19 | Jan 2024 | 8:30 | 8:35 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | WARD | NIBP NOT WORKING | ADJUSTED THE WELCROW AND CHECKED, NOW WORKNG GOOD | NO | NO | 2024-01-19 | January 2024 | 8:45 | 2024-01-19 08:30 | 2024-01-19 08:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1086 | 2024-01-19 | Jan 2024 | 9:35 | 9:37 | 0.0013889 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | ICU | PRINTER NOT WORKING | CLEANED THE PRINTER HEAD. REINSERTED THE PAPER ROLL. NOW WORKING GOOD. | NO | NO | 2024-01-19 | January 2024 | 9:55 | 2024-01-19 09:35 | 2024-01-19 09:55 | 0:20 | COMPLETED | CRITICAL
1087 | 2024-01-19 | Jan 2024 | 10:00 | 10:05 | 0.0034722 | RADIANT HEAT WARMER | WIPRO GE HEALTH CARE | GIRAFFE | HAFIZ | ICU | PROBE NOT WORKING | CHECKED AND ADJUSTED THE TEMPERATURE PROBE. NOW READY TO USE. | NO | NO | 2024-01-19 | January 2024 | 10:20 | 2024-01-19 10:00 | 2024-01-19 10:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1088 | 2024-01-19 | Jan 2024 | 21:05 | 21:10 | 0.0034722 | VENTILATOR, UNIVERSAL | MAQUET | SERVO-I | THANISH | ICU | EXPIRATORY CASETTE NOT WORKING | REPLACED THE MEMEBRANE FOR EXPIRATORY CASETTE.WORKING FINE. | EXPIRATORY CASSETTE | NO | 2024-01-19 | January 2024 | 21:40 | 2024-01-19 21:05 | 2024-01-19 21:40 | 0:35 | COMPLETED | CRITICAL
1089 | 2024-01-19 | Jan 2024 | 22:10 | 22:15 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | THANISH | ICU | NEED TO INSTALL NEW ABG CATRIDGE | INSTALLED THE NEW ABG CATRIDGE AND CVP TEST PERFORMED.WORKING FINE. | CARTRIDGE | NO | 2024-01-19 | January 2024 | 23:05 | 2024-01-19 22:10 | 2024-01-19 23:05 | 0:55 | COMPLETED | SEMI-CRITICAL
1090 | 2024-01-19 | Jan 2024 | 22:10 | 22:15 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | THANISH | ICU | NEED TO INSTALL NEW ABG CATRIDGE | INSTALLED THE NEW ABG CATRIDGE AND CVP TEST PERFORMED.WORKING FINE. | CARTRIDGE | NO | 2024-01-19 | January 2024 | 23:05 | 2024-01-19 22:10 | 2024-01-19 23:05 | 0:55 | COMPLETED | SEMI-CRITICAL
1091 | 2024-01-19 | Jan 2024 | 16:30 | 16:40 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 450 | JERLIN | ICU | PEDIATRIC TEMPERATURE PROBE NOT GIVEN | REPLACEMENT PROVIDE AND CHECKED | PEDIATRIC TEMP PROBE | NO | 2024-01-19 | January 2024 | 16:50 | 2024-01-19 16:30 | 2024-01-19 16:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1092 | 2024-01-19 | Jan 2024 | 15:00 | 15:15 | 0.010417 | OXYGEN THERAPY UNIT | INSPIRED | O2FLO | JERLIN | ICU | CHECKOUT NOT PASSED | CLEANED THE FILTER AND NOW ITS WORKING GOOD | NO | NO | 2024-01-19 | January 2024 | 15:30 | 2024-01-19 15:00 | 2024-01-19 15:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1093 | 2024-01-20 | Jan 2024 | 8:25 | 8:30 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | FOUND MOISTURE IN ABSORBER CANISTER. CLEANED IT AND REFIXED. RERUNNED THE TEST AND NOW CIRCUIT LEAK UNDER 100 ML/MIN. NOW READY TO USE. | NO | NO | 2024-01-20 | January 2024 | 8:45 | 2024-01-20 08:25 | 2024-01-20 08:45 | 0:20 | COMPLETED | CRITICAL
1094 | 2024-01-20 | Jan 2024 | 8:45 | 8:47 | 0.0013889 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | ICU | EQUIPMENT DISABLED, SYSTEM FAILURE | RESET DONE, RERUNNED THE OPERATIONAL CHECK. MACHINE READY TO USE. | NO | NO | 2024-01-20 | January 2024 | 9:00 | 2024-01-20 08:45 | 2024-01-20 09:00 | 0:15 | COMPLETED | CRITICAL
1095 | 2024-01-20 | Jan 2024 | 11:30 | 11:35 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | WARD | SPO2 PROBE NOT WORKING | CHECKED AND FOUND NO ISSUES WITH THE PROBE. RECONNECTED THE PROBE PROPERLY AT BOTH ENDS. NOW WORKING GOOD. | NO | NO | 2024-01-20 | January 2024 | 11:45 | 2024-01-20 11:30 | 2024-01-20 11:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1096 | 2024-01-20 | Jan 2024 | 12:50 | 12:55 | 0.0034722 | FETAL DOPPLER | BPL | FD9713N | HAFIZ | WARD | DISPLAY ISSUE | RESET DONE AND NOW READY TO USE. | NO | NO | 2024-01-20 | January 2024 | 13:10 | 2024-01-20 12:50 | 2024-01-20 13:10 | 0:20 | COMPLETED | SEMI-CRITICAL
1097 | 2024-01-20 | Jan 2024 | 15:30 | 15:40 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | HAFIZ | OT | O2 CELL ISSUE | FOUND MOISTURE IN O2 CELL MODULE, DRIED AND SWAPPED O2 CELL FROM 9100C NXT (/00214) AND AVANCE CS2 (/00142). NOW BOTH READY TO USE. | NO | NO | 2024-01-20 | January 2024 | 16:30 | 2024-01-20 15:30 | 2024-01-20 16:30 | 1:00 | COMPLETED | CRITICAL
1098 | 2024-01-20 | Jan 2024 | 20:10 | 20:15 | 0.0034722 | RADIANT HEAT WARMER | WIPRO GE HEALTH CARE | LULLABY | THANISH | ICU | WARMER NOT WORKING | REPLACED THE HEATER COIL BY THE COMPANY PERSON. | HEATER COIL | NO | 2024-01-20 | January 2024 | 21:30 | 2024-01-20 20:10 | 2024-01-20 21:30 | 1:20 | COMPLETED | SEMI-CRITICAL
1099 | 2024-01-20 | Jan 2024 | 21:10 | 21:15 | 0.0034722 | RADIANT HEAT WARMER | WIPRO GE HEALTH CARE | GIRAFFE | THANISH | ICU | SKIN TEMPERATURE PROBE FAILED | REPLACED THE SKIN TEMPERATURE PROBE | SKIN TEMPERATURE PROBE | NO | 2024-01-20 | January 2024 | 21:30 | 2024-01-20 21:10 | 2024-01-20 21:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1100 | 2024-01-20 | Jan 2024 | 11:10 | 11:20 | 0.0069444 | MULTIFUNCTIONAL BED | PARAMOUNT | A5 4M IA 54380 | RAMKIE | WARD | NOT WORKING | CHECKED, REMOTE GOT LOCKED, SO RESET DONE AND NOW IT WORKING FINE | NO | NO | 2024-01-20 | January 2024 | 11:35 | 2024-01-20 11:10 | 2024-01-20 11:35 | 0:25 | COMPLETED | NON-CRITICAL
1101 | 2024-01-21 | Jan 2024 | 0:20 | 0:24 | 0.0027778 | BIPAP MACHINE | RESMED | STELLAR 150 | HAFIZ | WARD | NOT WORKING | FOUND THAT BATTERY FULLY DEPLETED. PLUGGED TO AC SOURCE, SETTINGS DONE. AND NOW WORKING GOOD. | NO | NO | 2024-01-21 | January 2024 | 0:30 | 2024-01-21 00:20 | 2024-01-21 00:30 | 0:10 | COMPLETED | SEMI-CRITICAL
1102 | 2024-01-21 | Jan 2024 | 1:00 | 1:03 | 0.0020833 | BIPAP MACHINE | PHILIPS | A40 | HAFIZ | ICU | NOT WORKING | SWAPPED THE STANDBY BIPAP MACHINE (BME/01110) WITH THIS BIPAP MACHINE. | NO | NO | 2024-01-21 | January 2024 | 1:15 | 2024-01-21 01:00 | 2024-01-21 01:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1103 | 2024-01-21 | Jan 2024 | 2:00 | 2:05 | 0.0034722 | USG MACHINE | MINDRAY MEDICAL INDIA | MX7 | HAFIZ | WARD | NOT WORKING | FOUND THAT PROBE NOT CONNECTED PROPERLY. RECONNECTED AND NOW WORKING GOOD | NO | NO | 2024-01-21 | January 2024 | 2:10 | 2024-01-21 02:00 | 2024-01-21 02:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1104 | 2024-01-22 | Jan 2024 | 11:30 | 11:40 | 0.0069444 | ELECTRICAL SUCTION | ATMOS | S351 | RAMKIE | ICU | GLASS JAR BROKEN | CHECKED AND FOUND THAT JAR BROKEN, UNOBLE TO FIX, SO EXTERNOL JAR WAS PROVIDED, NOW IT WORKING | NO | NO | 2024-01-22 | January 2024 | 11:45 | 2024-01-22 11:30 | 2024-01-22 11:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1105 | 2024-01-23 | Jan 2024 | 20:35 | 20:40 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | WARD | COT NOT WORKING | FOUND BATTERY DEPLETED. CONNECTED ON MAINS AND ISSUE RECTIFIED. | NO | NO | 2024-01-23 | January 2024 | 20:45 | 2024-01-23 20:35 | 2024-01-23 20:45 | 0:10 | COMPLETED | NON-CRITICAL
1106 | 2024-01-23 | Jan 2024 | 12:20 | 12:30 | 0.0069444 | IVUS MACHINE | PHILIPS | CORE SERIES S5 | RAMKIE | CATHLAB | ERROR AND NO DATA AFTER INSERTING CATHETER TO PATIENT | COMPANY CHECKED AND ORDERED FFR MODULE, SO REPLACED IT AND NOW IT WORKING FINE | NO | FFR MODULE | 2024-01-23 | January 2024 | 12:45 | 2024-01-23 12:20 | 2024-01-23 12:45 | 0:25 | PENDING | SEMI-CRITICAL
1107 | 2024-01-23 | Jan 2024 | 11:40 | 11:45 | 0.0034722 | MICROSCOPE OPERATING, SURGICAL | CARL ZEISS | OPMI SENSARA | JERLIN | OT | NOT ABLE TO USE SLAVE MONITOR | CHANGED SETTINGS AND CONNECTED TO THE SLAVE MONITOR | NO | NO | 2024-01-23 | January 2024 | 11:55 | 2024-01-23 11:40 | 2024-01-23 11:55 | 0:15 | COMPLETED | SEMI-CRITICAL
1108 | 2024-01-24 | Jan 2024 | 10:00 | 10:04 | 0.0027778 | RADIANT HEAT WARMER | WIPRO GE HEALTH CARE | GIRAFFE | HAFIZ | ICU | TEMP PROBE NOT WORKING | REPLACED THE NEW TEMPERATURE WITH THE OLD ONE. NOW WORKING GOOD. | SKIN TEMPERATURE PROBE | NO | 2024-01-24 | January 2024 | 10:15 | 2024-01-24 10:00 | 2024-01-24 10:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1109 | 2024-01-24 | Jan 2024 | 13:05 | 13:10 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | THANISH | WARD | SPO2 PROBE NOT WORKING | OLD SPO2 PROBE PROBE FAILED/REPLACED WITH NEW ONE.WORKING FINE. | SPO2 PROBE | NO | 2024-01-24 | January 2024 | 13:30 | 2024-01-24 13:05 | 2024-01-24 13:30 | 0:25 | COMPLETED | SEMI-CRITICAL
1110 | 2024-01-24 | Jan 2024 | 16:10 | 16:15 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | THANISH | WARD | ECG BULB MISSING | REPLACED WITH NEW ECG BULB.WORKING FINE. | BULB ELECTRODE | NO | 2024-01-24 | January 2024 | 16:30 | 2024-01-24 16:10 | 2024-01-24 16:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1111 | 2024-01-24 | Jan 2024 | 12:00 | 12:10 | 0.0069444 | MICRO CENTRIFUGE | NEUATION | I FUGE D12 | RAMKIE | LABORATORY | LED DISPLAY NOT GOOD | CHECKED AND RECONNECTED, NOT IT IS GOOD, UNDER OBSERVATION | NO | NO | 2024-01-24 | January 2024 | 12:35 | 2024-01-24 12:00 | 2024-01-24 12:35 | 0:35 | COMPLETED | SEMI-CRITICAL
1112 | 2024-01-24 | Jan 2024 | 8:00 | 8:10 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 550 | JERLIN | OT | SLAVE MONITOR NOT WORKING | CHANGED THE SETTINGS AND RECTIFIED THE ISSUE. | NO | NO | 2024-01-24 | January 2024 | 8:20 | 2024-01-24 08:00 | 2024-01-24 08:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1113 | 2024-01-24 | Jan 2024 | 11:00 | 11:10 | 0.0069444 | MONITOR HIGH DEFINITION | KARL STORZ | TM220 | JERLIN | OT | DISPLAY COLOUR CHANGED | CAMERA CABLE CONNECTED PROPERLY AND RECTIFIED THE ISSUE. | NO | NO | 2024-01-24 | January 2024 | 11:20 | 2024-01-24 11:00 | 2024-01-24 11:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1114 | 2024-01-24 | Jan 2024 | 10:00 | 10:15 | 0.010417 | ACT MACHINE | HELENO LABORATORIES | ACTALYKE | JERLIN | CATHLAB | POL ERROR | CHECKED WITH ANOTHER SAMPLE AND IT'S WORKING FINE. | NO | NO | 2024-01-24 | January 2024 | 10:25 | 2024-01-24 10:00 | 2024-01-24 10:25 | 0:25 | COMPLETED | SEMI-CRITICAL
1115 | 2024-01-24 | Jan 2024 | 7:15 | 7:20 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | JERLIN | OT | LEAK IN BP CUFF | FIXED IT AND NOW IT'S WORKING GOOD | NO | NO | 2024-01-24 | January 2024 | 7:25 | 2024-01-24 07:15 | 2024-01-24 07:25 | 0:10 | COMPLETED | SEMI-CRITICAL
1116 | 2024-01-24 | Jan 2024 | 7:25 | 7:25 | 0 | PATIENT MONITOR | PHILIPS | CM12 | JERLIN | OT | LEAK IN BP CUFF | FIXED IT AND NOW IT'S WORKING GOOD | NO | NO | 2024-01-24 | January 2024 | 7:30 | 2024-01-24 07:25 | 2024-01-24 07:30 | 0:05 | COMPLETED | SEMI-CRITICAL
1117 | 2024-01-24 | Jan 2024 | 7:30 | 7:30 | 0 | PATIENT MONITOR | PHILIPS | CM12 | JERLIN | OT | NOT ABLE TO CONNECT THE BP CUFF | LUBRICATED THE CONNECTOR AND FIXED IT | NO | NO | 2024-01-24 | January 2024 | 7:35 | 2024-01-24 07:30 | 2024-01-24 07:35 | 0:05 | COMPLETED | SEMI-CRITICAL
1118 | 2024-01-25 | Jan 2024 | 1:00 | 1:05 | 0.0034722 | INCUBATOR, INTENSIVE CARE | WIPRO GE HEALTH CARE | GIRAFFE | SOORAJ | ICU | SOUND HEARD WHILE SWITCHING ON | REMOVED AND REFIXED INCUBATOR BEDS AND SENSORS, ISSUE RECTIFIED. | NO | NO | 2024-01-25 | January 2024 | 1:40 | 2024-01-25 01:00 | 2024-01-25 01:40 | 0:40 | COMPLETED | SEMI-CRITICAL
1119 | 2024-01-25 | Jan 2024 | 10:00 | 10:10 | 0.0069444 | MRI MACHINE | PHILIPS | INGENIA ELITON S | RAMKIE | RADIOLOGY | MRI GRADIENT COIL OVER HEATED | COMPANY ENGG CHECKED AND CONFIRMED THE PROBLEM WITH QIB | QIB | NO | 2024-01-25 | January 2024 | 10:50 | 2024-01-25 10:00 | 2024-01-25 10:50 | 0:50 | COMPLETED | CRITICAL
1120 | 2024-01-25 | Jan 2024 | 13:00 | 13:10 | 0.0069444 | VENTILATOR, ADULT | MAQUET | SERVO-I | RAMKIE | ICU | O2 SENSOR NOT DETECTED ERROR | CHECKED AND RECONNECTED THE O2 SESNOR SECTION, NOW IT WORKIN GOOD | NO | NO | 2024-01-25 | January 2024 | 13:35 | 2024-01-25 13:00 | 2024-01-25 13:35 | 0:35 | COMPLETED | CRITICAL
1121 | 2024-01-26 | Jan 2024 | 9:50 | 9:55 | 0.0034722 | MRI MACHINE | PHILIPS | INGENIA ELITON S | SOORAJ | RADIOLOGY | GRADIENT COIL OVERHEAT ALARM, UNOBLE TO PERFORM SCAN | FOUND ISSUE WITH QIB BOARD. REPLACED WITH NEW ONE, CALIBRATION DONE AND SOFTWARE UPGRADED. TEST SCAN DONE AND ISSUE RECTIFIED. | QUADRATURE INTERFACE BOARD | NO | 2024-01-26 | January 2024 | 11:00 | 2024-01-26 09:50 | 2024-01-26 11:00 | 1:10 | COMPLETED | CRITICAL
1122 | 2024-01-26 | Jan 2024 | 7:00 | 7:00 | 0 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | SOORAJ | RADIOLOGY | CONSOLE DISPLAY NOT SWITCHIN ON | SWAPPED AND CHECKED WITH ANOTHER POWER ADAPTOR, OBSERVED WORKING, FOUND POWER ADAPTOR DEFECTIVE, NEEDS REPLACEMENT. | NO | POWER ADAPTOR | 2024-01-26 | January 2024 | 7:30 | 2024-01-26 07:00 | 2024-01-26 07:30 | 0:30 | PENDING | SEMI-CRITICAL
1123 | 2024-01-26 | Jan 2024 | 9:30 | 9:35 | 0.0034722 | TREATMENT UNIT | SKANRAY DENTAL | SKANLECTICA OPAL | JERLIN | OPD | UNWANTED NOISE DURING USAGE | FIXED THE AIR COMPRESSOR HOSE RECTIFIED THE ISSUE | NO | NO | 2024-01-26 | January 2024 | 9:45 | 2024-01-26 09:30 | 2024-01-26 09:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1124 | 2024-01-27 | Jan 2024 | 20:00 | 20:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | SOORAJ | WARD | BP VALUE NOT SHWOING. | FOUND BP CUFF DEFECTIVE. CHANGED CUFF AND HOSE FROM ANOTHER WARD MONITOR AND ISSUE RECTIFIED. | BP CUFF | NO | 2024-01-27 | January 2024 | 20:20 | 2024-01-27 20:00 | 2024-01-27 20:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1125 | 2024-01-27 | Jan 2024 | 19:25 | 19:30 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | WARD | COT NOT SWITCHING ON | FOUND POWER CHORD DISCONNECTED FROM CONTROL BOX. RECONNECTED AND ISSUE RECTIFIED. | NO | NO | 2024-01-27 | January 2024 | 19:40 | 2024-01-27 19:25 | 2024-01-27 19:40 | 0:15 | COMPLETED | NON-CRITICAL
1126 | 2024-01-27 | Jan 2024 | 10:30 | 10:33 | 0.0020833 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | HAFIZ | OT | O2 CELL ISSUE | MOISTURE IN O2 CELL AND FLOW SENSOR. SWAPPED THE O2 CELL FROM 00149. NOW WORKING GOOD. | O2 CELL | NO | 2024-01-27 | January 2024 | 11:00 | 2024-01-27 10:30 | 2024-01-27 11:00 | 0:30 | COMPLETED | CRITICAL
1127 | 2024-01-27 | Jan 2024 | 17:00 | 17:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | HAFIZ | OT | TOUCH NOT WORKING AND AUTOMATICALLY OFF ISSUE | REPLACED THE INTERFACE BOARD BY THE COM[PANY PERSON. TOUCH NOW WORKING GOOD. | INTERFACE BOARD | NO | 2024-01-27 | January 2024 | 17:30 | 2024-01-27 17:00 | 2024-01-27 17:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1128 | 2024-01-27 | Jan 2024 | 13:40 | 13:45 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | RAMKIE | ICU | EXP CASSETTE ERROR | CHECKED AND FOUND THE PROBLEM WITH EXP CASSETTE, SO REPLACED IT AND USED WITH TEST LUNG CIRCUIT | EXPIRATORY CASSETTE | NO | 2024-01-27 | January 2024 | 14:10 | 2024-01-27 13:40 | 2024-01-27 14:10 | 0:30 | COMPLETED | CRITICAL
1129 | 2024-01-28 | Jan 2024 | 8:15 | 8:20 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | HAFIZ | ICU | O2 CONCENTRATION LOW ALARM | O2 CELL ADJUSTMENT DONE. ALSO O2 CELL CAPACITY FOUND OK (73%). NOW WORKING GOOD. | NO | NO | 2024-01-28 | January 2024 | 8:30 | 2024-01-28 08:15 | 2024-01-28 08:30 | 0:15 | COMPLETED | CRITICAL
1130 | 2024-01-29 | Jan 2024 | 12:20 | 12:30 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 450 | RAMKIE | ICU | NO POWER INDICATION | CHECKED AND FOUND THE PROBLEM WITH INTERFACE BOARD, SO REPLACED IT, NOW IT WORKING | INTERFACE BOARD | NO | 2024-01-29 | January 2024 | 12:50 | 2024-01-29 12:20 | 2024-01-29 12:50 | 0:30 | COMPLETED | SEMI-CRITICAL
1131 | 2024-01-29 | Jan 2024 | 12:50 | 13:00 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 450 | RAMKIE | ICU | NO POWER INDICATION | CHECKED AND CONFIRMED THE PROBLEM, WORKING INTERMITTENTLY, UNDER OBSERVATION | NO | NO | 2024-01-29 | January 2024 | 13:10 | 2024-01-29 12:50 | 2024-01-29 13:10 | 0:20 | COMPLETED | SEMI-CRITICAL
1132 | 2024-01-29 | Jan 2024 | 19:00 | 19:15 | 0.010417 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | JERLIN | WARD | NOT ABLE TO CONNECT THE BP CUFF | LUBRICATED THE CONNECTOR AND FIXED IT | NO | NO | 2024-01-29 | January 2024 | 19:20 | 2024-01-29 19:00 | 2024-01-29 19:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1133 | 2024-01-30 | Jan 2024 | 11:10 | 11:20 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 450 | RAMKIE | ICU | NO POWER INDICATION | CHECKED AND CONFIRMED THE PROBLEM, WORKING INTERMITTENTLY, UNDER OBSERVATION | NO | NO | 2024-01-30 | January 2024 | 11:40 | 2024-01-30 11:10 | 2024-01-30 11:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1134 | 2024-01-30 | Jan 2024 | 9:00 | 9:10 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | CARTRIDGE EXPIRED | CHANGED NEW CARTRIDGE AND CVP DONE. | CARTRIDGE | NO | 2024-01-30 | January 2024 | 9:30 | 2024-01-30 09:00 | 2024-01-30 09:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1135 | 2024-01-31 | Jan 2024 | 10:10 | 10:20 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | RAMKIE | ENDOSCOPY | NOT WORKING | CHECKED AND RESET DONE, NOW IT WORKING | NO | NO | 2024-01-31 | January 2024 | 10:30 | 2024-01-31 10:10 | 2024-01-31 10:30 | 0:20 | COMPLETED | NON-CRITICAL
1136 | 2024-02-01 | Feb 2024 | 22:00 | 22:10 | 0.0069444 | NEBULIZER | JK MEDICALS | ULTIMA | JERLIN | WARD | KNOB BROKEN | FIXED THE KNOB AND RECTIFIED THE ISSUE | NO | NO | 2024-02-01 | February 2024 | 22:30 | 2024-02-01 22:00 | 2024-02-01 22:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1137 | 2024-02-01 | Feb 2024 | 22:30 | 22:30 | 0 | NEBULIZER | JK MEDICALS | ULTIMA | JERLIN | WARD | KNOB BROKEN | FIXED THE KNOB AND RECTIFIED THE ISSUE | NO | NO | 2024-02-01 | February 2024 | 22:45 | 2024-02-01 22:30 | 2024-02-01 22:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1138 | 2024-02-01 | Feb 2024 | 22:45 | 22:45 | 0 | NEBULIZER | JK MEDICALS | ULTIMA | JERLIN | WARD | KNOB BROKEN | FIXED THE KNOB AND RECTIFIED THE ISSUE | NO | NO | 2024-02-01 | February 2024 | 23:05 | 2024-02-01 22:45 | 2024-02-01 23:05 | 0:20 | COMPLETED | SEMI-CRITICAL
1139 | 2024-02-01 | Feb 2024 | 23:10 | 23:10 | 0 | NEBULIZER | JK MEDICALS | ULTIMA | JERLIN | WARD | KNOB BROKEN | FIXED THE KNOB AND RECTIFIED THE ISSUE | NO | NO | 2024-02-01 | February 2024 | 23:15 | 2024-02-01 23:10 | 2024-02-01 23:15 | 0:05 | COMPLETED | SEMI-CRITICAL
1140 | 2024-02-01 | Feb 2024 | 8:05 | 8:10 | 0.0034722 | TABLE TOP STEAM STERILIZER | GETINGE | HS 22K7 | THANISH | CSSD | STEAM STERILIZER NOT WORKING | GENERATED NEW PASSWORD AND PERFORMED THE FRESH CYCLE.WORKING FINE | NO | NO | 2024-02-01 | February 2024 | 8:30 | 2024-02-01 08:05 | 2024-02-01 08:30 | 0:25 | COMPLETED | SEMI-CRITICAL
1141 | 2024-02-01 | Feb 2024 | 9:10 | 9:15 | 0.0034722 | DEFIBRILLATOR MONITOR WITH AED | PHILIPS | EFFICIA DFM 100 | THANISH | OT | SHOCK ABORTED ERROR | RECTIFIED THE PADDLES ISSUE.WORKING FINE. | NO | NO | 2024-02-01 | February 2024 | 9:25 | 2024-02-01 09:10 | 2024-02-01 09:25 | 0:15 | COMPLETED | CRITICAL
1142 | 2024-02-01 | Feb 2024 | 10:15 | 10:20 | 0.0034722 | CATHLAB - BIPLANE | PHILIPS | AZURION 7B20 | THANISH | CATHLAB | 3D MODALITY SLAVE NOT WORKING | SWITCHED ON THE CONSOLE.WORKING FINE. | NO | NO | 2024-02-01 | February 2024 | 10:30 | 2024-02-01 10:15 | 2024-02-01 10:30 | 0:15 | COMPLETED | CRITICAL
1143 | 2024-02-01 | Feb 2024 | 2:00 | 2:15 | 0.010417 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | JERLIN | ICU | O2 SENSOR FAILED | CHANGED O2 SENSOR AND RECTIFIED THE ISSUE | O2 CELL | NO | 2024-02-01 | February 2024 | 2:35 | 2024-02-01 02:00 | 2024-02-01 02:35 | 0:35 | COMPLETED | CRITICAL
1144 | 2024-02-01 | Feb 2024 | 13:30 | 13:35 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SANTHOSH | ENDOSCOPY | O2 CELL ISSUE | CALIBRATED O2 CELL, FOUND O2 CELL DEFECT | NO | NO | 2024-02-01 | February 2024 | 13:40 | 2024-02-01 13:30 | 2024-02-01 13:40 | 0:10 | COMPLETED | CRITICAL
1145 | 2024-02-01 | Feb 2024 | 11:20 | 11:30 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | RAMKIE | WARD | RIGHT SIDE RAIL PROBLEM | CHECKED AND PROBLEM NOT RECTIFIED, INFORMED TO COMPANY PERSON | NO | SIDE REEL | 2024-02-01 | February 2024 | 11:40 | 2024-02-01 11:20 | 2024-02-01 11:40 | 0:20 | PENDING | NON-CRITICAL
1146 | 2024-02-02 | Feb 2024 | 6:35 | 6:40 | 0.0034722 | PLASMA STERILIZER | JOHNSON&JOHNSON | STERRAD 100NX | JERLIN | CSSD | MACHINE NOT SWITCHING ON | RESTARTED THE MACHINE AND NOW IT'S WORKING GOOD | NO | NO | 2024-02-02 | February 2024 | 6:45 | 2024-02-02 06:35 | 2024-02-02 06:45 | 0:10 | COMPLETED | SEMI-CRITICAL
1147 | 2024-02-02 | Feb 2024 | 7:10 | 7:12 | 0.0013889 | PATIENT MONITOR | PHILIPS | MX 550 | JERLIN | OT | MONITOR SWITCHED OFF | RESTARTED THE MONITOR AND NOW IT'S WORKING GOOD | NO | NO | 2024-02-02 | February 2024 | 7:15 | 2024-02-02 07:10 | 2024-02-02 07:15 | 0:05 | COMPLETED | SEMI-CRITICAL
1148 | 2024-02-02 | Feb 2024 | 22:20 | 22:20 | 0 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | JERLIN | OT | ECHO PROBE NOT WORKING | FIXED IT PROPERLY AND NOW IT'S WORKING GOOD. | NO | NO | 2024-02-02 | February 2024 | 22:30 | 2024-02-02 22:20 | 2024-02-02 22:30 | 0:10 | COMPLETED | SEMI-CRITICAL
1149 | 2024-02-02 | Feb 2024 | 2:00 | 2:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | JERLIN | WARD | ARTIFACTS IN ECG WAVES | TURNED ON THE NOTCH FILTER AND RECTIFIED THE ISSUE. | NO | NO | 2024-02-02 | February 2024 | 2:15 | 2024-02-02 02:00 | 2024-02-02 02:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1150 | 2024-02-02 | Feb 2024 | 11:25 | 11:30 | 0.0034722 | BMI MACHINE | AKAS MEDICAL | QVS-100 | THANISH | OPD | NOT SWITCHING ON | REPLACED THE POWER ADAPTOR,AS OLD ADAPTOR FAILED.WORKING FINE. | POWER ADAPTOR | NO | 2024-02-02 | February 2024 | 11:50 | 2024-02-02 11:25 | 2024-02-02 11:50 | 0:25 | COMPLETED | NON-CRITICAL
1151 | 2024-02-02 | Feb 2024 | 12:05 | 12:10 | 0.0034722 | MRI MACHINE | PHILIPS | INGENIA ELITON S | THANISH | RADIOLOGY | WORK STATION UNIT MOUSE NOT WORKING | REPLACED WITH NEW MOUSE. WORKING FINE. | MOUSE | NO | 2024-02-02 | February 2024 | 12:30 | 2024-02-02 12:05 | 2024-02-02 12:30 | 0:25 | COMPLETED | CRITICAL
1152 | 2024-02-02 | Feb 2024 | 13:25 | 13:30 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | THANISH | OT | TIDAL VOLUME ABNORMAL ISSUE | PERFORMED THE FLOW SENSOR CALIBRATION.WORKING FINE. | NO | NO | 2024-02-02 | February 2024 | 13:50 | 2024-02-02 13:25 | 2024-02-02 13:50 | 0:25 | COMPLETED | CRITICAL
1153 | 2024-02-02 | Feb 2024 | 14:05 | 14:10 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | THANISH | OT | DISPLAY NOT WORKING | REFIXED THE DISPLAY CABLE.WORKING FINE | NO | NO | 2024-02-02 | February 2024 | 14:40 | 2024-02-02 14:05 | 2024-02-02 14:40 | 0:35 | COMPLETED | SEMI-CRITICAL
1154 | 2024-02-02 | Feb 2024 | 22:00 | 22:10 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | OT | NOT WORKING | POWER CORD WAS REMOVED FROM THE CONTROL BOX FIXED IT AND NOW IT'S WORKING GOOD. | NO | NO | 2024-02-02 | February 2024 | 22:20 | 2024-02-02 22:00 | 2024-02-02 22:20 | 0:20 | COMPLETED | NON-CRITICAL
1155 | 2024-02-02 | Feb 2024 | 13:40 | 13:45 | 0.0034722 | VENTILATOR, TRANSPORT | RESMED | ASTRAL 100 | SANTHOSH | ICU | NOT CHARGING ISSUE | CHECKED POWERCORD ISSUE RECTIFIED | NO | NO | 2024-02-02 | February 2024 | 13:50 | 2024-02-02 13:40 | 2024-02-02 13:50 | 0:10 | COMPLETED | CRITICAL
1156 | 2024-02-02 | Feb 2024 | 11:30 | 11:40 | 0.0069444 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | RAMKIE | OPD | NO PAPER OUTPUT | CHECKED AND REPLACED PAPER, PROBLEM RECTIFIED | NO | NO | 2024-02-02 | February 2024 | 11:50 | 2024-02-02 11:30 | 2024-02-02 11:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1157 | 2024-02-03 | Feb 2024 | 22:20 | 22:25 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | JERLIN | OT | NOT WORKING | FIXED THE TEMP PROBE PROPERLY AND NOW IT'S WORKING GOOD | NO | NO | 2024-02-03 | February 2024 | 22:40 | 2024-02-03 22:20 | 2024-02-03 22:40 | 0:20 | COMPLETED | SEMI-CRITICAL
1158 | 2024-02-03 | Feb 2024 | 9:10 | 9:15 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | THANISH | OT | COT NOT WORKING | RE-SETTED THE CONTROL BBOX | NO | NO | 2024-02-03 | February 2024 | 9:30 | 2024-02-03 09:10 | 2024-02-03 09:30 | 0:20 | COMPLETED | NON-CRITICAL
1159 | 2024-02-03 | Feb 2024 | 11:05 | 11:10 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | THANISH | OT | ECG NOT WORKING | REPAIRED THE 5 LEAD ECG CABLE.WORKING FINE. | NO | NO | 2024-02-03 | February 2024 | 11:35 | 2024-02-03 11:05 | 2024-02-03 11:35 | 0:30 | COMPLETED | SEMI-CRITICAL
1160 | 2024-02-03 | Feb 2024 | 11:35 | 11:40 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | THANISH | EMERGENCY | PRINTER NOT WORKING | CLEANED THE PRINTER HEAD. WORKING FINE. | NO | NO | 2024-02-03 | February 2024 | 11:55 | 2024-02-03 11:35 | 2024-02-03 11:55 | 0:20 | COMPLETED | CRITICAL
1161 | 2024-02-03 | Feb 2024 | 3:00 | 3:10 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | JERLIN | OT | CIRCUIT LEAK TEST FAILED | FIXED THE CANISTER PROPERLY AND RECTIFIED THE ISSUE. | NO | NO | 2024-02-03 | February 2024 | 3:25 | 2024-02-03 03:00 | 2024-02-03 03:25 | 0:25 | COMPLETED | CRITICAL
1162 | 2024-02-03 | Feb 2024 | 3:25 | 3:30 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | JERLIN | ICU | INTERNOL LEAK TEST FAILED. | FIXED THE EXPIRATORY CASSETTE PROPERLY AND RECTIFIED THE ISSUE. | NO | NO | 2024-02-03 | February 2024 | 3:35 | 2024-02-03 03:25 | 2024-02-03 03:35 | 0:10 | COMPLETED | CRITICAL
1163 | 2024-02-03 | Feb 2024 | 21:30 | 21:40 | 0.0069444 | IABP MACHINE | MAQUET | CS300 | JERLIN | OT | SLAVE CABLE NOT WORKING | FIXED IT AND IT'S WORKING GOOD. | NO | NO | 2024-02-03 | February 2024 | 21:50 | 2024-02-03 21:30 | 2024-02-03 21:50 | 0:20 | COMPLETED | CRITICAL
1164 | 2024-02-03 | Feb 2024 | 3:35 | 3:35 | 0 | VENTILATOR, ADULT | MAQUET | SERVO-I | JERLIN | ICU | CIRCUIT LEAKAGE | CHECKED WITH A STANDARD CIRCUIT AND FOUND NO ISSUES | NO | NO | 2024-02-03 | February 2024 | 3:40 | 2024-02-03 03:35 | 2024-02-03 03:40 | 0:05 | COMPLETED | CRITICAL
1165 | 2024-02-03 | Feb 2024 | 17:00 | 17:05 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SANTHOSH | WARD | COT NOT WORKING | REFIXED AND RESTARTED ISSUE RECTIFIED | NO | NO | 2024-02-03 | February 2024 | 17:10 | 2024-02-03 17:00 | 2024-02-03 17:10 | 0:10 | COMPLETED | NON-CRITICAL
1166 | 2024-02-03 | Feb 2024 | 9:00 | 9:10 | 0.0069444 | DIALYSIS MACHINE | FRESENIUS | 4008 S NG | RAMKIE | DIALYSIS | FLOW ALARM | CLEANED DEGASING FILTER. | NO | NO | 2024-02-03 | February 2024 | 9:40 | 2024-02-03 09:00 | 2024-02-03 09:40 | 0:40 | COMPLETED | CRITICAL
1167 | 2024-02-04 | Feb 2024 | 16:10 | 16:20 | 0.0069444 | CT MACHINE | PHILIPS | INGENUITY | HAFIZ | RADIOLOGY | CIRS ERROR | FOUND THAT CPU WAS SWITCHED OFF. TURNED ON THE CPU AND LOGGED-IN AGAIN. CONNECTION TO CIRS WAS FOUND OK NOW.ISSUE RECTIFIED AND MACHINE READY TO USE. | NO | NO | 2024-02-04 | February 2024 | 16:30 | 2024-02-04 16:10 | 2024-02-04 16:30 | 0:20 | COMPLETED | CRITICAL
1168 | 2024-02-04 | Feb 2024 | 16:30 | 16:35 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | HAFIZ | WARD | PRINTER NOT WORKING | FOUND ISSUES WITH THE PRINTER PAPER. REPLACED ANOTHER NEW PAPER ROLL AND CHECKED. NOW WORKING GOOD. | PAPER ROLL | NO | 2024-02-04 | February 2024 | 16:45 | 2024-02-04 16:30 | 2024-02-04 16:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1169 | 2024-02-04 | Feb 2024 | 17:30 | 17:35 | 0.0034722 | TRANSPORT VENTILATOR | RESMED | ASTRAL 100 | HAFIZ | ICU | NOT WORKING | FOUND THAT BATTERY FULLY DEPLETED. TAKEN TO DEPARTMENT AND CONNECTED WIRES IN THE PLUGTOP. PLUGGED TO AC SOURCE AND NOW WORKING GOOD. | NO | NO | 2024-02-04 | February 2024 | 18:00 | 2024-02-04 17:30 | 2024-02-04 18:00 | 0:30 | COMPLETED | SEMI-CRITICAL
1170 | 2024-02-04 | Feb 2024 | 18:10 | 18:15 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | WARD | PROBE NOT WORKING | CHECKED AND FOUND THAT SPO2 PROBE NOT WORKING. REFIXED AND CHECKED AND NOW WORKING GOOD. | NO | NO | 2024-02-04 | February 2024 | 18:30 | 2024-02-04 18:10 | 2024-02-04 18:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1171 | 2024-02-04 | Feb 2024 | 23:00 | 23:15 | 0.010417 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | JERLIN | WARD | SPO2 PROBE SLOT BROKEN | FIXED IT AND IT'S WORKING GOOD. | NO | NO | 2024-02-04 | February 2024 | 23:30 | 2024-02-04 23:00 | 2024-02-04 23:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1172 | 2024-02-04 | Feb 2024 | 21:00 | 21:10 | 0.0069444 | STEAM STERILIZER | GETINGE | GSS67H102E | JERLIN | CSSD | CONTACTOR ISSUE | CHANGED THE CONTACTOR AND RECTIFIED THE IISUE | CONTACTOR | NO | 2024-02-04 | February 2024 | 23:00 | 2024-02-04 21:00 | 2024-02-04 23:00 | 2:00 | COMPLETED | SEMI-CRITICAL
1173 | 2024-02-04 | Feb 2024 | 20:30 | 20:35 | 0.0034722 | BIPAP MACHINE | PHILIPS | A40 | JERLIN | WARD | MACHINE NOT SWITCHING ON | POWER CORD WAS NOT FIXED PROPERLY , FIXED IT AND IT'S WORKING GOOD | NO | NO | 2024-02-04 | February 2024 | 20:40 | 2024-02-04 20:30 | 2024-02-04 20:40 | 0:10 | COMPLETED | SEMI-CRITICAL
1174 | 2024-02-04 | Feb 2024 | 16:50 | 16:55 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | ICU | PRINTER NOT WORKING | CLEANED THE PRINTER HEAD AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-02-04 | February 2024 | 17:10 | 2024-02-04 16:50 | 2024-02-04 17:10 | 0:20 | COMPLETED | CRITICAL
1175 | 2024-02-04 | Feb 2024 | 13:10 | 13:15 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | O2 CELL ISSUE | FOUND MOISTURE IN O2 CELL, CLEANED, AND REFIXED. ISSUE RECTIFIED. | NO | NO | 2024-02-04 | February 2024 | 14:00 | 2024-02-04 13:10 | 2024-02-04 14:00 | 0:50 | COMPLETED | CRITICAL
1176 | 2024-02-05 | Feb 2024 | 2:00 | 2:10 | 0.0069444 | ACT MACHINE | HELENO LABORATORIES | ACTALYKE | JERLIN | OT | NOT WORKING | CHECKED WITH ANOTHER CATRIDGE AND NOW IT'S WORKING GOOD | NO | NO | 2024-02-05 | February 2024 | 2:25 | 2024-02-05 02:00 | 2024-02-05 02:25 | 0:25 | COMPLETED | SEMI-CRITICAL
1177 | 2024-02-05 | Feb 2024 | 6:00 | 6:10 | 0.0069444 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | JERLIN | RADIOLOGY | DISPLAY NOT WORKING | RESTARTED THE MACHINE AND NOW IT'S WORKING GOOD | NO | NO | 2024-02-05 | February 2024 | 6:30 | 2024-02-05 06:00 | 2024-02-05 06:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1178 | 2024-02-05 | Feb 2024 | 9:10 | 9:10 | 0 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | THANISH | RADIOLOGY | MACHINE NOT SWITCHING ON | PLUG TOP BROKEN/REPLACED WITH NEW 16AMPS PLUG TOP.WORKING FINE. | PLUG TOP | NO | 2024-02-05 | February 2024 | 9:30 | 2024-02-05 09:10 | 2024-02-05 09:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1179 | 2024-02-05 | Feb 2024 | 9:35 | 9:40 | 0.0034722 | TREATMENT UNIT | SKANRAY DENTAL | SKANLECTICA OPAL | THANISH | OPD | MOBILE LIGHT ISSUE | CLEANED THE MOISTURE INSIDE THE MOBILE LIGHT.WORKING FINE. | NO | NO | 2024-02-05 | February 2024 | 10:10 | 2024-02-05 09:35 | 2024-02-05 10:10 | 0:35 | COMPLETED | NON-CRITICAL
1180 | 2024-02-05 | Feb 2024 | 10:20 | 10:25 | 0.0034722 | OT LIGHT | MAQUET | ARDVCS209006A VCS64SF | THANISH | OT | OT LIGHT ALIGNMENT ISSUE | ADJUSTED THE OT LIGHT PENDANT ALIGNMENT.WORKING FINE. | NO | NO | 2024-02-05 | February 2024 | 10:45 | 2024-02-05 10:20 | 2024-02-05 10:45 | 0:25 | COMPLETED | NON-CRITICAL
1181 | 2024-02-05 | Feb 2024 | 11:10 | 11:15 | 0.0034722 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | THANISH | ICU | NOT WORKING | SYRINGE CLAMP HOLDER ISSUE RECTIFIED | NO | NO | 2024-02-05 | February 2024 | 11:30 | 2024-02-05 11:10 | 2024-02-05 11:30 | 0:20 | COMPLETED | NON-CRITICAL
1182 | 2024-02-05 | Feb 2024 | 12:05 | 12:10 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | THANISH | OPD | FREEZE BUTTON ISSUE | REPLACED THE FREEZE BUTTON CONTROLLER BOARD. | FREEZE KEY ASSEMBLY | NO | 2024-02-05 | February 2024 | 12:50 | 2024-02-05 12:05 | 2024-02-05 12:50 | 0:45 | COMPLETED | SEMI-CRITICAL
1183 | 2024-02-05 | Feb 2024 | 17:20 | 17:27 | 0.0048611 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SANTHOSH | WARD | HIGHT ADJUSTMENT NOT WORKING | RESET DONE ,NOW WORKING | NO | NO | 2024-02-05 | February 2024 | 17:35 | 2024-02-05 17:20 | 2024-02-05 17:35 | 0:15 | COMPLETED | NON-CRITICAL
1184 | 2024-02-05 | Feb 2024 | 13:20 | 13:30 | 0.0069444 | LAMINAR AIR FLOW | MICROFLOW DEVICES INDIA PVT LTD | MFD V 1200 | RAMKIE | LABORATORY | NO UV LIGHT | CHECKED AND FOUND THE BALANCING PLATE PROBLEM, RECTIFIED AND NOW WORKING | NO | NO | 2024-02-05 | February 2024 | 13:40 | 2024-02-05 13:20 | 2024-02-05 13:40 | 0:20 | COMPLETED | NON-CRITICAL
1185 | 2024-02-06 | Feb 2024 | 23:00 | 23:10 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 450 | HAFIZ | EMERGENCY | NOT ABLE TO PERFORM NIBP | CHECKED AND FOUND THAT NIBP DISABLED. ENABLED THE NIBP MEASUREMENT. NOW WORKING GOOD. | NO | NO | 2024-02-06 | February 2024 | 23:20 | 2024-02-06 23:00 | 2024-02-06 23:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1186 | 2024-02-06 | Feb 2024 | 13:40 | 13:50 | 0.0069444 | INCUBATOR | VANI SCIENTIFIC&EQUIPMENTS | NO | RAMKIE | LABORATORY | LOCK BROKEN | CHECKED AND NEED TO REPLACED NEW ONE, QUOTATION RECEIVED FROM VENDOR | NO | NO | 2024-02-06 | February 2024 | 14:00 | 2024-02-06 13:40 | 2024-02-06 14:00 | 0:20 | COMPLETED | NON-CRITICAL
1187 | 2024-02-06 | Feb 2024 | 1:10 | 1:15 | 0.0034722 | NITRIC OXIDE DELIVERY SYSTEM | WIPRO GE HEALTH CARE | SLE | HAFIZ | ICU | NOT WORKING | RESET DONE, RESTARTED THE MACHINE. NOW WORKING GOOD. | NO | NO | 2024-02-06 | February 2024 | 1:30 | 2024-02-06 01:10 | 2024-02-06 01:30 | 0:20 | COMPLETED | CRITICAL
1188 | 2024-02-07 | Feb 2024 | 1:00 | 1:00 | 0 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | NEED TO REPLACE A CARTRIDGE | REMOVED THE OLD CARTRIDGE (LOT:231206I) AND INSTALLED A NEW CARTRIDGE. WARMING UP DONE. READY TO USE. | CARTRIDGE | NO | 2024-02-07 | February 2024 | 1:45 | 2024-02-07 01:00 | 2024-02-07 01:45 | 0:45 | COMPLETED | SEMI-CRITICAL
1189 | 2024-02-07 | Feb 2024 | 6:00 | 6:05 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | HAFIZ | OPD | NOT WORKING | FOUND THAT BATTERY FULLY DEPLETED. PLUGGED TO AC SOURCE. ISSUE RECTIFIED AND READY TO USE. | NO | NO | 2024-02-07 | February 2024 | 6:15 | 2024-02-07 06:00 | 2024-02-07 06:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1190 | 2024-02-07 | Feb 2024 | 6:35 | 6:40 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | WARD | ALARM INDICATION | CHECKED AND FOUND THAT SPO2 PROBE GOT DETACHED FROM THE PATIENT. ASKED THEM TO REFIX AND ISSUE RECTIFIED. | NO | NO | 2024-02-07 | February 2024 | 6:50 | 2024-02-07 06:35 | 2024-02-07 06:50 | 0:15 | COMPLETED | SEMI-CRITICAL
1191 | 2024-02-07 | Feb 2024 | 20:00 | 20:13 | 0.0090278 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | JERLIN | ICU | NEED TO CHANGE A NEW CARTRIDGE | REPLACED A NEW CARTDRIGE AND CVP DONE. | CARTRIDGE | NO | 2024-02-07 | February 2024 | 20:30 | 2024-02-07 20:00 | 2024-02-07 20:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1192 | 2024-02-08 | Feb 2024 | 8:00 | 8:05 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | ICU | NOT WORKING | ULTRASONIC SENSOR BROKEN. REPLACED ANOTHER ULTRASONIC SENSOR FROM DEFECTIVE PUMP AND ISSUE RECTIFIED. | NO | NO | 2024-02-08 | February 2024 | 8:30 | 2024-02-08 08:00 | 2024-02-08 08:30 | 0:30 | COMPLETED | NON-CRITICAL
1193 | 2024-02-08 | Feb 2024 | 21:00 | 21:04 | 0.0027778 | BIPAP MACHINE | RESMED | STELLAR 150 | HAFIZ | WARD | SETTINGS TO BE DONE | CHANGED THE SETTINGS AS PER REQUIREMENTS. NOW WORKING GOOD. | NO | NO | 2024-02-08 | February 2024 | 21:20 | 2024-02-08 21:00 | 2024-02-08 21:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1194 | 2024-02-08 | Feb 2024 | 17:30 | 17:40 | 0.0069444 | IVUS MACHINE | PHILIPS | CORE SERIES S5 | RAMKIE | CATHLAB | ERROR CATHETER FAULT DETECTOR | CHECKED AND NEED TO REPLACE THE PIM MODULE | NO | NO | 2024-02-08 | February 2024 | 18:30 | 2024-02-08 17:30 | 2024-02-08 18:30 | 1:00 | COMPLETED | SEMI-CRITICAL
1195 | 2024-02-09 | Feb 2024 | 22:05 | 22:10 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | WARD | NOT WORKING | CHECKED AND FOUND THAT BATTERY FULLY DEPLETED. PLUGGED TO AC SOURCE AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-02-09 | February 2024 | 22:30 | 2024-02-09 22:05 | 2024-02-09 22:30 | 0:25 | COMPLETED | SEMI-CRITICAL
1196 | 2024-02-09 | Feb 2024 | 15:00 | 15:15 | 0.010417 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | JERLIN | RADIOLOGY | FREEZE KEY NOT WORKING | CHANGED THE FREEZE KEY ASSAY AND NOW IT'S WORKING GOOD | FREEZE KEY ASSEMBLY | NO | 2024-02-09 | February 2024 | 15:30 | 2024-02-09 15:00 | 2024-02-09 15:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1197 | 2024-02-09 | Feb 2024 | 14:20 | 14:30 | 0.0069444 | CENTRIFUGE TABLE TOP | REMI | NEYA 4 | RAMKIE | LABORATORY | UNOBLE TO UNLOCK | CHECKED AND FOUND THE PROBLEM WITH HINGE SOLENOID. REPLACED | HINGE SOLENOID | NO | 2024-02-09 | February 2024 | 14:40 | 2024-02-09 14:20 | 2024-02-09 14:40 | 0:20 | COMPLETED | SEMI-CRITICAL
1198 | 2024-02-09 | Feb 2024 | 16:15 | 16:20 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | MOISTURE INSIDE CANISTER, CLEANED AND CHECKOUT DONE. ISSUE RECTIFIED. | NO | NO | 2024-02-09 | February 2024 | 17:00 | 2024-02-09 16:15 | 2024-02-09 17:00 | 0:45 | COMPLETED | CRITICAL
1199 | 2024-02-10 | Feb 2024 | 14:40 | 14:45 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | RAMKIE | RADIOLOGY | WORKLIST NOT UPDATING | CHECKED AND RECONNECTED THE LAN CABLE | NO | NO | 2024-02-10 | February 2024 | 14:55 | 2024-02-10 14:40 | 2024-02-10 14:55 | 0:15 | COMPLETED | SEMI-CRITICAL
1200 | 2024-02-10 | Feb 2024 | 18:10 | 18:20 | 0.0069444 | IVUS MACHINE | PHILIPS | CORE SERIES S5 | RAMKIE | CATHLAB | ERROR CATHETER FAULT DETECTOR | REPLACED THE PIM MODULE, PROBLEM RECTIFIED | PIM MODULE | NO | 2024-02-10 | February 2024 | 18:40 | 2024-02-10 18:10 | 2024-02-10 18:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1201 | 2024-02-10 | Feb 2024 | 8:15 | 8:20 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | CANISTER CHANGED AND OBSERVED. ISSUE RECTIFIED | CANISTER | NO | 2024-02-10 | February 2024 | 8:45 | 2024-02-10 08:15 | 2024-02-10 08:45 | 0:30 | COMPLETED | CRITICAL
1202 | 2024-02-10 | Feb 2024 | 9:00 | 9:00 | 0 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | CLEANED CANISTER, SODALIME QUANTITY REDUCED. WORKING GOOD. | NO | NO | 2024-02-10 | February 2024 | 9:30 | 2024-02-10 09:00 | 2024-02-10 09:30 | 0:30 | COMPLETED | CRITICAL
1203 | 2024-02-12 | Feb 2024 | 14:50 | 15:00 | 0.0069444 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | RAMKIE | OT | LOCKING CLAMP PROBLEM | CHECKED AND ADJUSTED AND RECTIFIED PROBLEM | NO | NO | 2024-02-12 | February 2024 | 15:15 | 2024-02-12 14:50 | 2024-02-12 15:15 | 0:25 | COMPLETED | NON-CRITICAL
1204 | 2024-02-13 | Feb 2024 | 14:35 | 14:40 | 0.0034722 | PATIENT MONITOR,VITAL SIGN | WELCH ALLYN | 75CT | HAFIZ | OPD | SPO2 PROBE NOT WORKING | CLEANED AND REFIXED THE PROBE. INTERMITTENTLY WORKING, NEED TO REPLACE THE PROBE. | NO | NO | 2024-02-13 | February 2024 | 14:45 | 2024-02-13 14:35 | 2024-02-13 14:45 | 0:10 | COMPLETED | SEMI-CRITICAL
1205 | 2024-02-13 | Feb 2024 | 15:20 | 15:30 | 0.0069444 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | RAMKIE | OT | OCCLUSION ERROR | CHECKED AND CHANGED TO M AND OBSERVED, IT WORKING | NO | NO | 2024-02-13 | February 2024 | 15:40 | 2024-02-13 15:20 | 2024-02-13 15:40 | 0:20 | COMPLETED | NON-CRITICAL
1206 | 2024-02-14 | Feb 2024 | 14:30 | 14:40 | 0.0069444 | EXAMINATION LIGHT | BURTON | NXF103 | JERLIN | OPD | NOT WORKING | RESTARTED AND NOW IT'S WORKING GOOD. | NO | NO | 2024-02-14 | February 2024 | 14:45 | 2024-02-14 14:30 | 2024-02-14 14:45 | 0:15 | COMPLETED | NON-CRITICAL
1207 | 2024-02-14 | Feb 2024 | 15:30 | 15:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | WARD | NOT WORKING | REMOTE LOCKED , UNLOCKED AND NOW IT'S WORKING GOOD. | NO | NO | 2024-02-14 | February 2024 | 15:40 | 2024-02-14 15:30 | 2024-02-14 15:40 | 0:10 | COMPLETED | NON-CRITICAL
1208 | 2024-02-14 | Feb 2024 | 12:20 | 12:25 | 0.0034722 | GO FLO PUMP, ARTHROSCOPY | SMITH & NEPHEW | 72204968 | SANTHOSH | OT | NOT WORKING | CHECKED AND RESET DONE NOW WORKING | NO | NO | 2024-02-14 | February 2024 | 12:35 | 2024-02-14 12:20 | 2024-02-14 12:35 | 0:15 | COMPLETED | SEMI-CRITICAL
1209 | 2024-02-14 | Feb 2024 | 12:35 | 12:40 | 0.0034722 | LITHOCLAST | SURGIMEDIK HEALTHCARE | DIGI NOVA | SANTHOSH | OT | NOT WORKING | CHECKED AND OBSERVED HAND PIECE NEED TO REPLACE | HANDPIECE | NO | 2024-02-14 | February 2024 | 12:50 | 2024-02-14 12:35 | 2024-02-14 12:50 | 0:15 | COMPLETED | SEMI-CRITICAL
1210 | 2024-02-14 | Feb 2024 | 10:00 | 10:04 | 0.0027778 | DIALYSIS MACHINE | FRESENIUS | 4008 S NG | RAMKIE | DIALYSIS | FLOW ALARM | CLEANED DEGASING FILTER. | NO | NO | 2024-02-14 | February 2024 | 10:40 | 2024-02-14 10:00 | 2024-02-14 10:40 | 0:40 | COMPLETED | CRITICAL
1211 | 2024-02-15 | Feb 2024 | 13:15 | 13:20 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | JERLIN | OT | NOT WORKING | FIXED THE TEMP PROBE PROPERLY AND NOW IT'S WORKING GOOD | NO | NO | 2024-02-15 | February 2024 | 13:30 | 2024-02-15 13:15 | 2024-02-15 13:30 | 0:15 | COMPLETED | SEMI-CRITICAL
1212 | 2024-02-15 | Feb 2024 | 0:45 | 0:50 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | GLUCOSE IQM FAILED | NEW CADRIDGE REPLACED | CARTRIDGE | NO | 2024-02-15 | February 2024 | 1:00 | 2024-02-15 00:45 | 2024-02-15 01:00 | 0:15 | COMPLETED | SEMI-CRITICAL
1213 | 2024-02-15 | Feb 2024 | 0:30 | 0:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SANTHOSH | WARD | NOT WORKING | CHECKED ,CONTROL PANEL DEFECT,NEED TO REPLACE | NO | NO | 2024-02-15 | February 2024 | 0:45 | 2024-02-15 00:30 | 2024-02-15 00:45 | 0:15 | COMPLETED | NON-CRITICAL
1214 | 2024-02-15 | Feb 2024 | 15:20 | 15:30 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | RAMKIE | OT | CIRCUIT LEAK TEST FAILED | CHECKED AND SODA LIME CHAMBER RECONNECTED, RECTIFIED | NO | NO | 2024-02-15 | February 2024 | 15:50 | 2024-02-15 15:20 | 2024-02-15 15:50 | 0:30 | COMPLETED | CRITICAL
1215 | 2024-02-16 | Feb 2024 | 1:15 | 1:20 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | SANTHOSH | ICU | SERVICE INDICATION ALARM | RESET DONE NOW WORKING GOOD | NO | NO | 2024-02-16 | February 2024 | 1:30 | 2024-02-16 01:15 | 2024-02-16 01:30 | 0:15 | COMPLETED | SEMI-CRITICAL
1216 | 2024-02-17 | Feb 2024 | 0:45 | 0:50 | 0.0034722 | HEATER UNIT | GETINGE | HU35 | SANTHOSH | OT | HOSE LEAKING ISSUE | LEAK ARRESTED, | NO | NO | 2024-02-17 | February 2024 | 1:00 | 2024-02-17 00:45 | 2024-02-17 01:00 | 0:15 | COMPLETED | SEMI-CRITICAL
1217 | 2024-02-17 | Feb 2024 | 2:00 | 2:05 | 0.0034722 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | SANTHOSH | RADIOLOGY | SUDDENLY TURNED OFF | FOUND PLUG WAS LOOSENED THEN FIXED | NO | NO | 2024-02-17 | February 2024 | 2:10 | 2024-02-17 02:00 | 2024-02-17 02:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1218 | 2024-02-17 | Feb 2024 | 16:20 | 16:30 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | OT | NOT WORKING | CHECKED AND RESET DONE, PROBLEM RECTIFIED | NO | NO | 2024-02-17 | February 2024 | 16:50 | 2024-02-17 16:20 | 2024-02-17 16:50 | 0:30 | COMPLETED | NON-CRITICAL
1219 | 2024-02-17 | Feb 2024 | 15:00 | 15:15 | 0.010417 | MRI MACHINE | PHILIPS | INGENIA ELITON S | JERLIN | RADIOLOGY | MOUSE NOT WORKING | CHANGED IT AND NOW IT'S WORKING GOOD | NO | NO | 2024-02-17 | February 2024 | 15:30 | 2024-02-17 15:00 | 2024-02-17 15:30 | 0:30 | COMPLETED | CRITICAL
1220 | 2024-02-18 | Feb 2024 | 23:40 | 23:45 | 0.0034722 | HEATER COOLER | GETINGE | HCU40 | SANTHOSH | OT | PLUG TOP BURNT AND STUCKED | NEW PLUG TOP REPLACED | PLUG TOP | NO | 2024-02-18 | February 2024 | 23:50 | 2024-02-18 23:40 | 2024-02-18 23:50 | 0:10 | COMPLETED | CRITICAL
1221 | 2024-02-19 | Feb 2024 | 15:30 | 15:33 | 0.0020833 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | HAFIZ | EMERGENCY | UNOBLE TO TAKE NIBP MEASUREMENTS | CHECKED THE CUFFS AND FOUND OK. REPLACED ANOTHER BP HOSE AND VERIFIED. WORKING GOOD. | NIBP HOSE | NO | 2024-02-19 | February 2024 | 15:45 | 2024-02-19 15:30 | 2024-02-19 15:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1222 | 2024-02-19 | Feb 2024 | 16:30 | 16:40 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 450 | RAMKIE | ICU | NIBP TIMING PROBLEM | CHECKED AND RECTIFIED | NO | NO | 2024-02-19 | February 2024 | 16:50 | 2024-02-19 16:30 | 2024-02-19 16:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1223 | 2024-02-20 | Feb 2024 | 12:20 | 12:24 | 0.0027778 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 9000 | HAFIZ | WARD | NOT WORKING | FOUND THAT BED LOCKED. UNLOCKED AND OBSERVED. WORKING GOOD. | NO | NO | 2024-02-20 | February 2024 | 12:30 | 2024-02-20 12:20 | 2024-02-20 12:30 | 0:10 | COMPLETED | NON-CRITICAL
1224 | 2024-02-20 | Feb 2024 | 16:50 | 17:00 | 0.0069444 | VENTILATOR, ADULT | MAQUET | SERVO-I | RAMKIE | ICU | TEST NOT PASSED | CHECKED AND RECONNECTED, PROBLEM CLEARED | NO | NO | 2024-02-20 | February 2024 | 17:00 | 2024-02-20 16:50 | 2024-02-20 17:00 | 0:10 | COMPLETED | CRITICAL
1225 | 2024-02-21 | Feb 2024 | 9:55 | 10:00 | 0.0034722 | BABY WEIGHING SCALE WITH INFANTOMETER | SECA | 334 | HAFIZ | ICU | VARIATION IN WEIGH OF BABY | VERIFIED (CALIBRATED) WITH THE STANDARD WEIGH. WORKING GOOD. | STANDARD WEIGH | NO | 2024-02-21 | February 2024 | 10:10 | 2024-02-21 09:55 | 2024-02-21 10:10 | 0:15 | COMPLETED | NON-CRITICAL
1226 | 2024-02-21 | Feb 2024 | 9:30 | 9:35 | 0.0034722 | CENTRIFUGE TABLE TOP | REMI | NEYA 4 | HAFIZ | BLOODBANK | DOOR NOT CLOSING | FOUND THAT DOOR BACKEND SCREW LOOSENED. TIGHTENED IT AND CHECKED. NOW WORKING GOOD. | NO | NO | 2024-02-21 | February 2024 | 9:50 | 2024-02-21 09:30 | 2024-02-21 09:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1227 | 2024-02-21 | Feb 2024 | 14:30 | 14:35 | 0.0034722 | PATIENT MONITOR ,TRANSPORT | PHILIPS | MX 100 | HAFIZ | EMERGENCY | NOT ABLE TO PERFORM NIBP | CLEANED AND REFIXED THE HOSE, REEXAMINED AND WORKING GOOD. | NO | NO | 2024-02-21 | February 2024 | 14:45 | 2024-02-21 14:30 | 2024-02-21 14:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1228 | 2024-02-21 | Feb 2024 | 15:00 | 15:03 | 0.0020833 | ECHO MACHINE | GE HEALTHCARE | VIVID S70 | HAFIZ | ICU | POWERCORD ISSUE | CLIPPED THE POWERCORD CABLE TO THE HOLDER SO THAT IT CANNOT DISTURB CASTOR WHEEL. ALSO INSULATED THE DAMAGED PART OF THE POWERCORD. CHECKED AND WORKING GOOD. | NO | NO | 2024-02-21 | February 2024 | 15:30 | 2024-02-21 15:00 | 2024-02-21 15:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1229 | 2024-02-21 | Feb 2024 | 17:00 | 17:05 | 0.0034722 | MAMMOGRAPH UNIT | KIRAN MEDICAL SYSTEM | FELICIA | HAFIZ | RADIOLOGY | BUTTONS NOT FUNCTIONING | REMOVED AND REFIXED THE CONNECTOR. OBSERVED AND WORKING GOOD. | NO | NO | 2024-02-21 | February 2024 | 17:30 | 2024-02-21 17:00 | 2024-02-21 17:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1230 | 2024-02-21 | Feb 2024 | 10:15 | 10:20 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | JERLIN | OT | NOT WORKING PROPERLY | CLEARED THE STORAGE AND NOW IT'S WORKINF GOOD. | NO | NO | 2024-02-21 | February 2024 | 10:40 | 2024-02-21 10:15 | 2024-02-21 10:40 | 0:25 | COMPLETED | SEMI-CRITICAL
1231 | 2024-02-21 | Feb 2024 | 17:00 | 17:10 | 0.0069444 | DIALYSIS MACHINE | FRESENIUS | 4008 S NG | RAMKIE | DIALYSIS | V99 FAILURE | CHECKED AND CLEANED THE FILTER, PROBLEM RECTIFIED | NO | NO | 2024-02-21 | February 2024 | 17:20 | 2024-02-21 17:00 | 2024-02-21 17:20 | 0:20 | COMPLETED | CRITICAL
1232 | 2024-02-22 | Feb 2024 | 10:30 | 10:35 | 0.0034722 | MAMMOGRAPH UNIT | KIRAN MEDICAL SYSTEM | FELICIA | HAFIZ | RADIOLOGY | FACEPLATE DETACHED | ATTACHED THE FACE PLATE TO THE CONSOLE. NOW READY TO USE. | NO | NO | 2024-02-22 | February 2024 | 11:00 | 2024-02-22 10:30 | 2024-02-22 11:00 | 0:30 | COMPLETED | SEMI-CRITICAL
1233 | 2024-02-22 | Feb 2024 | 16:00 | 16:04 | 0.0027778 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC V2 | HAFIZ | EMERGENCY | NOT SWITCHING ON | CHECKED AND FOUND THAT BATTERY FULLY DEPLETED. PLUGGED TO AC SOURCE AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-02-22 | February 2024 | 16:15 | 2024-02-22 16:00 | 2024-02-22 16:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1234 | 2024-02-22 | Feb 2024 | 15:10 | 15:20 | 0.0069444 | REVERSE OSMOSIS SYSTEM | ENGINEERING ALLIANCE | 500 LPH | JERLIN | DIALYSIS | NOT WORKING | RESTARTED THE MACHINE AND NOW IT'S WORKING GOOD | NO | NO | 2024-02-22 | February 2024 | 15:25 | 2024-02-22 15:10 | 2024-02-22 15:25 | 0:15 | COMPLETED | SEMI-CRITICAL
1235 | 2024-02-22 | Feb 2024 | 17:30 | 17:40 | 0.0069444 | EXAMINATION COUCH | MIDMARK JANAK | MX400A | RAMKIE | DIALYSIS | HEAD POSITION NOT WORKING | PROBLEM WITH GAS SPRING, NEED TO REPLACE | NO | GAS SPRING | 2024-02-22 | February 2024 | 17:50 | 2024-02-22 17:30 | 2024-02-22 17:50 | 0:20 | PENDING | NON-CRITICAL
1236 | 2024-02-22 | Feb 2024 | 9:00 | 9:03 | 0.0020833 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | HAFIZ | ICU | O2 CELL DEFECTIVE | REPLACED THE O2 CELL AND CHECK-OUT PASSED. NOW WORKING GOOD. | O2 CELL | NO | 2024-02-22 | February 2024 | 9:30 | 2024-02-22 09:00 | 2024-02-22 09:30 | 0:30 | COMPLETED | CRITICAL
1237 | 2024-02-22 | Feb 2024 | 9:00 | 9:10 | 0.0069444 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | JERLIN | ICU | MEMBRANE EXPIRED | CHANGED A NEW MEMBRANE AND RECTIFIED THE ISSUE. | MEMBRANE | NO | 2024-02-22 | February 2024 | 9:20 | 2024-02-22 09:00 | 2024-02-22 09:20 | 0:20 | COMPLETED | CRITICAL
1238 | 2024-02-23 | Feb 2024 | 9:20 | 9:25 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | WARD | AIR-IN LINE ALARM | FOUND THAT INFUSION TUBING NOT PROPERLY FIXED. REFIXED AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-02-23 | February 2024 | 9:45 | 2024-02-23 09:20 | 2024-02-23 09:45 | 0:25 | COMPLETED | NON-CRITICAL
1239 | 2024-02-23 | Feb 2024 | 9:50 | 9:55 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | HAFIZ | ICU | UNOBLE TO SEND PACS | CLEARED THE SPOOLER, RESTARTED THE MACHINE. RESEND THE PACS AND NOW ISSUE RECTIFIED. | NO | NO | 2024-02-23 | February 2024 | 10:10 | 2024-02-23 09:50 | 2024-02-23 10:10 | 0:20 | COMPLETED | SEMI-CRITICAL
1240 | 2024-02-24 | Feb 2024 | 10:00 | 10:10 | 0.0069444 | MOVEMENT THERAPHY UNIT, UPPER AND LOWER LIMBS | RECK | MOTOMED LETTO2 | HAFIZ | PHYSIOTHERAPY | LOCK ISSUE | FABRICATED THE DYE CAST RINGS AND FIXED IT BY THE VENDOR. NOW WORKING FINE. | NO | NO | 2024-02-24 | February 2024 | 11:00 | 2024-02-24 10:00 | 2024-02-24 11:00 | 1:00 | COMPLETED | SEMI-CRITICAL
1241 | 2024-02-24 | Feb 2024 | 13:00 | 13:07 | 0.0048611 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | HAFIZ | WARD | NOT SWITCHING ON | FOUND THAT BATTERY FULLY DEPLETED. PLUGGED TO AC SOURCE AND OBSERVED. ISSUE RECTIFIED AND READY TO USE. | NO | NO | 2024-02-24 | February 2024 | 13:30 | 2024-02-24 13:00 | 2024-02-24 13:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1242 | 2024-02-24 | Feb 2024 | 17:55 | 18:00 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | CVP TEST DUE | ALL CVP TESTS DONE. MACHINE NOW READY TO USE. | CARTRIDGE | NO | 2024-02-24 | February 2024 | 18:25 | 2024-02-24 17:55 | 2024-02-24 18:25 | 0:30 | COMPLETED | SEMI-CRITICAL
1243 | 2024-02-26 | Feb 2024 | 18:20 | 18:30 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | RAMKIE | ICU | PRINTER PROBLEM | CHECKED AND RECTIFIED | NO | NO | 2024-02-26 | February 2024 | 18:40 | 2024-02-26 18:20 | 2024-02-26 18:40 | 0:20 | COMPLETED | CRITICAL
1244 | 2024-02-27 | Feb 2024 | 22:00 | 22:04 | 0.0027778 | BIPAP MACHINE | PHILIPS | A40 | HAFIZ | WARD | SETTINGS TO BE DONE | CHANGED THE SETTINGS AS PER REQUIREMENTS. NOW WORKING GOOD. | NO | NO | 2024-02-27 | February 2024 | 22:15 | 2024-02-27 22:00 | 2024-02-27 22:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1245 | 2024-02-27 | Feb 2024 | 11:50 | 11:55 | 0.0034722 | HARMONIC GENERATOR | LOTUS | UBERSONIC TM | SANTHOSH | OT | NOT WORKING | CHECKED WITH ANOTHER HANDPIECE ,ISSUE RECTIFIED | NO | NO | 2024-02-27 | February 2024 | 12:05 | 2024-02-27 11:50 | 2024-02-27 12:05 | 0:15 | COMPLETED | SEMI-CRITICAL
1246 | 2024-02-27 | Feb 2024 | 18:00 | 18:20 | 0.013889 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | RAMKIE | ICU | AIR IN LINE | CHECKED ULTRA SOUND SENSOR | NO | NO | 2024-02-27 | February 2024 | 18:40 | 2024-02-27 18:00 | 2024-02-27 18:40 | 0:40 | COMPLETED | NON-CRITICAL
1247 | 2024-02-28 | Feb 2024 | 22:00 | 22:03 | 0.0020833 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID E95 | HAFIZ | OPD | PROBE STUCK IN CASTOR WHEEL | RELEASED THE PROBE WIRE FROM THE WHEEL. APPLIED TAPE TO THE DAMAGED PART OF THE PROBE. CHECKED THE PROBE AND WORKING GOOD. | NO | NO | 2024-02-28 | February 2024 | 22:30 | 2024-02-28 22:00 | 2024-02-28 22:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1248 | 2024-02-28 | Feb 2024 | 11:00 | 11:10 | 0.0069444 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | JERLIN | ICU | NEED TO CHANGE A NEW CARTRIDGE | REPLACED A NEW CARTDRIGE AND CVP DONE. | CARTRIDGE | NO | 2024-02-28 | February 2024 | 11:40 | 2024-02-28 11:00 | 2024-02-28 11:40 | 0:40 | COMPLETED | SEMI-CRITICAL
1249 | 2024-02-28 | Feb 2024 | 11:45 | 11:55 | 0.0069444 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | JERLIN | ICU | NEED TO CHANGE A NEW CARTRIDGE | REPLACED A NEW CARTDRIGE AND CVP DONE. | CARTRIDGE | NO | 2024-02-28 | February 2024 | 12:05 | 2024-02-28 11:45 | 2024-02-28 12:05 | 0:20 | COMPLETED | SEMI-CRITICAL
1250 | 2024-02-28 | Feb 2024 | 11:00 | 11:05 | 0.0034722 | DIALYSIS MACHINE | FRESENIUS | 4008S NG | SOORAJ | DIALYSIS | FLOW ALARM | CLEANED DEGASING FILTER. | NO | NO | 2024-02-28 | February 2024 | 12:00 | 2024-02-28 11:00 | 2024-02-28 12:00 | 1:00 | COMPLETED | CRITICAL
1251 | 2024-02-29 | Feb 2024 | 2:00 | 2:03 | 0.0020833 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | HAFIZ | WARD | NOT WORKING | RESET DONE, AND CHECKED. WORKING GOOD. | NO | NO | 2024-02-29 | February 2024 | 2:10 | 2024-02-29 02:00 | 2024-02-29 02:10 | 0:10 | COMPLETED | NON-CRITICAL
1252 | 2024-02-29 | Feb 2024 | 20:00 | 20:06 | 0.0041667 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | NEED TO REPLACE A CARTRIDGE | REMOVED THE OLD CARTRIDGE AND INSTALLED A NEW CARTRIDGE. CVP-2 TEST ONLY FAILED. | CARTRIDGE | NO | 2024-02-29 | February 2024 | 20:50 | 2024-02-29 20:00 | 2024-02-29 20:50 | 0:50 | COMPLETED | SEMI-CRITICAL
1253 | 2024-02-29 | Feb 2024 | 14:00 | 14:15 | 0.010417 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | JERLIN | ICU | PRINTER NOT WORKING | CLEARED THE QUEUE AND NOW IT'S WORKING GOOD | NO | NO | 2024-02-29 | February 2024 | 14:25 | 2024-02-29 14:00 | 2024-02-29 14:25 | 0:25 | COMPLETED | SEMI-CRITICAL
1254 | 2024-02-29 | Feb 2024 | 9:00 | 9:11 | 0.0076389 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SANTHOSH | OT | AC POWER FAIL ISSUE | CHECK OUT AND CAKIBRATION DONE ISSUE RECTIFIED | NO | NO | 2024-02-29 | February 2024 | 9:20 | 2024-02-29 09:00 | 2024-02-29 09:20 | 0:20 | COMPLETED | CRITICAL
1255 | 2024-03-01 | Mar 2024 | 6:30 | 6:35 | 0.0034722 | FETAL HEART MONITOR | WIPRO GE HEALTH CARE | COROMETRICS 259 | HAFIZ | WARD | AUTOMATICALLY SWITCHED OFF | CHECKED AND FOUND THAT AC POWER FAILED. CROSSCHECKED WITH ANOTHER SOCKET AND NO SUPPLY IN ANY SOCKET. INFORMED MAINTENONCE. | NO | NO | 2024-03-01 | March 2024 | 6:45 | 2024-03-01 06:30 | 2024-03-01 06:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1256 | 2024-03-01 | Mar 2024 | 19:00 | 19:02 | 0.0013889 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | CVP-2 TEST FAILED MULTIPLE TIMES | ADDED THE CVP-2 LOT TO THE MACHINEBY THE COMPANY PERSON. RERUNNED THE TEST AND NOW PASSED. MACHINE IS READY TO USE. | NO | NO | 2024-03-01 | March 2024 | 20:00 | 2024-03-01 19:00 | 2024-03-01 20:00 | 1:00 | COMPLETED | SEMI-CRITICAL
1257 | 2024-03-01 | Mar 2024 | 20:00 | 20:00 | 0 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | CVP-2 TEST FAILED MULTIPLE TIMES | ADDED THE CVP-2 LOT TO THE MACHINEBY THE COMPANY PERSON. RERUNNED THE TEST AND NOW PASSED. MACHINE IS READY TO USE. | NO | NO | 2024-03-01 | March 2024 | 20:30 | 2024-03-01 20:00 | 2024-03-01 20:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1258 | 2024-03-01 | Mar 2024 | 7:05 | 7:18 | 0.0090278 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | THANISH | ICU | ABG CATRIDGE EXPIRED | REPLACED THE NEW ABG CATRIDGE AND DONE QC.WORKING FINE | CARTRIDGE | NO | 2024-03-01 | March 2024 | 7:20 | 2024-03-01 07:05 | 2024-03-01 07:20 | 0:15 | COMPLETED | SEMI-CRITICAL
1259 | 2024-03-01 | Mar 2024 | 20:25 | 20:30 | 0.0034722 | ECMO MACHINE | GETINGE | CARDIO HELP | THANISH | ICU | VENOUS PRESSURE LOW | REPLACED NEW VENOUS PRESSURE CABLE.WORKING FINE. | VENOUS PRESSURE CABLE | NO | 2024-03-01 | March 2024 | 20:45 | 2024-03-01 20:25 | 2024-03-01 20:45 | 0:20 | COMPLETED | CRITICAL
1260 | 2024-03-01 | Mar 2024 | 8:00 | 8:10 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | NEED TO CHANGE A NEW CARTRIDGE | CHANGED A NEW CARTRIDGE AND CVP DONE | CARTRIDGE | NO | 2024-03-01 | March 2024 | 8:30 | 2024-03-01 08:00 | 2024-03-01 08:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1261 | 2024-03-01 | Mar 2024 | 10:10 | 10:30 | 0.013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | EMERGENCY | NOT WORKING | CHECKED AND POWER INSERTED PROPERLY AND RECTIFIED | NO | NO | 2024-03-01 | March 2024 | 10:40 | 2024-03-01 10:10 | 2024-03-01 10:40 | 0:30 | COMPLETED | NON-CRITICAL
1262 | 2024-03-02 | Mar 2024 | 10:10 | 10:15 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | THANISH | ICU | COT BRAKE PEDAL NOT WORKING | SCREW FIXED. ISSUE RECTIFIED | NO | NO | 2024-03-02 | March 2024 | 10:40 | 2024-03-02 10:10 | 2024-03-02 10:40 | 0:30 | COMPLETED | NON-CRITICAL
1263 | 2024-03-02 | Mar 2024 | 15:15 | 15:20 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | THANISH | OT | VOLUME CONTROL MODE ISSUE | CLEANED THE FLOW SENSOR AND SODA LIME JAR & INSPIRATORY, EXPIRATORY VALVES. WORKING FINE. | NO | NO | 2024-03-02 | March 2024 | 15:50 | 2024-03-02 15:15 | 2024-03-02 15:50 | 0:35 | COMPLETED | CRITICAL
1264 | 2024-03-02 | Mar 2024 | 17:10 | 17:15 | 0.0034722 | EXAMINATION COUCH | MIDMARK JANOK | E122012 | THANISH | OPD | HEADREST BROKEN | REPLACED WITH ANOTHER HEAD REST.WORKING FINE. | HEADREST | NO | 2024-03-02 | March 2024 | 17:45 | 2024-03-02 17:10 | 2024-03-02 17:45 | 0:35 | COMPLETED | NON-CRITICAL
1265 | 2024-03-02 | Mar 2024 | 9:10 | 9:20 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | ICU | CVP DUE | CVP DONE | NO | NO | 2024-03-02 | March 2024 | 9:30 | 2024-03-02 09:10 | 2024-03-02 09:30 | 0:20 | COMPLETED | NON-CRITICAL
1266 | 2024-03-02 | Mar 2024 | 10:30 | 10:40 | 0.0069444 | MONITOR HIGH DEFINITION,3D | KARL STORZ | TM330 | JERLIN | OT | RECORDING NOT WORKING | CHANGED THE CONNECTOR CABLES AND NOW IT'S WORKING PROPERLY | HDMI CABLE | NO | 2024-03-02 | March 2024 | 11:00 | 2024-03-02 10:30 | 2024-03-02 11:00 | 0:30 | COMPLETED | SEMI-CRITICAL
1267 | 2024-03-02 | Mar 2024 | 15:30 | 15:45 | 0.010417 | LITHOCLAST | SURGIMEDIK HEALTHCARE | DIGI NOVA | JERLIN | OT | NOT WORKING PROPERLY | CHANGED THE HANDPIECE AND CHECKED IT'S WORKING GOOD | HAND PIECE | NO | 2024-03-02 | March 2024 | 16:15 | 2024-03-02 15:30 | 2024-03-02 16:15 | 0:45 | COMPLETED | SEMI-CRITICAL
1268 | 2024-03-02 | Mar 2024 | 11:10 | 11:30 | 0.013889 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | RAMKIE | EMERGENCY | ECG DISTURBANCE | CLEANED THE ECG CONNECTING PORT AND ISSUE RECTIFIED | NO | NO | 2024-03-02 | March 2024 | 11:40 | 2024-03-02 11:10 | 2024-03-02 11:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1269 | 2024-03-03 | Mar 2024 | 22:00 | 22:04 | 0.0027778 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | ICU | KEY STUCK | REPLACED THE CONTROL BOARD FROM ANOTHER DEFECTIVE PUMP (00446). CHECKED AND NOW WORKING GOOD. | CONTROL BOARD | NO | 2024-03-03 | March 2024 | 22:40 | 2024-03-03 22:00 | 2024-03-03 22:40 | 0:40 | COMPLETED | NON-CRITICAL
1270 | 2024-03-03 | Mar 2024 | 22:40 | 22:40 | 0 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | ICU | CONTROL BOARD ISSUE | CLEANED THE CONTROL BOARD AND CHECKED. WORKING GOOD. | NO | NO | 2024-03-03 | March 2024 | 23:00 | 2024-03-03 22:40 | 2024-03-03 23:00 | 0:20 | COMPLETED | NON-CRITICAL
1271 | 2024-03-03 | Mar 2024 | 13:00 | 13:10 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 450 | JERLIN | EMERGENCY | NIBP LEAK | LEAK ARRESTED AND ISSUE RECTIFIED | NO | NO | 2024-03-03 | March 2024 | 13:20 | 2024-03-03 13:00 | 2024-03-03 13:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1272 | 2024-03-03 | Mar 2024 | 13:20 | 13:20 | 0 | PATIENT MONITOR | PHILIPS | CM12 | JERLIN | EMERGENCY | NIBP LEAK | LEAK ARRESTED AND ISSUE RECTIFIED | NO | NO | 2024-03-03 | March 2024 | 13:30 | 2024-03-03 13:20 | 2024-03-03 13:30 | 0:10 | COMPLETED | SEMI-CRITICAL
1273 | 2024-03-03 | Mar 2024 | 10:00 | 10:10 | 0.0069444 | TRACTION WITH TREATMENT TABLE | CHATTANOOGA | TRITON DTS | RAMKIE | PHYSIOTHERAPY | NOT SWITCHING ON | RECONNECTED AND ISSUE RECTIFIED | NO | NO | 2024-03-03 | March 2024 | 10:40 | 2024-03-03 10:00 | 2024-03-03 10:40 | 0:40 | COMPLETED | SEMI-CRITICAL
1274 | 2024-03-04 | Mar 2024 | 14:30 | 14:35 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | SANTHOSH | WARD | ALARM IN HIGH VOLUME | CHECKED AND MUTED ALARM | NO | NO | 2024-03-04 | March 2024 | 14:40 | 2024-03-04 14:30 | 2024-03-04 14:40 | 0:10 | COMPLETED | SEMI-CRITICAL
1275 | 2024-03-04 | Mar 2024 | 13:40 | 13:45 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | SANTHOSH | WARD | PRINTER ISSUE | CHECKED WITH ANOTHER PAPER ISSUE RECTIFIED | NO | NO | 2024-03-05 | March 2024 | 13:50 | 2024-03-04 13:40 | 2024-03-05 13:50 | 24:10 | COMPLETED | SEMI-CRITICAL
1276 | 2024-03-04 | Mar 2024 | 10:05 | 10:15 | 0.0069444 | EXAMINATION COUCH | MIDMARK JANOK | E122012 | THANISH | OPD | NOT WORKING | REFIXED THE CABLE CONNECTIONS IN CONTROL BOX.WORKING FINE. | NO | NO | 2024-03-04 | March 2024 | 10:40 | 2024-03-04 10:05 | 2024-03-04 10:40 | 0:35 | COMPLETED | NON-CRITICAL
1277 | 2024-03-04 | Mar 2024 | 10:50 | 10:55 | 0.0034722 | BMI MACHINE | AKAS MEDICAL | QVS-100 | THANISH | OPD | NIBP NOT WORKING | REPLACED THE NIBP HOSE.WORKING FINE. | NIBP HOSE | NO | 2024-03-04 | March 2024 | 11:20 | 2024-03-04 10:50 | 2024-03-04 11:20 | 0:30 | COMPLETED | NON-CRITICAL
1278 | 2024-03-04 | Mar 2024 | 14:00 | 14:05 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | THANISH | ICU | ULTRASOUND NOT WORKING | RE-ATTACHED THE BATTERY MODULE.WORKING FINE. | NO | NO | 2024-03-04 | March 2024 | 14:15 | 2024-03-04 14:00 | 2024-03-04 14:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1279 | 2024-03-04 | Mar 2024 | 14:00 | 14:15 | 0.010417 | BMI MACHINE | AKAS MEDICAL | QVS-100 | JERLIN | OPD | NIBP NOT WORKING | REFIXED THE CUFF PROPERLY AND CHECKED IT'S WORKING GOOD. | NO | NO | 2024-03-04 | March 2024 | 14:25 | 2024-03-04 14:00 | 2024-03-04 14:25 | 0:25 | COMPLETED | NON-CRITICAL
1280 | 2024-03-05 | Mar 2024 | 1:20 | 1:22 | 0.0013889 | PATIENT MONITOR | PHILIPS | MX 550 | SOORAJ | OT | MONITOR OUTPUT NOT SHOWING ON SLAVE MONITOR | FOUND INPUT SOURCE NOT SELECTED, ISSUE RECTIFIED | NO | NO | 2024-03-05 | March 2024 | 1:45 | 2024-03-05 01:20 | 2024-03-05 01:45 | 0:25 | COMPLETED | SEMI-CRITICAL
1281 | 2024-03-05 | Mar 2024 | 11:20 | 11:27 | 0.0048611 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SANTHOSH | WARD | HIGHT ADJUSTMENT NOT WORKING | RESET DONE ,NOW WORKING | NO | NO | 2024-03-05 | March 2024 | 11:30 | 2024-03-05 11:20 | 2024-03-05 11:30 | 0:10 | COMPLETED | NON-CRITICAL
1282 | 2024-03-05 | Mar 2024 | 12:20 | 12:25 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | SANTHOSH | ICU | CANT ABLE TO SAVE IMAGES | CLEARED DATA NOW SAVING IMAGE | NO | NO | 2024-03-05 | March 2024 | 12:30 | 2024-03-05 12:20 | 2024-03-05 12:30 | 0:10 | COMPLETED | SEMI-CRITICAL
1283 | 2024-03-05 | Mar 2024 | 13:35 | 13:40 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | VOLUSON E10 | SANTHOSH | OPD | IMAGES NOT PUSHING TO SYATEM | RESTARTED CHECKED CONNECTION,WORKING GOOD | NO | NO | 2024-03-05 | March 2024 | 13:50 | 2024-03-05 13:35 | 2024-03-05 13:50 | 0:15 | COMPLETED | SEMI-CRITICAL
1284 | 2024-03-05 | Mar 2024 | 18:30 | 18:35 | 0.0034722 | MOVEMENT THERAPHY UNIT, UPPER AND LOWER LIMBS | RECK | MOTOMED LETTO2 | SANTHOSH | PHYSIOTHERAPY | PLUG TOP BROKEN | REPLACED NEW PLUGTOP | PLUG TOP | NO | 2024-03-05 | March 2024 | 18:45 | 2024-03-05 18:30 | 2024-03-05 18:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1285 | 2024-03-05 | Mar 2024 | 17:00 | 17:03 | 0.0020833 | PATIENT MONITOR ,TRANSPORT | PHILIPS | MX 100 | HAFIZ | EMERGENCY | NIBP MEASUREMENT FAILED | CHECKED MOTOR AND FOUND FINE. REMOVED, CLEANED AND REFIXED THE CABLE INSIDE. CHECKED AND NOW WORKING GOOD. | NO | NO | 2024-03-05 | March 2024 | 17:30 | 2024-03-05 17:00 | 2024-03-05 17:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1286 | 2024-03-05 | Mar 2024 | 18:00 | 18:05 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S7 | HAFIZ | RADIOLOGY | C1-6-D PROBE ARTIFACTS ISSUE | THROUGH COMPANY PERSON'S GUIDELINES, CHECKED THE CONDITION OF PROBE AND REPLACEMENT GIVEN. REPLACED THE PROBE AND OBSERVED, WORKING GOOD. | C1-6-D PROBE | NO | 2024-03-05 | March 2024 | 18:30 | 2024-03-05 18:00 | 2024-03-05 18:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1287 | 2024-03-05 | Mar 2024 | 9:10 | 9:15 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | THANISH | OPD | SETTINGS ISSUE | RE-STORED THE DEFAULT IN LOGIQ P9.WORKING FINE. | NO | NO | 2024-03-05 | March 2024 | 9:20 | 2024-03-05 09:10 | 2024-03-05 09:20 | 0:10 | COMPLETED | SEMI-CRITICAL
1288 | 2024-03-05 | Mar 2024 | 10:05 | 10:10 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | THANISH | OT | TEMPERATURE PROBE NOT WORKING. | REPLACED WITH NEW TEMPERATURE PROBE.WORKING FINE. | ESOPHAGEAL/RECTAL PROBE | NO | 2024-03-05 | March 2024 | 10:20 | 2024-03-05 10:05 | 2024-03-05 10:20 | 0:15 | COMPLETED | SEMI-CRITICAL
1289 | 2024-03-05 | Mar 2024 | 10:30 | 10:35 | 0.0034722 | RADIOGRAPHIC UNIT | CARESTREAM | DR ASCEND CMD | THANISH | RADIOLOGY | DETECTOR NOT WORKING. | EXISTING BATTERY FAILED/REPLACED WITH NEW BATTERY.WORKING FINE | BATTERY | NO | 2024-03-05 | March 2024 | 10:50 | 2024-03-05 10:30 | 2024-03-05 10:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1290 | 2024-03-05 | Mar 2024 | 11:10 | 11:30 | 0.013889 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | RAMKIE | OT | CLAMP THREAD NOT GOOD | CHANGED THE CLAMP | CLAMP | NO | 2024-03-05 | March 2024 | 11:55 | 2024-03-05 11:10 | 2024-03-05 11:55 | 0:45 | COMPLETED | NON-CRITICAL
1291 | 2024-03-06 | Mar 2024 | 17:00 | 17:05 | 0.0034722 | VENTILATOR, PEDIATRIC | MINDRAY MEDICAL INDIA | SV 300 | SANTHOSH | ICU | BATTERY BACKUP ISSUE | CHECKED BATTERY AND NEED TO REPLACE THE BATTERY | NO | BATTERY | 2024-03-06 | March 2024 | 17:10 | 2024-03-06 17:00 | 2024-03-06 17:10 | 0:10 | PENDING | CRITICAL
1292 | 2024-03-06 | Mar 2024 | 0:45 | 0:50 | 0.0034722 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | SANTHOSH | WARD | POWERPORT ISSUE | ISSUE RECTIFIED | NO | NO | 2024-03-06 | March 2024 | 1:00 | 2024-03-06 00:45 | 2024-03-06 01:00 | 0:15 | COMPLETED | SEMI-CRITICAL
1293 | 2024-03-06 | Mar 2024 | 11:15 | 11:20 | 0.0034722 | URODYNOMIC SYSTEM | AYMED MEDICAL TECHNOLOGY | LOCUM | SANTHOSH | OPD | PRESSURE TRANSDUCER FAILURE | REPLACED NEW TRANSDUCER | TRANSDUCER | NO | 2024-03-06 | March 2024 | 11:30 | 2024-03-06 11:15 | 2024-03-06 11:30 | 0:15 | COMPLETED | SEMI-CRITICAL
1294 | 2024-03-06 | Mar 2024 | 9:10 | 9:15 | 0.0034722 | TRACTION WITH TREATMENT TABLE | CHATTANOOGA | TRITON DTS | THANISH | PHYSIOTHERAPY | BRAKE NOT WORKING | REFIXED THE SCREW IN THE BRAKE.WORKING FINE | NO | NO | 2024-03-06 | March 2024 | 9:30 | 2024-03-06 09:10 | 2024-03-06 09:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1295 | 2024-03-06 | Mar 2024 | 11:05 | 11:10 | 0.0034722 | DIODE LASER,DERMATOLOGY | ALMA | SOPRANO ICE | THANISH | OPD | LASER GOOGLES BROKEN | REPLACED WITH NEW LASER GOOGLES.WORKING FINE. | LASER GOGGLES | NO | 2024-03-06 | March 2024 | 11:30 | 2024-03-06 11:05 | 2024-03-06 11:30 | 0:25 | COMPLETED | SEMI-CRITICAL
1296 | 2024-03-06 | Mar 2024 | 14:10 | 14:15 | 0.0034722 | CELL SAVER | HAEMONETICS | ELITE PLUS | THANISH | OT | FLUID DETECTED IN CENTRIFUGE WALL | CLEANED THE TUBING SENSOR AND CENTRIFUGE WELL.WORKIN FINE. | NO | NO | 2024-03-06 | March 2024 | 14:55 | 2024-03-06 14:10 | 2024-03-06 14:55 | 0:45 | COMPLETED | CRITICAL
1297 | 2024-03-06 | Mar 2024 | 15:00 | 15:10 | 0.0069444 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | ePM 12c | JERLIN | ICU | PATIENT DATA STORAGE IS FULL | CLEARED THE DATA AND RECTIFIED THE ISSUE. | NO | NO | 2024-03-06 | March 2024 | 15:20 | 2024-03-06 15:00 | 2024-03-06 15:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1298 | 2024-03-06 | Mar 2024 | 11:45 | 11:55 | 0.0069444 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | RAMKIE | OT | OCCLUSION ERROR | CHECKED AND FOUND FEEDING, SO CHANGED TO COMMON AND ASKED TO DILUTE | NO | NO | 2024-03-06 | March 2024 | 12:15 | 2024-03-06 11:45 | 2024-03-06 12:15 | 0:30 | COMPLETED | NON-CRITICAL
1299 | 2024-03-07 | Mar 2024 | 0:45 | 0:50 | 0.0034722 | ORTHOPANTOMOGRAPHY | SKANRAY DENTAL | MYRAY HYPERION X5 | SANTHOSH | OPD | POSITIONING ISSUE | TEETH STICK BROKEN NEED A REPLACEMENT | NO | NO | 2024-03-07 | March 2024 | 1:00 | 2024-03-07 00:45 | 2024-03-07 01:00 | 0:15 | COMPLETED | SEMI-CRITICAL
1300 | 2024-03-07 | Mar 2024 | 10:05 | 10:10 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | THANISH | ICU | TEMPERATURE PROBE NOT WORKING. | REPLACED WITH NEW TEMPERATURE PROBE.WORKING FINE. | ESOPHAGEAL/RECTAL PROBE | NO | 2024-03-07 | March 2024 | 10:30 | 2024-03-07 10:05 | 2024-03-07 10:30 | 0:25 | COMPLETED | SEMI-CRITICAL
1301 | 2024-03-07 | Mar 2024 | 12:05 | 12:10 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | THANISH | ICU | MONITOR DISPLAY NOT WORKING. | DISPLAY PANEL FAILED/SWAPPED WITH ANOTHER WORKING DISPLAY PANEL.WORKING FINE. | DISPLAY PANEL | NO | 2024-03-07 | March 2024 | 12:40 | 2024-03-07 12:05 | 2024-03-07 12:40 | 0:35 | COMPLETED | SEMI-CRITICAL
1302 | 2024-03-07 | Mar 2024 | 12:10 | 12:20 | 0.0069444 | PATIENT MONITOR | PHILIPS | CM12 | RAMKIE | OT | ECG DISTURBANCE | CHECKED AND CHANGED ECG CABLES | 5-LEADS ECG CABLE | NO | 2024-03-07 | March 2024 | 12:40 | 2024-03-07 12:10 | 2024-03-07 12:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1303 | 2024-03-08 | Mar 2024 | 14:00 | 14:05 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | WARD | NOT WORKING | CHECKED AND FOUND THAT THE POWERCORD WAS DETACHED FROM THE CONTROL BOX. REFIXED AND CHECKED, ISSUE RECTIFIED. | NO | NO | 2024-03-08 | March 2024 | 14:20 | 2024-03-08 14:00 | 2024-03-08 14:20 | 0:20 | COMPLETED | NON-CRITICAL
1304 | 2024-03-08 | Mar 2024 | 10:10 | 10:15 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | THANISH | ICU | WARMER NOT WORKING | CLEANED THE HEATER CABLE CONNECTOR PIN.WORKING FINE. | NO | NO | 2024-03-08 | March 2024 | 10:30 | 2024-03-08 10:10 | 2024-03-08 10:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1305 | 2024-03-08 | Mar 2024 | 11:15 | 11:20 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | THANISH | OPD | ECG NOT WORKING | REPLACED THE 5-LEAD ECG CABLE.WORKING FINE. | 5-LEADS ECG CABLE | NO | 2024-03-08 | March 2024 | 11:30 | 2024-03-08 11:15 | 2024-03-08 11:30 | 0:15 | COMPLETED | SEMI-CRITICAL
1306 | 2024-03-08 | Mar 2024 | 14:00 | 14:10 | 0.0069444 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | JERLIN | OT | NEED TO REPLACE NEW HINDGES | NEW HINGES CHANGED | HINGES | NO | 2024-03-08 | March 2024 | 14:35 | 2024-03-08 14:00 | 2024-03-08 14:35 | 0:35 | COMPLETED | SEMI-CRITICAL
1307 | 2024-03-09 | Mar 2024 | 0:35 | 0:40 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | WARD | COT FUNCTIONS NOT WORKING | FOUND BATTERY DEPLETED. CONNECTED ON MAINS AND ISSUE RECTIFIED. | NO | NO | 2024-03-09 | March 2024 | 0:45 | 2024-03-09 00:35 | 2024-03-09 00:45 | 0:10 | COMPLETED | NON-CRITICAL
1308 | 2024-03-09 | Mar 2024 | 1:50 | 1:55 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | WARD | COT LOCKED | RESET DONE AND ISSUE RECTIFIED | NO | NO | 2024-03-09 | March 2024 | 2:00 | 2024-03-09 01:50 | 2024-03-09 02:00 | 0:10 | COMPLETED | NON-CRITICAL
1309 | 2024-03-09 | Mar 2024 | 2:00 | 2:05 | 0.0034722 | LABOUR BED | ARJO HUNTLEIGH | HCB-04LY/BXA | SOORAJ | WARD | IV STAND STUCK | FOUND STAND SCREWS TIGHTENED. LUBRICATED AND ISSUE RECTIFIED | NO | NO | 2024-03-09 | March 2024 | 2:10 | 2024-03-09 02:00 | 2024-03-09 02:10 | 0:10 | COMPLETED | NON-CRITICAL
1310 | 2024-03-09 | Mar 2024 | 23:50 | 23:53 | 0.0020833 | BIPAP MACHINE | PHILIPS | A40 | SOORAJ | WARD | PRESSURE VALUE ABNORMAL | ADJUSTED EPAP SETTINGS AND ISSUE RECTIFIED. | NO | NO | 2024-03-09 | March 2024 | 23:55 | 2024-03-09 23:50 | 2024-03-09 23:55 | 0:05 | COMPLETED | SEMI-CRITICAL
1311 | 2024-03-09 | Mar 2024 | 8:00 | 8:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SANTHOSH | RADIOLOGY | CIRCUIT LEAK TEST FAILED | ISSUE RECTIFIED | NO | NO | 2024-03-09 | March 2024 | 8:10 | 2024-03-09 08:00 | 2024-03-09 08:10 | 0:10 | COMPLETED | CRITICAL
1312 | 2024-03-09 | Mar 2024 | 23:40 | 23:45 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | SANTHOSH | OPD | OPTION NEED TO CHANGE | ISSUE RECTIFIED | NO | NO | 2024-03-09 | March 2024 | 23:50 | 2024-03-09 23:40 | 2024-03-09 23:50 | 0:10 | COMPLETED | SEMI-CRITICAL
1313 | 2024-03-09 | Mar 2024 | 12:30 | 12:35 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | SHUTDOWN THEN CATRIDGE EJECTED | NEED TO REPLACE A CATRIDGE | NO | NO | 2024-03-09 | March 2024 | 12:40 | 2024-03-09 12:30 | 2024-03-09 12:40 | 0:10 | COMPLETED | SEMI-CRITICAL
1314 | 2024-03-09 | Mar 2024 | 14:00 | 14:05 | 0.0034722 | ECMO MACHINE | GETINGE | ROTAFLOW | HAFIZ | ICU | DEFECTIVE BATTERY | REPLACED THE NEW BATTERY WITH THE COMPANY PERSON. ISSUE RECTIFIED. | ECMO BATTERY | NO | 2024-03-09 | March 2024 | 14:30 | 2024-03-09 14:00 | 2024-03-09 14:30 | 0:30 | COMPLETED | CRITICAL
1315 | 2024-03-09 | Mar 2024 | 14:30 | 14:35 | 0.0034722 | ECMO MACHINE | GETINGE | ROTAFLOW | HAFIZ | ICU | DEFECTIVE BATTERY | REPLACED THE NEW BATTERY WITH THE COMPANY PERSON. ISSUE RECTIFIED. | ECMO BATTERY | NO | 2024-03-09 | March 2024 | 15:00 | 2024-03-09 14:30 | 2024-03-09 15:00 | 0:30 | COMPLETED | CRITICAL
1316 | 2024-03-09 | Mar 2024 | 10:10 | 10:15 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | THANISH | OT | TIDAL VOLUME ABNORMAL ISSUE | CLEANED THE FLOW SENSOR AND PERFORMED THE FLOW SENSOR CALIBRATION.WORKING FINE. | NO | NO | 2024-03-09 | March 2024 | 10:30 | 2024-03-09 10:10 | 2024-03-09 10:30 | 0:20 | COMPLETED | CRITICAL
1317 | 2024-03-09 | Mar 2024 | 11:05 | 11:10 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | THANISH | ICU | FLOW TRANSDUCER TEST FAILED | CLEANED AIR FILTER AND REPLACED THE EXPIRATORY CASSETTE.WORKING FINE. | NO | NO | 2024-03-09 | March 2024 | 11:45 | 2024-03-09 11:05 | 2024-03-09 11:45 | 0:40 | COMPLETED | CRITICAL
1318 | 2024-03-09 | Mar 2024 | 12:30 | 12:40 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 450 | RAMKIE | ICU | CENTRAL MONITOR ISSUE | CHECKED AND CONFIGURATION CHANGED. ISSUE RECTIFIED. | NO | NO | 2024-03-09 | March 2024 | 12:50 | 2024-03-09 12:30 | 2024-03-09 12:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1319 | 2024-03-10 | Mar 2024 | 10:00 | 10:02 | 0.0013889 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | ICU | PRINTER NOT WORKING | CLEANED THE PRINTER HEAD. OBSERVED AND WORKING GOOD. | NO | NO | 2024-03-10 | March 2024 | 10:10 | 2024-03-10 10:00 | 2024-03-10 10:10 | 0:10 | COMPLETED | CRITICAL
1320 | 2024-03-10 | Mar 2024 | 10:15 | 10:17 | 0.0013889 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | CATHLAB | ECG EQUIPMENT MALFUNCTION | CLEANED THE CABLE AND REFIXED THE CABLE. PASSED THE OPERATIONAL CHECK. NOW READY TO USE. | NO | NO | 2024-03-10 | March 2024 | 10:30 | 2024-03-10 10:15 | 2024-03-10 10:30 | 0:15 | COMPLETED | CRITICAL
1321 | 2024-03-11 | Mar 2024 | 17:35 | 17:40 | 0.0034722 | HOLTER,RECORDER | WIPRO GE HEALTH CARE | SEER 1000 | SANTHOSH | OPD | DDNT GET SIGNOLS | CHECKED WITH CABLES NOW WORKING | NO | NO | 2024-03-11 | March 2024 | 17:50 | 2024-03-11 17:35 | 2024-03-11 17:50 | 0:15 | COMPLETED | SEMI-CRITICAL
1322 | 2024-03-11 | Mar 2024 | 7:30 | 7:45 | 0.010417 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | CLEANED MOISTURE FROM ABSORBER CANISTER. PASSED THE CHECK NOW READY TO USE. | NO | NO | 2024-03-11 | March 2024 | 7:50 | 2024-03-11 07:30 | 2024-03-11 07:50 | 0:20 | COMPLETED | CRITICAL
1323 | 2024-03-11 | Mar 2024 | 8:00 | 8:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | OT | NOT WORKING | FOUND THAT BATTERY FULLY DEPLETED. PLUGGED TO AC SOURCE BUT NO POWER SOCKET. AFTER POWER SUPPLY, OBSERVED AND ISSUE RECTIFIED. | NO | NO | 2024-03-11 | March 2024 | 8:20 | 2024-03-11 08:00 | 2024-03-11 08:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1324 | 2024-03-11 | Mar 2024 | 15:50 | 15:53 | 0.0020833 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | HAFIZ | OT | NOT WORKING | CHECKED AND FOUND THAT THERMISTOR RECEPTACLE FOUND DETACHED. REFIXED AND ISSUE RECTIFIED. | NO | NO | 2024-03-11 | March 2024 | 16:00 | 2024-03-11 15:50 | 2024-03-11 16:00 | 0:10 | COMPLETED | SEMI-CRITICAL
1325 | 2024-03-11 | Mar 2024 | 16:00 | 16:02 | 0.0013889 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | HAFIZ | OT | MECHANICAL MOVEMENT ISSUE | CHECKED ANGULATION OF THE TEE PROBE AND MECHANICAL ROTATION ALSO GOOD. ISSUE RECTIFIED. | NO | NO | 2024-03-11 | March 2024 | 16:15 | 2024-03-11 16:00 | 2024-03-11 16:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1326 | 2024-03-11 | Mar 2024 | 9:30 | 9:40 | 0.0069444 | VENTILATOR, ADULT | MAQUET | SERVO-I | RAMKIE | ICU | PREOP TEST FAILED | CHECKED AND RECONNECTED CASETTE. WORKING | NO | NO | 2024-03-11 | March 2024 | 10:00 | 2024-03-11 09:30 | 2024-03-11 10:00 | 0:30 | COMPLETED | CRITICAL
1327 | 2024-03-12 | Mar 2024 | 11:50 | 11:55 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID E95 | SANTHOSH | OPD | HANGING ISSUE | CLEARED DATA NOW WORKING GOOD | NO | NO | 2024-03-12 | March 2024 | 12:05 | 2024-03-12 11:50 | 2024-03-12 12:05 | 0:15 | COMPLETED | SEMI-CRITICAL
1328 | 2024-03-12 | Mar 2024 | 14:00 | 14:11 | 0.0076389 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | SANTHOSH | ICU | COLOR DOPPLER ISSUE | CHECKED CURVED LINEAR PROBE PROBLEM | NO | NO | 2024-03-12 | March 2024 | 14:20 | 2024-03-12 14:00 | 2024-03-12 14:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1329 | 2024-03-12 | Mar 2024 | 8:00 | 8:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | HAFIZ | OT | O2 SUPPLY LOW ALARM | RESTARTED THE MACHINE, CHECKOUT DONE AND OBSERVED. NOW NO ALARMS DETECTED. | NO | NO | 2024-03-12 | March 2024 | 8:20 | 2024-03-12 08:00 | 2024-03-12 08:20 | 0:20 | COMPLETED | CRITICAL
1330 | 2024-03-12 | Mar 2024 | 10:30 | 10:33 | 0.0020833 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | CLEANED MOISTURE FROM ABSORBER CANISTER. PASSED THE CHECK NOW READY TO USE. | NO | NO | 2024-03-12 | March 2024 | 10:45 | 2024-03-12 10:30 | 2024-03-12 10:45 | 0:15 | COMPLETED | CRITICAL
1331 | 2024-03-12 | Mar 2024 | 15:00 | 15:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | OT | NIBP MEASUREMENT FAILED | FOUND LEAK IN TWO BP CUFF. BLOCKED THE LEAK. CHECKED AND NOW WORKING GOOD. | NO | NO | 2024-03-12 | March 2024 | 15:20 | 2024-03-12 15:00 | 2024-03-12 15:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1332 | 2024-03-12 | Mar 2024 | 15:15 | 15:20 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | HAFIZ | OT | AUTOMATICALLY SWITCHED OFF | SOFTWARE UPDATED BY THE VENDOR. NOW READY TO USE. | NO | NO | 2024-03-12 | March 2024 | 15:40 | 2024-03-12 15:15 | 2024-03-12 15:40 | 0:25 | COMPLETED | SEMI-CRITICAL
1333 | 2024-03-12 | Mar 2024 | 22:30 | 22:40 | 0.0069444 | FLUID WARMER SYSTEM | MATRIX | MM 5 | JERLIN | OT | LEAK IN FLU | LEAK ARRESTED AND ISSUE RECTIFIED | NO | NO | 2024-03-12 | March 2024 | 23:00 | 2024-03-12 22:30 | 2024-03-12 23:00 | 0:30 | COMPLETED | SEMI-CRITICAL
1334 | 2024-03-12 | Mar 2024 | 13:00 | 13:10 | 0.0069444 | NEBULIZER | PARI | PARIBOY SX | RAMKIE | ICU | PRESSURE IS NOT GOOD | CHANGED THE CONNECTOR AND ISSUE RECTIFIED | NEBULIZER CONNECTOR | NO | 2024-03-12 | March 2024 | 13:20 | 2024-03-12 13:00 | 2024-03-12 13:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1335 | 2024-03-13 | Mar 2024 | 11:00 | 11:11 | 0.0076389 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SANTHOSH | ICU | BEEP SOUND ISSUE | ISSUE RECTIFIED | NO | NO | 2024-03-13 | March 2024 | 11:10 | 2024-03-13 11:00 | 2024-03-13 11:10 | 0:10 | COMPLETED | NON-CRITICAL
1336 | 2024-03-13 | Mar 2024 | 16:00 | 16:10 | 0.0069444 | OXYGEN THERAPY UNIT | INSPIRED | O2FLO | SANTHOSH | ICU | CANT CHANGE SATURATION LEVEL | ADJUST SETTIGS ISSUE RECTIFIED | NO | NO | 2024-03-13 | March 2024 | 16:10 | 2024-03-13 16:00 | 2024-03-13 16:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1337 | 2024-03-13 | Mar 2024 | 8:40 | 8:45 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | HAFIZ | ENDOSCOPY | CHECKOUT FAILED | RECONNECTED THE O2 CELL, RERUNNED THE CHECKOUT AND PASSED. NOW READY TO USE. | NO | NO | 2024-03-13 | March 2024 | 9:00 | 2024-03-13 08:40 | 2024-03-13 09:00 | 0:20 | COMPLETED | CRITICAL
1338 | 2024-03-13 | Mar 2024 | 11:20 | 11:23 | 0.0020833 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | WARD | ARTIFACTS IN ECG WAVES | CHECKED AND FOUND NO ISSUES WITH ECG CABLES. INTERFERENCE CAUSED BY THE POWERCORD. REMOVED IT AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-03-13 | March 2024 | 11:40 | 2024-03-13 11:20 | 2024-03-13 11:40 | 0:20 | COMPLETED | SEMI-CRITICAL
1339 | 2024-03-13 | Mar 2024 | 13:00 | 13:03 | 0.0020833 | OT TABLE | MAQUET | ALPHA CLASSIC PRO | HAFIZ | OT | NOT ABLE TO FIX ABSORBER CANISTER | CHECKED AND FOUND ONE SCREW MISSING. TEMPORARILY TAKEN FROM ONE DOWN ANOESTHESIA MACHINE (00140), FIXED IN THIS MACHINE. NOW ISSUE RECTIFIED. | SCREW | NO | 2024-03-13 | March 2024 | 13:20 | 2024-03-13 13:00 | 2024-03-13 13:20 | 0:20 | COMPLETED | NON-CRITICAL
1340 | 2024-03-13 | Mar 2024 | 13:45 | 13:50 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AESTIVA MRI 7900 | HAFIZ | RADIOLOGY | TIDAL VOLUME ABNORMAL ISSUE | REINSERTED THE O2 CELL AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-03-13 | March 2024 | 14:05 | 2024-03-13 13:45 | 2024-03-13 14:05 | 0:20 | COMPLETED | CRITICAL
1341 | 2024-03-13 | Mar 2024 | 1:00 | 1:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | JERLIN | WARD | NOT WORKING | SETTINGS ADJUSTED. ISSUE RECTIFIED. | NO | NO | 2024-03-13 | March 2024 | 1:15 | 2024-03-13 01:00 | 2024-03-13 01:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1342 | 2024-03-13 | Mar 2024 | 14:10 | 14:20 | 0.0069444 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | RAMKIE | ICU | NO ECG WAVEFORM | CLEANED THE CONNECTOR AND ISSUE RECTIFIED | NO | NO | 2024-03-13 | March 2024 | 14:40 | 2024-03-13 14:10 | 2024-03-13 14:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1343 | 2024-03-14 | Mar 2024 | 19:00 | 19:10 | 0.0069444 | BLOOD GAS ANALYSER | GEM 3500 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | CATRIDGE NEED TO REPLACE | NEW CATRIDGE REPLACED | CARTRIDGE | NO | 2024-03-14 | March 2024 | 19:10 | 2024-03-14 19:00 | 2024-03-14 19:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1344 | 2024-03-14 | Mar 2024 | 7:15 | 7:20 | 0.0034722 | OT TABLE | MAQUET | MEERA | HAFIZ | OT | NOT WORKING | FOUND REMOTE LOCKED, UNLOCKED IT AND RESET DONE. NOW READY TO USE. | NO | NO | 2024-03-14 | March 2024 | 7:30 | 2024-03-14 07:15 | 2024-03-14 07:30 | 0:15 | COMPLETED | NON-CRITICAL
1345 | 2024-03-14 | Mar 2024 | 9:10 | 9:15 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | DIALYSIS | ECG EQUIPMENT MALFUNCTION | CLEANED THE CABLE AND REFIXED THE CABLE. PASSED THE OPERATIONAL CHECK. NOW READY TO USE. | NO | NO | 2024-03-14 | March 2024 | 9:25 | 2024-03-14 09:10 | 2024-03-14 09:25 | 0:15 | COMPLETED | CRITICAL
1346 | 2024-03-14 | Mar 2024 | 12:20 | 12:22 | 0.0013889 | RADIANT HEAT WARMER | WIPRO GE HEALTH CARE | LULLABY | HAFIZ | OT | FRONT WINDOW PANEL BROKEN | TEMPORARILY FIXED THE PANEL. NOW READY TO USE. | NO | NO | 2024-03-14 | March 2024 | 12:45 | 2024-03-14 12:20 | 2024-03-14 12:45 | 0:25 | COMPLETED | SEMI-CRITICAL
1347 | 2024-03-14 | Mar 2024 | 15:00 | 15:10 | 0.0069444 | ANESTHESIA MACHINE | MINDRAY | WATO EX-35 | HAFIZ | OT | CALIBRATION DUE | O2 CALIBRATION DONE. NOW READY TO USE. | NO | NO | 2024-03-14 | March 2024 | 15:20 | 2024-03-14 15:00 | 2024-03-14 15:20 | 0:20 | COMPLETED | CRITICAL
1348 | 2024-03-14 | Mar 2024 | 10:00 | 10:10 | 0.0069444 | C-ARM MACHINE | WIPRO GE HEALTH CARE | OEC ONE | HAFIZ | OT | MONOBLOCK ISSUE | REPLACED MONOBLOCK BY VENDOR. NOW ISSUE RECTIFIED. | MONOBLOCK | NO | 2024-03-14 | March 2024 | 14:00 | 2024-03-14 10:00 | 2024-03-14 14:00 | 4:00 | COMPLETED | CRITICAL
1349 | 2024-03-15 | Mar 2024 | 11:40 | 11:50 | 0.0069444 | PATIENT MONITOR | PHILIPS | CM12 | SANTHOSH | WARD | CONTINEOUS ALAMING | ALARM MUTED | NO | NO | 2024-03-16 | March 2024 | 18:20 | 2024-03-15 11:40 | 2024-03-16 18:20 | 30:40 | COMPLETED | SEMI-CRITICAL
1350 | 2024-03-15 | Mar 2024 | 14:45 | 14:55 | 0.0069444 | ACT MACHINE | HELENO LABORATORIES | ACTALYKE | RAMKIE | CATHLAB | ERROR | CHECKED AND SUSPECTED THE TUBE ISSUES | NO | NO | 2024-03-15 | March 2024 | 15:10 | 2024-03-15 14:45 | 2024-03-15 15:10 | 0:25 | COMPLETED | SEMI-CRITICAL
1351 | 2024-03-16 | Mar 2024 | 8:00 | 8:05 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | HAFIZ | OT | POLE CLAMP LOOSENED | TIGHTENED IT AND CHECKED. NOW READY TO USE. | NO | NO | 2024-03-16 | March 2024 | 8:30 | 2024-03-16 08:00 | 2024-03-16 08:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1352 | 2024-03-16 | Mar 2024 | 8:25 | 8:30 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | HAFIZ | OPD | PACS ISSUE | RECONNECTED THE LAN CABLE AND ISSUE RECTIFIED. | NO | NO | 2024-03-16 | March 2024 | 8:35 | 2024-03-16 08:25 | 2024-03-16 08:35 | 0:10 | COMPLETED | SEMI-CRITICAL
1353 | 2024-03-16 | Mar 2024 | 8:35 | 8:35 | 0 | BMI MACHINE | AKAS MEDICAL | QVS-100 | HAFIZ | OPD | UNOBLE TO MEASURE NIBP | SWAPPED THE BP HOSE FROM PROCEDURE ROOM. NOW ISSUE RECTIFIED. | NIBP HOSE | NO | 2024-03-16 | March 2024 | 8:45 | 2024-03-16 08:35 | 2024-03-16 08:45 | 0:10 | COMPLETED | NON-CRITICAL
1354 | 2024-03-16 | Mar 2024 | 10:30 | 10:32 | 0.0013889 | BLOOD GAS ANALYSER | GEM 3500 | INSTRUMENTATION LABORATORY | HAFIZ | ICU | CARTRIDGE FAILED | THROUGH COMPANY PERSON'S GUIDELINES, BASIC TROUBLESHOOTING DONE. ISSUE RECTIFIED. | NO | NO | 2024-03-16 | March 2024 | 11:00 | 2024-03-16 10:30 | 2024-03-16 11:00 | 0:30 | COMPLETED | SEMI-CRITICAL
1355 | 2024-03-16 | Mar 2024 | 15:00 | 15:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | RAMKIE | RADIOLOGY | ECG MALFUNCTION | RECONNECTED AND CHECK PASSED. NOW WORKING | NO | NO | 2024-03-16 | March 2024 | 15:30 | 2024-03-16 15:00 | 2024-03-16 15:30 | 0:30 | COMPLETED | CRITICAL
1356 | 2024-03-17 | Mar 2024 | 9:30 | 9:35 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SOORAJ | ICU | PRINTER NOT WORKING | CLEANED PRINTER HEAD AND ISSUE RECTIFIED | NO | NO | 2024-03-17 | March 2024 | 9:50 | 2024-03-17 09:30 | 2024-03-17 09:50 | 0:20 | COMPLETED | CRITICAL
1357 | 2024-03-17 | Mar 2024 | 10:30 | 10:35 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SOORAJ | ICU | CARTRIDGE EMPTY | REPLACED NEW CARTRIDGE AND ISSUE RECTIFIED | CARTRIDGE | NO | 2024-03-17 | March 2024 | 13:00 | 2024-03-17 10:30 | 2024-03-17 13:00 | 2:30 | COMPLETED | SEMI-CRITICAL
1358 | 2024-03-17 | Mar 2024 | 2:00 | 2:05 | 0.0034722 | AUTOMATED MICROBIAL ANALYSER | BIOMERIUX | VITEK 2 COMPACT 30 | SOORAJ | LABORATORY | SERVICE ALARM | CLEARED LOGS AND RESTARTED THE SYSTEM, ISSUE RECTIFIED | NO | NO | 2024-03-17 | March 2024 | 2:25 | 2024-03-17 02:00 | 2024-03-17 02:25 | 0:25 | COMPLETED | SEMI-CRITICAL
1359 | 2024-03-17 | Mar 2024 | 23:20 | 23:30 | 0.0069444 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | SANTHOSH | EMERGENCY | BULB ELECRODE BROKEN | REPLACED ANOTHER BULB | BULB ELECTRODE | NO | 2024-03-17 | March 2024 | 23:30 | 2024-03-17 23:20 | 2024-03-17 23:30 | 0:10 | COMPLETED | SEMI-CRITICAL
1360 | 2024-03-18 | Mar 2024 | 9:00 | 9:02 | 0.0013889 | USG MACHINE | WIPRO GE HEALTH CARE | VOLUSON S8 | HAFIZ | OPD | PRINTER NOT WORKING | FOUND ISSUE WITH POWER SUPPLY. PLUGGED TO AC SOURCE AND RESTARTED THE MACHINE. ISSUE RECTIFIED. | NO | NO | 2024-03-18 | March 2024 | 21:15 | 2024-03-18 09:00 | 2024-03-18 21:15 | 12:15 | COMPLETED | SEMI-CRITICAL
1361 | 2024-03-18 | Mar 2024 | 15:55 | 16:00 | 0.0034722 | LITHOCLAST | SURGIMEDIK HEALTHCARE | DIGI NOVA | HAFIZ | OT | HANDPIECE NOT WORKING | CONNECTED A WASHER AND CHECKED. ISSUE RECTIFIED. | WASHER | NO | 2024-03-18 | March 2024 | 16:10 | 2024-03-18 15:55 | 2024-03-18 16:10 | 0:15 | COMPLETED | SEMI-CRITICAL
1362 | 2024-03-18 | Mar 2024 | 15:20 | 15:30 | 0.0069444 | DIALYSIS MACHINE | FRESENIUS | 4008 S NG | RAMKIE | DIALYSIS | V99 FAILURE | CLEANED THE FILTER AND ISSUE RECTIFIED. | NO | NO | 2024-03-18 | March 2024 | 15:50 | 2024-03-18 15:20 | 2024-03-18 15:50 | 0:30 | COMPLETED | CRITICAL
1363 | 2024-03-19 | Mar 2024 | 7:00 | 7:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | REDUCED SODA LIME QTY, CHECKOUT DONE ISSUE RECTIFIED. | NO | NO | 2024-03-19 | March 2024 | 7:50 | 2024-03-19 07:00 | 2024-03-19 07:50 | 0:50 | COMPLETED | CRITICAL
1364 | 2024-03-19 | Mar 2024 | 10:30 | 10:35 | 0.0034722 | MICROSCOPE OPERATING, SURGICAL | LEICA MICROSYSTEMS | LEICA OHX | SOORAJ | OT | SLAVE MONITOR NOT SHOWING | SYSTEM RESET DONE USING RECOVERY MODE AND ISSUE RECTIFIED | NO | NO | 2024-03-19 | March 2024 | 11:00 | 2024-03-19 10:30 | 2024-03-19 11:00 | 0:30 | COMPLETED | SEMI-CRITICAL
1365 | 2024-03-19 | Mar 2024 | 13:00 | 13:10 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | SANTHOSH | OT | CIRCUIT LEAK TEST FAILED | ISSUE RECTIFIED | NO | NO | 2024-03-19 | March 2024 | 13:20 | 2024-03-19 13:00 | 2024-03-19 13:20 | 0:20 | COMPLETED | CRITICAL
1366 | 2024-03-19 | Mar 2024 | 8:00 | 8:05 | 0.0034722 | EXAMINATION COUCH | MIDMARK JANOK | E122012 | HAFIZ | OPD | NOT WORKING | RESET DONE AND OBSERVED. NOW READY TO USE. | NO | NO | 2024-03-19 | March 2024 | 8:15 | 2024-03-19 08:00 | 2024-03-19 08:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1367 | 2024-03-19 | Mar 2024 | 16:20 | 16:30 | 0.0069444 | FREEZER,DEEP FREEZING -40 | THERMO FISHER SCIENTIFIC | FORMA 7000 SERIES | RAMKIE | BLOODBANK | CHART RECORDER NOT WORKING | MOTOR SERVICED BY VENDOR AND ISSUE RECTIFIED | NO | NO | 2024-03-19 | March 2024 | 16:40 | 2024-03-19 16:20 | 2024-03-19 16:40 | 0:20 | COMPLETED | SEMI-CRITICAL
1368 | 2024-03-20 | Mar 2024 | 16:00 | 16:10 | 0.0069444 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX 8 | SANTHOSH | OT | ERROR 212 | RESET DONE ,NOW WORKING | NO | NO | 2024-03-20 | March 2024 | 21:20 | 2024-03-20 16:00 | 2024-03-20 21:20 | 5:20 | COMPLETED | SEMI-CRITICAL
1369 | 2024-03-20 | Mar 2024 | 10:00 | 10:03 | 0.0020833 | CAMERA SYSTEM | KARL STORZ | TC200 | HAFIZ | OT | UNABLE TO RECORD | ISSUE WITH STORAGE DEVICE, SO TEMPORARILY ARRANGED PENDRIVE FOR RECORDING. NOW ISSUE RECTIFIED. | PENDRIVE | NO | 2024-03-20 | March 2024 | 10:15 | 2024-03-20 10:00 | 2024-03-20 10:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1370 | 2024-03-20 | Mar 2024 | 11:00 | 11:05 | 0.0034722 | CAMERA SYSTEM | KARL STORZ | TP100EN | HAFIZ | OPD | DISPLAY ISSUE | SETTINGS ADJUSTED AND ISSUE RECTIFIED. | NO | NO | 2024-03-20 | March 2024 | 11:15 | 2024-03-20 11:00 | 2024-03-20 11:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1371 | 2024-03-20 | Mar 2024 | 15:00 | 15:03 | 0.0020833 | DYONICS POWER CONTROL UNIT,ARTHROSCOPY | SMITH & NEPHEW | 72200873 | HAFIZ | OT | OSCILLATE BUTTON NOT WORKING | CLEANED THE DUST IN THE BUTTONS. CHECKED AND ISSUE RECTIFIED. | NO | NO | 2024-03-20 | March 2024 | 15:30 | 2024-03-20 15:00 | 2024-03-20 15:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1372 | 2024-03-20 | Mar 2024 | 14:00 | 14:05 | 0.0034722 | OT TABLE | MAQUET | ALPHA MAX | HAFIZ | OT | SIDE ARM LOOSENED | TIGHTENED THE SIDE ARM AND ISSUE RECTIFIED. | NO | NO | 2024-03-20 | March 2024 | 14:20 | 2024-03-20 14:00 | 2024-03-20 14:20 | 0:20 | COMPLETED | NON-CRITICAL
1373 | 2024-03-20 | Mar 2024 | 15:00 | 15:10 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | CVP FAILURE | CHANGES DONE ACCORDINGLY AND ISSUE RECTIFIED | NO | NO | 2024-03-20 | March 2024 | 15:30 | 2024-03-20 15:00 | 2024-03-20 15:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1374 | 2024-03-20 | Mar 2024 | 10:20 | 10:30 | 0.0069444 | FREEZER,DEEP FREEZING -80 | THERMO FISHER SCIENTIFIC | FORMA 900 SERIES | RAMKIE | BLOODBANK | CHART RECORDER NOT WORKING | MOTOR SERVICED BY VENDOR AND ISSUE RECTIFIED | NO | NO | 2024-03-20 | March 2024 | 10:40 | 2024-03-20 10:20 | 2024-03-20 10:40 | 0:20 | COMPLETED | SEMI-CRITICAL
1375 | 2024-03-21 | Mar 2024 | 7:45 | 7:50 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | HAFIZ | OPD | NOT WORKING | SETTINGS ADJUSTED AND ISSUE RECTIFIED. | NO | NO | 2024-03-21 | March 2024 | 8:00 | 2024-03-21 07:45 | 2024-03-21 08:00 | 0:15 | COMPLETED | SEMI-CRITICAL
1376 | 2024-03-21 | Mar 2024 | 9:00 | 9:04 | 0.0027778 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | HAFIZ | OT | NOT WORKING | CHECKED AND FOUND THAT THERMISTOR RECEPTACLE FOUND DETACHED. REFIXED AND ISSUE RECTIFIED. | NO | NO | 2024-03-21 | March 2024 | 9:15 | 2024-03-21 09:00 | 2024-03-21 09:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1377 | 2024-03-21 | Mar 2024 | 9:40 | 9:45 | 0.0034722 | CAVITRON ULTRASONIC SURGICAL ASPIRATOR | INTEGRA | CUSA EXCEL + | HAFIZ | OT | ALARM INDICATION | CHANGED THE DISTILLED WATER. NOW ISSUE RECTIFIED | NO | NO | 2024-03-21 | March 2024 | 10:00 | 2024-03-21 09:40 | 2024-03-21 10:00 | 0:20 | COMPLETED | CRITICAL
1378 | 2024-03-21 | Mar 2024 | 11:00 | 11:05 | 0.0034722 | EXAMINATION COUCH | MIDMARK JANOK | E122012 | HAFIZ | OPD | NOT WORKING | RESET DONE AND OBSERVED. NOW READY TO USE. | NO | NO | 2024-03-21 | March 2024 | 11:15 | 2024-03-21 11:00 | 2024-03-21 11:15 | 0:15 | COMPLETED | NON-CRITICAL
1379 | 2024-03-21 | Mar 2024 | 11:10 | 11:20 | 0.0069444 | OPTICAL COHERENCE TOMOGRAPHY, 3 DIMENSIONOL | TOPCON | 3D OCT-1 | RAMKIE | OPD | PRINTING NOT GOOD | CHECKED AND PRINTING RECOVERY PROCESS REPEATED AND ISSUE RECTIFIED | NO | NO | 2024-03-21 | March 2024 | 11:40 | 2024-03-21 11:10 | 2024-03-21 11:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1380 | 2024-03-22 | Mar 2024 | 8:30 | 8:35 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | HAFIZ | WARD | WAVEFORM NOT SHOWING | CABLE FIXED PROPERLY. ISSUE RECTIFIED. | NO | NO | 2024-03-22 | March 2024 | 8:45 | 2024-03-22 08:30 | 2024-03-22 08:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1381 | 2024-03-22 | Mar 2024 | 10:30 | 10:35 | 0.0034722 | OXYGEN THERAPY UNIT | FISHER AND PAYKEL | AIRVO 2 | HAFIZ | ICU | POWERCORD ISSUE | REPLACED THE POWERCORD AND ISSUE RECTIFIED. | POWERCORD | NO | 2024-03-22 | March 2024 | 11:00 | 2024-03-22 10:30 | 2024-03-22 11:00 | 0:30 | COMPLETED | SEMI-CRITICAL
1382 | 2024-03-23 | Mar 2024 | 7:15 | 7:20 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | REMOVED AND REFIXED VAPORISER PROPERLY. ISSUE RECTIFEID | NO | NO | 2024-03-23 | March 2024 | 7:50 | 2024-03-23 07:15 | 2024-03-23 07:50 | 0:35 | COMPLETED | CRITICAL
1383 | 2024-03-23 | Mar 2024 | 8:00 | 8:04 | 0.0027778 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | ICU | SYSTEM FAILURE; EQUIPMENT DISBALED | RERUNNED THE OPERATIONAL CHECK IN SERVICE MODE. PASSED AND ISSUE RECTIFIED. | NO | NO | 2024-03-23 | March 2024 | 8:20 | 2024-03-23 08:00 | 2024-03-23 08:20 | 0:20 | COMPLETED | CRITICAL
1384 | 2024-03-23 | Mar 2024 | 10:00 | 10:04 | 0.0027778 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | HAFIZ | ICU | NOT WORKING | HEATER WIRE REPAIRED. NOW WORKING GOOD. | NO | NO | 2024-03-23 | March 2024 | 10:20 | 2024-03-23 10:00 | 2024-03-23 10:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1385 | 2024-03-23 | Mar 2024 | 10:25 | 10:30 | 0.0034722 | NEBULIZER | PARI | PARI BOY SX | HAFIZ | ICU | NOT WORKING | CONNECTOR ISSUE RECTIFIED. | NO | NO | 2024-03-23 | March 2024 | 10:40 | 2024-03-23 10:25 | 2024-03-23 10:40 | 0:15 | COMPLETED | SEMI-CRITICAL
1386 | 2024-03-23 | Mar 2024 | 21:00 | 21:05 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | NEED TO CHANGE A NEW CARTRIDGE | CHANGED A NEW CARTRIDGE AND CVP DONE | CARTRIDGE | NO | 2024-03-23 | March 2024 | 21:30 | 2024-03-23 21:00 | 2024-03-23 21:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1387 | 2024-03-23 | Mar 2024 | 1:00 | 1:03 | 0.0020833 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | JERLIN | WARD | SPO2 PROBE ISSUE | PIN STRAIGHTENED. ISSUE RECTIFIED. | NO | NO | 2024-03-23 | March 2024 | 1:25 | 2024-03-23 01:00 | 2024-03-23 01:25 | 0:25 | COMPLETED | SEMI-CRITICAL
1388 | 2024-03-23 | Mar 2024 | 1:25 | 1:30 | 0.0034722 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | JERLIN | WARD | SPO2 CONNECTOR ISSUE | ADJUSTED THE CONNECTOR. ISSUE RECTIFIED. | NO | NO | 2024-03-23 | March 2024 | 1:50 | 2024-03-23 01:25 | 2024-03-23 01:50 | 0:25 | COMPLETED | SEMI-CRITICAL
1389 | 2024-03-23 | Mar 2024 | 15:20 | 15:30 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | RAMKIE | CATHLAB | NOT READY TO USE | RERUNNED THE OPERATIONAL CHECK IN SERVICE MODE. PASSED AND ISSUE RECTIFIED. | NO | NO | 2024-03-23 | March 2024 | 15:40 | 2024-03-23 15:20 | 2024-03-23 15:40 | 0:20 | COMPLETED | CRITICAL
1390 | 2024-03-24 | Mar 2024 | 23:20 | 23:25 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | JERLIN | WARD | NOT WORKING | BATTERY DRAINED. PLUGGED TO AC SOURCE AND RECTIFIED. | NO | NO | 2024-03-24 | March 2024 | 23:35 | 2024-03-24 23:20 | 2024-03-24 23:35 | 0:15 | COMPLETED | SEMI-CRITICAL
1391 | 2024-03-24 | Mar 2024 | 1:00 | 1:10 | 0.0069444 | BIPAP MACHINE | PHILIPS | A40 | JERLIN | WARD | NOT WORKING | SETTINGS ADJUSTED. ISSUE RECTIFIED. | NO | NO | 2024-03-24 | March 2024 | 1:15 | 2024-03-24 01:00 | 2024-03-24 01:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1392 | 2024-03-25 | Mar 2024 | 14:00 | 14:03 | 0.0020833 | PATIENT MONITOR | PHILIPS | MX 450 | HAFIZ | ICU | UNOBLE TO MEASURE ECG AND RESP RATE | REFIXED THE PROBE AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-03-25 | March 2024 | 14:10 | 2024-03-25 14:00 | 2024-03-25 14:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1393 | 2024-03-25 | Mar 2024 | 14:10 | 14:10 | 0 | VENTILATOR, ADULT | MAQUET | SERVO-I | HAFIZ | ICU | PRE-USE CHECK FAILED | CLEANED THE EXPIRATORY CASSETTE AND RERUNNED THE CHECK. ISSUE RECTIFIED. | NO | NO | 2024-03-25 | March 2024 | 14:30 | 2024-03-25 14:10 | 2024-03-25 14:30 | 0:20 | COMPLETED | CRITICAL
1394 | 2024-03-26 | Mar 2024 | 16:20 | 16:30 | 0.0069444 | CATHLAB - MONOPLANE | PHILIPS | AZURION 7M12 | RAMKIE | CATHLAB | TABLE FOOT SWITCH PROBLEM | CHANGED THE CONNECTORS AND CABLES AND ISSUE RECTIFIED | CONNECTORS AND CABLES | NO | 2024-03-26 | March 2024 | 16:40 | 2024-03-26 16:20 | 2024-03-26 16:40 | 0:20 | COMPLETED | CRITICAL
1395 | 2024-03-27 | Mar 2024 | 7:55 | 8:00 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | FOUND MOISTURE IN O2 CELL, CLEANED AND ISSUE RECTIFIED. | NO | NO | 2024-03-27 | March 2024 | 8:20 | 2024-03-27 07:55 | 2024-03-27 08:20 | 0:25 | COMPLETED | CRITICAL
1396 | 2024-03-27 | Mar 2024 | 9:40 | 9:45 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | BELLOW ISSUE | REMOVED AND REFIXED BELLOWS, CHECKOUT DONE AND ISSUE RECTIFIED | NO | NO | 2024-03-27 | March 2024 | 10:00 | 2024-03-27 09:40 | 2024-03-27 10:00 | 0:20 | COMPLETED | CRITICAL
1397 | 2024-03-27 | Mar 2024 | 10:35 | 10:40 | 0.0034722 | GO FLO PUMP, ARTHROSCOPY | SMITH & NEPHEW | 72204968 | SOORAJ | OT | SWITCHED OFF AUTOMATICALLY | FOUND TUBING COIL HEATED, INTERNAL BOARD ISSUE | NO | INTERNAL BOARD | 2024-03-27 | March 2024 | 12:30 | 2024-03-27 10:35 | 2024-03-27 12:30 | 1:55 | PENDING | SEMI-CRITICAL
1398 | 2024-03-27 | Mar 2024 | 12:00 | 12:05 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SOORAJ | OT | NOT CHARGING. | FOUND ISSUE WITH POWER SOCKET, MACHINE WORKING FINE | NO | NO | 2024-03-27 | March 2024 | 12:15 | 2024-03-27 12:00 | 2024-03-27 12:15 | 0:15 | COMPLETED | CRITICAL
1399 | 2024-03-27 | Mar 2024 | 15:00 | 15:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SOORAJ | OT | CANNOT PRESSURIZE CIRCUIT ALARM | FOUND GAS ANALYZER CHECK SKIPPED. CHECKOUT DONE AND ISSUE RECTIFIED. | NO | NO | 2024-03-27 | March 2024 | 15:30 | 2024-03-27 15:00 | 2024-03-27 15:30 | 0:30 | COMPLETED | CRITICAL
1400 | 2024-03-27 | Mar 2024 | 15:30 | 15:30 | 0 | CORE CONSOLE WITH HANDPIECE | STRYKER | 5400-052-000 | SOORAJ | OT | PERFORATOR CHUCK NOT WORKING | FOUND INNTERANLY WORN OUT, NOT REPAIRABLE. | NO | NO | 2024-03-27 | March 2024 | 15:50 | 2024-03-27 15:30 | 2024-03-27 15:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1401 | 2024-03-27 | Mar 2024 | 15:00 | 15:10 | 0.0069444 | TREADMILL TEST | AFTON | AFTON | SANTHOSH | OPD | DDNT GET SIGNOLS | ISSUE RECTIFIED | NO | NO | 2024-03-28 | March 2024 | 15:20 | 2024-03-27 15:00 | 2024-03-28 15:20 | 24:20 | COMPLETED | NON-CRITICAL
1402 | 2024-03-29 | Mar 2024 | 11:10 | 11:20 | 0.0069444 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | RAMKIE | OPD | NOT WORKING | FOUND PROBLEM IN SSD. | NO | NO | 2024-03-29 | March 2024 | 11:40 | 2024-03-29 11:10 | 2024-03-29 11:40 | 0:30 | COMPLETED | NON-CRITICAL
1403 | 2024-03-30 | Mar 2024 | 11:00 | 11:19 | 0.013194 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SANTHOSH | WARD | PRINTER ISSUE | CLEANED THE PRINTER HEAD AND RECTIFIED | NO | NO | 2024-03-30 | March 2024 | 11:20 | 2024-03-30 11:00 | 2024-03-30 11:20 | 0:20 | COMPLETED | CRITICAL
1404 | 2024-03-30 | Mar 2024 | 13:35 | 13:40 | 0.0034722 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | SANTHOSH | OPD | HARD DISK FAILIURE | REPLACED NEW SSD | HARDDISK | NO | 2024-03-30 | March 2024 | 13:40 | 2024-03-30 13:35 | 2024-03-30 13:40 | 0:05 | COMPLETED | NON-CRITICAL
1405 | 2024-03-30 | Mar 2024 | 17:00 | 17:10 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | CATRIDGE DEFECT | NEW CATRIDGE REPLACED | CARTRIDGE | NO | 2024-03-30 | March 2024 | 17:10 | 2024-03-30 17:00 | 2024-03-30 17:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1406 | 2024-03-31 | Mar 2024 | 10:00 | 10:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SANTHOSH | DIALYSIS | ECG FAILIURE | OP CHECK DONE | NO | NO | 2024-03-31 | March 2024 | 10:05 | 2024-03-31 10:00 | 2024-03-31 10:05 | 0:05 | COMPLETED | CRITICAL
1407 | 2024-03-31 | Mar 2024 | 11:00 | 11:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SANTHOSH | ICU | THERAPHY KNOB FAILURE | OP CHECK DONE | NO | NO | 2024-03-31 | March 2024 | 11:05 | 2024-03-31 11:00 | 2024-03-31 11:05 | 0:05 | COMPLETED | CRITICAL
1408 | 2024-04-01 | Apr 2024 | 9:10 | 9:15 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | THANISH | OPD | DISPLAY NOT WORKING | CLEANED THE DISPLAY CABLE AND RECONNECTED IT.WORKING FINE. | NO | NO | 2024-04-01 | April 2024 | 9:40 | 2024-04-01 09:10 | 2024-04-01 09:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1409 | 2024-04-01 | Apr 2024 | 10:20 | 10:25 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | THANISH | ICU | NOT SWITCHING ON | RE-ATTACHED THE POWER CABLE TO CONTROL BOX.WORKING FINE. | NO | NO | 2024-04-01 | April 2024 | 10:45 | 2024-04-01 10:20 | 2024-04-01 10:45 | 0:25 | COMPLETED | NON-CRITICAL
1410 | 2024-04-01 | Apr 2024 | 14:00 | 14:10 | 0.0069444 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | HAFIZ | OPD | NOT SWITCHING ON | FOUND ISSUE WITH THE HARDDISK. REPLACED THE STANDBY HARDDISK AND ISSUE RECTIFIED. | HARDDISK | NO | 2024-04-01 | April 2024 | 16:45 | 2024-04-01 14:00 | 2024-04-01 16:45 | 2:45 | COMPLETED | NON-CRITICAL
1411 | 2024-04-01 | Apr 2024 | 17:00 | 17:02 | 0.0013889 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | ICU | HOLDER CLAMP NOT WORKING | REPLACED THE HOLDER CLAMP. ISSUE RECTIFIED. | CLAMP | NO | 2024-04-01 | April 2024 | 17:15 | 2024-04-01 17:00 | 2024-04-01 17:15 | 0:15 | COMPLETED | NON-CRITICAL
1412 | 2024-04-01 | Apr 2024 | 17:15 | 17:15 | 0 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | ICU | PRESSURE SENSOR BROKEN | REPLACED THE PRESSURE SENSOR. OBSERVED AND WORKING GOOD. | PRESSURE SENSOR | NO | 2024-04-01 | April 2024 | 17:35 | 2024-04-01 17:15 | 2024-04-01 17:35 | 0:20 | COMPLETED | NON-CRITICAL
1413 | 2024-04-01 | Apr 2024 | 12:30 | 12:35 | 0.0034722 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | SANTHOSH | OPD | SMALL WAVES ISSUE | REPLACED WITH 1ST FLOOR CABLE ISSUE RECTIFIED | NO | NO | 2024-04-01 | April 2024 | 12:40 | 2024-04-01 12:30 | 2024-04-01 12:40 | 0:10 | COMPLETED | NON-CRITICAL
1414 | 2024-04-01 | Apr 2024 | 13:40 | 13:45 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SANTHOSH | WARD | NOT WORKING | RESET DONE ,NOW WORKING | NO | NO | 2024-04-01 | April 2024 | 13:50 | 2024-04-01 13:40 | 2024-04-01 13:50 | 0:10 | COMPLETED | NON-CRITICAL
1415 | 2024-04-01 | Apr 2024 | 17:00 | 17:05 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | CADRIDGE EXPIRED | NEW CATRIDGE REPLACED | CARTRIDGE | NO | 2024-04-01 | April 2024 | 17:10 | 2024-04-01 17:00 | 2024-04-01 17:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1416 | 2024-04-01 | Apr 2024 | 9:40 | 9:50 | 0.0069444 | HOLTER,RECORDER | WIPRO GE HEALTH CARE | SEER 1000 | RAMKIE | OPD | ECG DISTURBANCE | CLEANED ECG CONNECTOR AND WORKING | NO | NO | 2024-04-01 | April 2024 | 10:10 | 2024-04-01 09:40 | 2024-04-01 10:10 | 0:30 | COMPLETED | SEMI-CRITICAL
1417 | 2024-04-01 | Apr 2024 | 16:10 | 16:15 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | NEED TO REPLACE CARTRIDGE | REPLACED THE CARTRIDGE AND RECTIFIED THE ISSUE | CARTRIDGE | NO | 2024-04-01 | April 2024 | 16:50 | 2024-04-01 16:10 | 2024-04-01 16:50 | 0:40 | COMPLETED | SEMI-CRITICAL
1418 | 2024-04-01 | Apr 2024 | 9:10 | 9:15 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | THANISH | OPD | DISPLAY NOT WORKING | CLEANED THE DISPLAY CABLE AND RECONNECTED IT.WORKING FINE. | NO | NO | 2024-04-01 | April 2024 | 9:40 | 2024-04-01 09:10 | 2024-04-01 09:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1419 | 2024-04-01 | Apr 2024 | 10:20 | 10:25 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | THANISH | ICU | NOT SWITCHING ON | RE-ATTACHED THE POWER CABLE TO CONTROL BOX.WORKING FINE. | NO | NO | 2024-04-01 | April 2024 | 10:45 | 2024-04-01 10:20 | 2024-04-01 10:45 | 0:25 | COMPLETED | NON-CRITICAL
1420 | 2024-04-02 | Apr 2024 | 10:05 | 10:10 | 0.0034722 | VENTILATOR, ADULT | MAQUET | SERVO-I | THANISH | ICU | O2 CELL NOT WORKING | CLEANED THE O2 CELL/RECALIBRATED IT.WORKING FINE. | NO | NO | 2024-04-02 | April 2024 | 10:35 | 2024-04-02 10:05 | 2024-04-02 10:35 | 0:30 | COMPLETED | CRITICAL
1421 | 2024-04-02 | Apr 2024 | 12:05 | 12:10 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | THANISH | OPD | SWITCHING OFF AUTOMATICALLY | RE-INSATLLED THE SOFTWARE | NO | NO | 2024-04-02 | April 2024 | 12:50 | 2024-04-02 12:05 | 2024-04-02 12:50 | 0:45 | COMPLETED | SEMI-CRITICAL
1422 | 2024-04-02 | Apr 2024 | 15:10 | 15:15 | 0.0034722 | RADIOGRAPHIC UNIT | CARESTREAM | DRF ASCEND ANOLOG HF | THANISH | RADIOLOGY | AUTOMODE NOT WORKING | RE-STARTED THE GENERATOR,DETECTOR&SYSTEM | NO | NO | 2024-04-02 | April 2024 | 15:50 | 2024-04-02 15:10 | 2024-04-02 15:50 | 0:40 | COMPLETED | SEMI-CRITICAL
1423 | 2024-04-02 | Apr 2024 | 15:20 | 15:25 | 0.0034722 | INTRA ORAL XRAY | SKANRAY DENTAL | INTRA SKANDC | SANTHOSH | OPD | CONTINEOUS XRAY PROBLEM | CHECKED WITH ANOTHER PORT WORKING | NO | NO | 2024-04-02 | April 2024 | 15:30 | 2024-04-02 15:20 | 2024-04-02 15:30 | 0:10 | COMPLETED | SEMI-CRITICAL
1424 | 2024-04-02 | Apr 2024 | 10:10 | 10:20 | 0.0069444 | AUTOMATED MICROBIAL ANALYSER | BIOMERIUX | VITEK 2 COMPACT 30 | RAMKIE | LABORATORY | NOT WORKING | SAMPLE CARD REMOVED, ADDED MANUAL AND PROBLEM RECTIFIED. | NO | NO | 2024-04-02 | April 2024 | 10:50 | 2024-04-02 10:10 | 2024-04-02 10:50 | 0:40 | COMPLETED | SEMI-CRITICAL
1425 | 2024-04-02 | Apr 2024 | 15:00 | 15:10 | 0.0069444 | MONITOR HIGH DEFINITION,3D | KARL STORZ | TM330 | JERLIN | OT | POWER ERROR | RESTARTED THE MACHINE AND RECTIFIED THE ISSUE | NO | NO | 2024-04-02 | April 2024 | 15:30 | 2024-04-02 15:00 | 2024-04-02 15:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1426 | 2024-04-02 | Apr 2024 | 10:00 | 10:15 | 0.010417 | QUANTUM MACHINE | SMITH & NEPHEW | RF12000 | JERLIN | OT | FOOT PEDAL NOT WORKING | SERVICED THE FOOT PEDAL AND NOW IT'S WORKING GOOD. | NO | NO | 2024-04-02 | April 2024 | 10:35 | 2024-04-02 10:00 | 2024-04-02 10:35 | 0:35 | COMPLETED | SEMI-CRITICAL
1427 | 2024-04-02 | Apr 2024 | 10:00 | 10:03 | 0.0020833 | VENTILATOR, ADULT | MAQUET | SERVO-I | THANISH | ICU | O2 CELL NOT WORKING | CLEANED THE O2 CELL/RECALIBRATED IT.WORKING FINE. | NO | NO | 2024-04-02 | April 2024 | 10:20 | 2024-04-02 10:00 | 2024-04-02 10:20 | 0:20 | COMPLETED | CRITICAL
1428 | 2024-04-02 | Apr 2024 | 12:05 | 12:10 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | THANISH | OPD | SWITCHING OFF AUTOMATICALLY | RE-INSATLLED THE SOFTWARE | NO | NO | 2024-04-02 | April 2024 | 12:50 | 2024-04-02 12:05 | 2024-04-02 12:50 | 0:45 | COMPLETED | SEMI-CRITICAL
1429 | 2024-04-02 | Apr 2024 | 15:10 | 15:15 | 0.0034722 | RADIOGRAPHIC UNIT | CARESTREAM | DRF ASCEND ANOLOG HF | THANISH | RADIOLOGY | AUTOMODE NOT WORKING | RE-STARTED THE GENERATOR,DETECTOR&SYSTEM | NO | NO | 2024-04-02 | April 2024 | 15:50 | 2024-04-02 15:10 | 2024-04-02 15:50 | 0:40 | COMPLETED | SEMI-CRITICAL
1430 | 2024-04-02 | Apr 2024 | 8:00 | 8:02 | 0.0013889 | MRI MACHINE | PHILIPS | INGENIA ELITON S | JERLIN | RADIOLOGY | GRADIENT MALFUNCTION ERROR | ERROR LOGS SEND TO TECHNICAL TEAM, CONFIRMED THE ERROR IS DUE TO SAFETY PRECAUTIONS AND NO HARDWARE ISSUES FOUND. NOW CLEARED THE ERROR FILES AND REBOOTED THE SYSTEM. WORKING GOOD. | NO | NO | 2024-04-02 | April 2024 | 11:00 | 2024-04-02 08:00 | 2024-04-02 11:00 | 3:00 | COMPLETED | CRITICAL
1431 | 2024-04-03 | Apr 2024 | 10:10 | 10:15 | 0.0034722 | EXAMINATION COUCH | MIDMARK JANOK | E122012 | THANISH | OPD | COUCH NOT WORKING | RE-SETTED THE CONTROL BOX | NO | NO | 2024-04-03 | April 2024 | 22:40 | 2024-04-03 10:10 | 2024-04-03 22:40 | 12:30 | COMPLETED | NON-CRITICAL
1432 | 2024-04-03 | Apr 2024 | 16:10 | 16:15 | 0.0034722 | BLOOD GAS ANALYSER | GEM 3500 | INSTRUMENTATION LABORATORY | THANISH | ICU | ABG CATRIDGE FAILED | REPLACED WITH NEW ABG CATRIDGE | CARTRIDGE | NO | 2024-04-03 | April 2024 | 16:50 | 2024-04-03 16:10 | 2024-04-03 16:50 | 0:40 | COMPLETED | SEMI-CRITICAL
1433 | 2024-04-03 | Apr 2024 | 21:00 | 21:05 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | SOORAJ | OT | LINEAR PROBE HANGING ISSUE | CHECKED PROBE SELF AND PASSED, OBSERVED AND FOUND NO HANGING ISSUE. | NO | NO | 2024-04-03 | April 2024 | 21:45 | 2024-04-03 21:00 | 2024-04-03 21:45 | 0:45 | COMPLETED | SEMI-CRITICAL
1434 | 2024-04-03 | Apr 2024 | 21:50 | 21:55 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC V2 | SOORAJ | EMERGENCY | SCAN PAGE NOT BOOTING UP | FOUND PROBE INSERTED INCORRECTLY. REMOVED AND INSERTED PROPERLY. ISSUE RECTIFIED. | NO | NO | 2024-04-03 | April 2024 | 22:30 | 2024-04-03 21:50 | 2024-04-03 22:30 | 0:40 | COMPLETED | SEMI-CRITICAL
1435 | 2024-04-03 | Apr 2024 | 10:20 | 10:30 | 0.0069444 | ELECTROTHERAPHY WITHOUT LASER | CHATTANOOGA | INTELECT ADVANCED | RAMKIE | PHYSIOTHERAPY | NO PROPER OUTPUT | CHECKED AND CONFIRMED PROBLEM WITH US CABLE, TAKEN TO COMPANY | NO | NO | 2024-04-03 | April 2024 | 10:45 | 2024-04-03 10:20 | 2024-04-03 10:45 | 0:25 | COMPLETED | SEMI-CRITICAL
1436 | 2024-04-03 | Apr 2024 | 10:40 | 10:45 | 0.0034722 | ELECTROTHERAPHY WITHOUT LASER | CHATTANOOGA | INTELECT ADVANCED | RAMKIE | PHYSIOTHERAPY | LOW US OUTPUT | CHECKED AND NEED TO DO CALIBRATION AND CRYSTAL REPLACEMENT | NO | NO | 2024-04-03 | April 2024 | 11:00 | 2024-04-03 10:40 | 2024-04-03 11:00 | 0:20 | COMPLETED | SEMI-CRITICAL
1437 | 2024-04-03 | Apr 2024 | 12:00 | 12:30 | 0.020833 | CAMERA SYSTEM,ICG | KARL STORZ | TC200 | JERLIN | OT | CAMERA HEAD NOT WORKING | REPLACED THE CAMERA CABLE'S SLEEVE AND RECTIFIED THE ISSUE | CAMERA CABLE | NO | 2024-04-03 | April 2024 | 12:40 | 2024-04-03 12:00 | 2024-04-03 12:40 | 0:40 | COMPLETED | SEMI-CRITICAL
1438 | 2024-04-03 | Apr 2024 | 15:00 | 15:20 | 0.013889 | INTELLIGENT HEARING SYSTEM | STARKEY | IHS DUET | JERLIN | OPD | NOT WORKING PROPERLY | CHECKED AND CLEARED THE AIRWAY AND NOW IT'S WORKING GOOD. | NO | NO | 2024-04-03 | April 2024 | 15:45 | 2024-04-03 15:00 | 2024-04-03 15:45 | 0:45 | COMPLETED | SEMI-CRITICAL
1439 | 2024-04-03 | Apr 2024 | 10:10 | 10:15 | 0.0034722 | EXAMINATION COUCH | MIDMARK JANOK | E122012 | THANISH | OPD | COUCH NOT WORKING | RE-SETTED THE CONTROL BOX | NO | NO | 2024-04-03 | April 2024 | 22:40 | 2024-04-03 10:10 | 2024-04-03 22:40 | 12:30 | COMPLETED | NON-CRITICAL
1440 | 2024-04-03 | Apr 2024 | 16:10 | 16:15 | 0.0034722 | BLOOD GAS ANALYSER | GEM 3500 | INSTRUMENTATION LABORATORY | THANISH | ICU | ABG CATRIDGE FAILED | REPLACED WITH NEW ABG CATRIDGE | CARTRIDGE | NO | 2024-04-03 | April 2024 | 16:50 | 2024-04-03 16:10 | 2024-04-03 16:50 | 0:40 | COMPLETED | SEMI-CRITICAL
1441 | 2024-04-04 | Apr 2024 | 9:10 | 9:15 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | THANISH | CATHLAB | NIBP NOT WORKING | RECTIFIED THE NIBP CUFF.WORKING FINE. | NO | NO | 2024-04-04 | April 2024 | 9:30 | 2024-04-04 09:10 | 2024-04-04 09:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1442 | 2024-04-04 | Apr 2024 | 11:10 | 11:15 | 0.0034722 | MAMMOGRAPH UNIT | HOLOGIC | SELENIA DIMENSIONS | THANISH | RADIOLOGY | DETECTOR NOT WORKING | RE-STARTED THE CPU.WORKING FINE | NO | NO | 2024-04-04 | April 2024 | 11:30 | 2024-04-04 11:10 | 2024-04-04 11:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1443 | 2024-04-04 | Apr 2024 | 16:10 | 16:15 | 0.0034722 | RADIANT HEAT WARMER | WIPRO GE HEALTH CARE | LULLABY | THANISH | OT | MACHINE NOT SWITCHING ON | REPLACED THE 2AMPS FUSE.WORKING FINE. | FUSE | NO | 2024-04-04 | April 2024 | 16:40 | 2024-04-04 16:10 | 2024-04-04 16:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1444 | 2024-04-04 | Apr 2024 | 1:00 | 1:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | O2 CELL ISSUE | CHECKED AND FOUND O2 SENSOR CABLE DEFECTIVE. SWAPPED FROM STANDBY MACHINE AND NOW WORKING FINE. | NO | NO | 2024-04-04 | April 2024 | 1:20 | 2024-04-04 01:00 | 2024-04-04 01:20 | 0:20 | COMPLETED | CRITICAL
1445 | 2024-04-04 | Apr 2024 | 2:00 | 2:05 | 0.0034722 | OT TABLE | MAQUET | ALPHA MAX | SOORAJ | OT | REMOTE NOT WORKING | FOUND INTERNOL CONNECTOR BROKEN. TAKEN TO DEPARTMENT FOR SERVICE. AS OF NOW TABLE FUNCTIONING ON MANUAL MODE | NO | NO | 2024-04-04 | April 2024 | 3:00 | 2024-04-04 02:00 | 2024-04-04 03:00 | 1:00 | COMPLETED | NON-CRITICAL
1446 | 2024-04-04 | Apr 2024 | 6:55 | 7:00 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SOORAJ | ICU | CARTRIDGE EMPTY MESSAGE | NEW CARTRIDGE REPLACED AND CP DONE. | CARTRIDGE | NO | 2024-04-04 | April 2024 | 8:15 | 2024-04-04 06:55 | 2024-04-04 08:15 | 1:20 | COMPLETED | SEMI-CRITICAL
1447 | 2024-04-04 | Apr 2024 | 14:04 | 14:10 | 0.0041667 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | HAFIZ | WARD | HOLDER CLAMP ISSUE | REPLACED THE HOLDER CLAMP. ISSUE RECTIFIED. | HOLDER CLAMP | NO | 2024-04-04 | April 2024 | 14:25 | 2024-04-04 14:04 | 2024-04-04 14:25 | 0:21 | COMPLETED | NON-CRITICAL
1448 | 2024-04-04 | Apr 2024 | 14:30 | 14:32 | 0.0013889 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | WARD | HOLDER CLAMP ISSUE | REPLACED THE HOLDER CLAMP. ISSUE RECTIFIED. | HOLDER CLAMP | NO | 2024-04-04 | April 2024 | 14:45 | 2024-04-04 14:30 | 2024-04-04 14:45 | 0:15 | COMPLETED | NON-CRITICAL
1449 | 2024-04-04 | Apr 2024 | 9:10 | 9:20 | 0.0069444 | ELECTROTHERAPHY WITH LASER | CHATTANOOGA | INTELECT ADVANCED | RAMKIE | PHYSIOTHERAPY | HIGH US OUTPUT | CHECKED AND NEED TO DO CALIBRATION AND CRYSTAL REPLACEMENT | NO | NO | 2024-04-04 | April 2024 | 9:35 | 2024-04-04 09:10 | 2024-04-04 09:35 | 0:25 | COMPLETED | SEMI-CRITICAL
1450 | 2024-04-04 | Apr 2024 | 9:10 | 9:15 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | THANISH | CATHLAB | NIBP NOT WORKING | RECTIFIED THE NIBP CUFF.WORKING FINE. | NO | NO | 2024-04-04 | April 2024 | 9:30 | 2024-04-04 09:10 | 2024-04-04 09:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1451 | 2024-04-04 | Apr 2024 | 11:10 | 11:15 | 0.0034722 | MAMMOGRAPH UNIT | HOLOGIC | SELENIA DIMENSIONS | THANISH | RADIOLOGY | DETECTOR NOT WORKING | RE-STARTED THE CPU.WORKING FINE | NO | NO | 2024-04-04 | April 2024 | 11:30 | 2024-04-04 11:10 | 2024-04-04 11:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1452 | 2024-04-04 | Apr 2024 | 16:10 | 16:15 | 0.0034722 | RADIANT HEAT WARMER | WIPRO GE HEALTH CARE | LULLABY | THANISH | OT | MACHINE NOT SWITCHING ON | REPLACED THE 2AMPS FUSE.WORKING FINE. | FUSE | NO | 2024-04-04 | April 2024 | 16:40 | 2024-04-04 16:10 | 2024-04-04 16:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1453 | 2024-04-05 | Apr 2024 | 20:30 | 20:35 | 0.0034722 | CT MACHINE | PHILIPS | INGENUITY | SOORAJ | RADIOLOGY | ECG WAVEFORM NOT SHWOING PROPERLY | BASE BOARD ASSEMBLY REPLACED IN CT AND CHECKED, NOW WAVEFORM SHOPWING PROPERLY. | BASE BOARD ASSEMBLY | NO | 2024-04-05 | April 2024 | 22:00 | 2024-04-05 20:30 | 2024-04-05 22:00 | 1:30 | COMPLETED | CRITICAL
1454 | 2024-04-05 | Apr 2024 | 8:25 | 8:30 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | HAFIZ | EMERGENCY | BUTTONS NOT WORKING | POWERCORD DETACHED FROM ACTUATOR.REFIXED IT AND ISSUE RECTIFIED. | NO | NO | 2024-04-05 | April 2024 | 8:45 | 2024-04-05 08:25 | 2024-04-05 08:45 | 0:20 | COMPLETED | NON-CRITICAL
1455 | 2024-04-05 | Apr 2024 | 8:45 | 8:47 | 0.0013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | HAFIZ | ICU | ONE SIDE REEL NOT WORKING | NEED CONNECTOR REPLACEMENT. | NO | CONNECTOR | 2024-04-05 | April 2024 | 8:50 | 2024-04-05 08:45 | 2024-04-05 08:50 | 0:05 | PENDING | NON-CRITICAL
1456 | 2024-04-05 | Apr 2024 | 10:10 | 10:15 | 0.0034722 | BMD MACHINE | HOLOGIC | ASY-05119 | HAFIZ | RADIOLOGY | DAILY QC TEST FAILED | CONSULTED THE SERVICE ENGINEER AND TEST AFTER BASIC TROUBLESHOOTING, QC TEST PASSED. NOW WORKING GOOD. | NO | NO | 2024-04-05 | April 2024 | 10:50 | 2024-04-05 10:10 | 2024-04-05 10:50 | 0:40 | COMPLETED | SEMI-CRITICAL
1457 | 2024-04-05 | Apr 2024 | 16:20 | 16:25 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | HAFIZ | WARD | PRINTER NOT WORKING | ISSUE WITH PAPER ROLL. REPLACED THE PAPER ROLL. ISSUE RECTIFIED. | PAPER ROLL | NO | 2024-04-05 | April 2024 | 16:35 | 2024-04-05 16:20 | 2024-04-05 16:35 | 0:15 | COMPLETED | SEMI-CRITICAL
1458 | 2024-04-05 | Apr 2024 | 14:20 | 14:25 | 0.0034722 | DIGITAL PLATFORM SCALE | SECA | 676 | SANTHOSH | DIALYSIS | 1.16 ,1.17 ERROR | TERMPORARILY WORKING | NO | NO | 2024-04-05 | April 2024 | 14:30 | 2024-04-05 14:20 | 2024-04-05 14:30 | 0:10 | COMPLETED | NON-CRITICAL
1459 | 2024-04-05 | Apr 2024 | 9:20 | 9:30 | 0.0069444 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | RAMKIE | OT | CLAMP NOT GOOD | CLAMP REPLACED AND WORKING | CLAMP | NO | 2024-04-05 | April 2024 | 9:35 | 2024-04-05 09:20 | 2024-04-05 09:35 | 0:15 | COMPLETED | NON-CRITICAL
1460 | 2024-04-05 | Apr 2024 | 7:30 | 7:40 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | JERLIN | OT | CALIBRATION DUE | RECHECKED AND NOW IT'S WORKING GOOD | NO | NO | 2024-04-05 | April 2024 | 7:50 | 2024-04-05 07:30 | 2024-04-05 07:50 | 0:20 | COMPLETED | CRITICAL
1461 | 2024-04-05 | Apr 2024 | 8:40 | 8:50 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | JERLIN | OT | CIRCUIT LEAK TEST FAILED | FIXED THE CANISTER PROPERLY AND NOW IT'S WORKING GOOD | NO | NO | 2024-04-05 | April 2024 | 9:00 | 2024-04-05 08:40 | 2024-04-05 09:00 | 0:20 | COMPLETED | CRITICAL
1462 | 2024-04-05 | Apr 2024 | 10:00 | 10:10 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | JERLIN | OT | CIRCUIT LEAK TEST FAILED | CHECKED WITH ANOTHER CIRCUIT, REPLACED IT AND NOW IT'S WORKING GOOD. | NO | NO | 2024-04-05 | April 2024 | 10:20 | 2024-04-05 10:00 | 2024-04-05 10:20 | 0:20 | COMPLETED | CRITICAL
1463 | 2024-04-05 | Apr 2024 | 15:00 | 15:10 | 0.0069444 | OT TABLE | MAQUET | ALPHA MAX | JERLIN | OT | REMOTE NOT WORKING | REPLACED THE CONNECTOR JACK AND NOW IT'S WORKING GOOD. | CABLE CONNECTOR | NO | 2024-04-05 | April 2024 | 15:24 | 2024-04-05 15:00 | 2024-04-05 15:24 | 0:24 | COMPLETED | NON-CRITICAL
1464 | 2024-04-06 | Apr 2024 | 10:05 | 10:10 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | THANISH | OT | THERAPY DIS-ABLED ISSUE | PERFORMED THE OP CHECK | NO | NO | 2024-04-06 | April 2024 | 10:30 | 2024-04-06 10:05 | 2024-04-06 10:30 | 0:25 | COMPLETED | CRITICAL
1465 | 2024-04-06 | Apr 2024 | 14:00 | 14:04 | 0.0027778 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | WARD | PRINTER NOT WORKING | FOUND ISSUE WITH THE PRINTER HEAD. REPLACED THE PRINTER HEAD FROM DEFECTIVE MONITOR. OBSERVED AND WORKING GOOD. | PRINTER HEAD | NO | 2024-04-06 | April 2024 | 14:10 | 2024-04-06 14:00 | 2024-04-06 14:10 | 0:10 | COMPLETED | CRITICAL
1466 | 2024-04-06 | Apr 2024 | 8:20 | 8:20 | 0 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | HAFIZ | ICU | NEED TO REPLACE CARTRIDGE | INSTALLED NEW CARTRDIGE (S.NO: 301863073). USED CARTRIDGE (S.NO: 301863067). WARMING UP DONE. READY TO USE. | CARTRIDGE | NO | 2024-04-06 | April 2024 | 8:50 | 2024-04-06 08:20 | 2024-04-06 08:50 | 0:30 | COMPLETED | SEMI-CRITICAL
1467 | 2024-04-06 | Apr 2024 | 9:04 | 9:06 | 0.0013889 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | HAFIZ | OPD | PROBE NOT WORKING | PROPERLY FIXED THE PROBE AND ISSUE RECTIFIED. | NO | NO | 2024-04-06 | April 2024 | 9:10 | 2024-04-06 09:04 | 2024-04-06 09:10 | 0:06 | COMPLETED | SEMI-CRITICAL
1468 | 2024-04-06 | Apr 2024 | 12:05 | 12:10 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | HAFIZ | OT | NIBP ISSUE | SETTINGS CHANGED AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-04-06 | April 2024 | 12:20 | 2024-04-06 12:05 | 2024-04-06 12:20 | 0:15 | COMPLETED | SEMI-CRITICAL
1469 | 2024-04-06 | Apr 2024 | 13:35 | 13:40 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | VOLUSON E10 | SANTHOSH | OPD | IMAGES NOT PUSHING TO SYATEM | RESTARTED CHECKED CONNECTION,WORKING GOOD | NO | NO | 2024-04-06 | April 2024 | 13:50 | 2024-04-06 13:35 | 2024-04-06 13:50 | 0:15 | COMPLETED | SEMI-CRITICAL
1470 | 2024-04-06 | Apr 2024 | 12:30 | 12:35 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | SANTHOSH | OPD | ECG WAVE PROBLEM | REPLACED NEW ECG CABLE | 10-LEADS ECG CABLE | NO | 2024-04-06 | April 2024 | 12:45 | 2024-04-06 12:30 | 2024-04-06 12:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1471 | 2024-04-06 | Apr 2024 | 7:30 | 7:45 | 0.010417 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | JERLIN | OT | NOT SWITCHING ON | REFIXED THE DISPLAY CABLE AND RECTIFIED THE ISSUE. | NO | NO | 2024-04-06 | April 2024 | 7:50 | 2024-04-06 07:30 | 2024-04-06 07:50 | 0:20 | COMPLETED | CRITICAL
1472 | 2024-04-06 | Apr 2024 | 14:00 | 14:20 | 0.013889 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | JERLIN | OPD | ARTIFACTS IN THE ECG SIGNOLS | REPLACED THE 10 LEAD ECG CABLE, ECG BULB AND RECTIFIED THE ISSUE. | 10-LEADS ECG CABLE | NO | 2024-04-06 | April 2024 | 14:40 | 2024-04-06 14:00 | 2024-04-06 14:40 | 0:40 | COMPLETED | SEMI-CRITICAL
1473 | 2024-04-06 | Apr 2024 | 15:30 | 15:35 | 0.0034722 | CAMERA SYSTEM | KARL STORZ | TC200 | JERLIN | OT | DISPLAY NOT WORKING PROPERLY. | REFIXED THE DISPLAY CABLE AND RECTIFIED THE ISSUE. | NO | NO | 2024-04-06 | April 2024 | 15:50 | 2024-04-06 15:30 | 2024-04-06 15:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1474 | 2024-04-06 | Apr 2024 | 10:05 | 10:10 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | THANISH | OT | THERAPY DIS-ABLED ISSUE | PERFORMED THE OP CHECK | NO | NO | 2024-04-06 | April 2024 | 10:15 | 2024-04-06 10:05 | 2024-04-06 10:15 | 0:10 | COMPLETED | CRITICAL
1475 | 2024-04-07 | Apr 2024 | 20:30 | 20:35 | 0.0034722 | DIALYSIS MACHINE | FRESENIUS | 4008S NG | SOORAJ | DIALYSIS | V99 ERROR | CHECKED AND FOUND CLOG IN BICARBONOTE FILTERS. CLEANED FILTERS AND RAN RINSE ROUINE, ISSUE RECTIFIED. | NO | NO | 2024-04-07 | April 2024 | 20:50 | 2024-04-07 20:30 | 2024-04-07 20:50 | 0:20 | COMPLETED | CRITICAL
1476 | 2024-04-07 | Apr 2024 | 12:05 | 12:10 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | CLEANED THE ABSORBER CANISTER AND RERUNNED THE TEST. ISSUE RECTIFIED AND IS READY TO USE, | NO | NO | 2024-04-07 | April 2024 | 12:25 | 2024-04-07 12:05 | 2024-04-07 12:25 | 0:20 | COMPLETED | CRITICAL
1477 | 2024-04-08 | Apr 2024 | 12:00 | 12:05 | 0.0034722 | PRESSURE INJECTOR,DUAL | LF OPTI VONTAGE | 846700A | HAFIZ | RADIOLOGY | NOT WORKING | CLEANED THE POWER PCB BOARD AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-04-08 | April 2024 | 14:00 | 2024-04-08 12:00 | 2024-04-08 14:00 | 2:00 | COMPLETED | SEMI-CRITICAL
1478 | 2024-04-08 | Apr 2024 | 15:45 | 15:50 | 0.0034722 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | SANTHOSH | OPD | DDNT GET ON | RELEASED EMERGENCY BUTTON | NO | NO | 2024-04-08 | April 2024 | 15:50 | 2024-04-08 15:45 | 2024-04-08 15:50 | 0:05 | COMPLETED | NON-CRITICAL
1479 | 2024-04-08 | Apr 2024 | 11:20 | 11:30 | 0.0069444 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | RAMKIE | OT | NOT WORKING | RECONNECTED THE CABLE. ISSUE RECTIFIED | NO | NO | 2024-04-08 | April 2024 | 11:40 | 2024-04-08 11:20 | 2024-04-08 11:40 | 0:20 | COMPLETED | SEMI-CRITICAL
1480 | 2024-04-09 | Apr 2024 | 14:45 | 14:50 | 0.0034722 | STEAM STERILIZER | GETINGE | GSS67H102E | SOORAJ | CSSD | CPU POWER FAILURE | RESTARTED THE SYSTEM AND CHECKED ERROR LOGS, FOUND JACKET KNOB NOT OPENING. | NO | JACKET KNOB | 2024-04-09 | April 2024 | 16:05 | 2024-04-09 14:45 | 2024-04-09 16:05 | 1:20 | PENDING | SEMI-CRITICAL
1481 | 2024-04-09 | Apr 2024 | 11:45 | 11:50 | 0.0034722 | TREADMILL TEST | WIPRO GE HEALTH CARE | T1000ST | SANTHOSH | OPD | ECG PRINTER ISSUE | ISSUE RECTIFIED | NO | NO | 2024-04-09 | April 2024 | 12:00 | 2024-04-09 11:45 | 2024-04-09 12:00 | 0:15 | COMPLETED | NON-CRITICAL
1482 | 2024-04-09 | Apr 2024 | 16:15 | 16:20 | 0.0034722 | COBLATOR | ARTHO CARE CORPORATION | COBLATOR II | SANTHOSH | OT | LOOSE CONNECTION CONNECTOR | THEY LOST CAP ,AND ISSUE RECTIFIED | NO | NO | 2024-04-09 | April 2024 | 16:30 | 2024-04-09 16:15 | 2024-04-09 16:30 | 0:15 | COMPLETED | SEMI-CRITICAL
1483 | 2024-04-09 | Apr 2024 | 12:00 | 12:10 | 0.0069444 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | RAMKIE | OT | OCCLUSION ERROR | CHECKED AND IT WORKING GOOD, FOUND FEED. | NO | NO | 2024-04-09 | April 2024 | 12:30 | 2024-04-09 12:00 | 2024-04-09 12:30 | 0:30 | COMPLETED | NON-CRITICAL
1484 | 2024-04-09 | Apr 2024 | 23:00 | 23:15 | 0.010417 | PNEUMATIC DRILL | STRYKER | 5407-200-000 | JERLIN | OT | CHUCK HEATING DURING THE PROCEDURE. | CHECKED WITH ANOTHER CHUCK AND CONFIRMED ISSUE WITH THE MOTOR . SERVIVED AND RECTIFIED THE ISSUE. | NO | NO | 2024-04-09 | April 2024 | 23:40 | 2024-04-09 23:00 | 2024-04-09 23:40 | 0:40 | COMPLETED | SEMI-CRITICAL
1485 | 2024-04-10 | Apr 2024 | 15:00 | 15:05 | 0.0034722 | STEAM STERILIZER | GETINGE | GSS67H102E | SOORAJ | CSSD | CPU POWER FAILURE, JACKET TEMP NOT RAISING | COMPANY PERSON CHECKED AND FOUND AIR PRESSURE LEAK ON THE CONTACTORS. TEMPORILY ISSUE RECTIFIED BY COMPANY PERSON. | NO | NO | 2024-04-10 | April 2024 | 17:00 | 2024-04-10 15:00 | 2024-04-10 17:00 | 2:00 | COMPLETED | SEMI-CRITICAL
1486 | 2024-04-10 | Apr 2024 | 8:10 | 8:15 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | ICU | OPRATIONOL CHECK FAILED | RERUNNED THE OPERATIONAL CHECK IN SERVICE MODE. NOW READY TO USE. | NO | NO | 2024-04-10 | April 2024 | 8:20 | 2024-04-10 08:10 | 2024-04-10 08:20 | 0:10 | COMPLETED | CRITICAL
1487 | 2024-04-10 | Apr 2024 | 8:45 | 8:50 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | HAFIZ | ICU | NO POWER SUPPLY | ISSUE WITH POWER SOCKET. CHANGED THE SOCKET AND READY TO USE. | NO | NO | 2024-04-10 | April 2024 | 9:00 | 2024-04-10 08:45 | 2024-04-10 09:00 | 0:15 | COMPLETED | SEMI-CRITICAL
1488 | 2024-04-10 | Apr 2024 | 11:00 | 11:02 | 0.0013889 | TREADMILL TEST | WIPRO GE HEALTH CARE | T1000-ST | HAFIZ | OPD | PRINTER NOT WORKING | CARTRIDGE RELOADED AND OBSERVED. ISSUE RECTIFIED. | CARTRIDGE | NO | 2024-04-10 | April 2024 | 11:30 | 2024-04-10 11:00 | 2024-04-10 11:30 | 0:30 | COMPLETED | NON-CRITICAL
1489 | 2024-04-10 | Apr 2024 | 10:30 | 10:35 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | HAFIZ | ICU | PRE-USE CHECK FAILED | ISSUE WITH O2 PENDANT. FIXED IN ANOTHER PENDANT. AND RERUN THE PRE-USE CHECK. NOW MACHINE READY TO USE. | NO | NO | 2024-04-10 | April 2024 | 10:40 | 2024-04-10 10:30 | 2024-04-10 10:40 | 0:10 | COMPLETED | CRITICAL
1490 | 2024-04-10 | Apr 2024 | 16:30 | 16:35 | 0.0034722 | CAMERA SYSTEM | KARL STORZ | TC200 | SANTHOSH | OT | RECORDER NOT WORKING | FORMATED NTFS ,ISSUE RECTIFIED | NO | NO | 2024-04-10 | April 2024 | 16:40 | 2024-04-10 16:30 | 2024-04-10 16:40 | 0:10 | COMPLETED | SEMI-CRITICAL
1491 | 2024-04-10 | Apr 2024 | 14:00 | 14:10 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | OT | NOT WORKING | CHECKED AND POWER CABLE RECONNECTED AND RECTIFIED. | NO | NO | 2024-04-10 | April 2024 | 14:30 | 2024-04-10 14:00 | 2024-04-10 14:30 | 0:30 | COMPLETED | NON-CRITICAL
1492 | 2024-04-10 | Apr 2024 | 2:00 | 2:15 | 0.010417 | PATIENT MONITOR | PHILIPS | MX 450 | JERLIN | ICU | MONITOR NOT SWITCHING ON | REFIXED THE POWER CORD PROPERLY AND RECTIFIED THE ISSUE | NO | NO | 2024-04-10 | April 2024 | 2:35 | 2024-04-10 02:00 | 2024-04-10 02:35 | 0:35 | COMPLETED | SEMI-CRITICAL
1493 | 2024-04-11 | Apr 2024 | 10:10 | 10:15 | 0.0034722 | MRI MACHINE | PHILIPS | INGENIA ELITON S | THANISH | RADIOLOGY | GRADIENT COIL ERROR | RE-STORED THE BACKUP DATA.WORKING FINE. | NO | NO | 2024-04-11 | April 2024 | 11:05 | 2024-04-11 10:10 | 2024-04-11 11:05 | 0:55 | COMPLETED | CRITICAL
1494 | 2024-04-11 | Apr 2024 | 14:30 | 14:40 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | ICU | NOT WORKING | RESET DONE ,NOW WORKING | NO | NO | 2024-04-11 | April 2024 | 14:50 | 2024-04-11 14:30 | 2024-04-11 14:50 | 0:20 | COMPLETED | NON-CRITICAL
1495 | 2024-04-11 | Apr 2024 | 2:30 | 2:35 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | JERLIN | ICU | WARMER NOT WORKING | FIXED THE TORN WARMER HOSE AND RECTIFIED THE ISSUE. | NO | NO | 2024-04-11 | April 2024 | 2:55 | 2024-04-11 02:30 | 2024-04-11 02:55 | 0:25 | COMPLETED | SEMI-CRITICAL
1496 | 2024-04-11 | Apr 2024 | 10:10 | 10:15 | 0.0034722 | MRI MACHINE | PHILIPS | INGENIA ELITON S | THANISH | RADIOLOGY | GRADIENT COIL ERROR | RE-STORED THE BACKUP DATA.WORKING FINE. | NO | NO | 2024-04-11 | April 2024 | 11:20 | 2024-04-11 10:10 | 2024-04-11 11:20 | 1:10 | COMPLETED | CRITICAL
1497 | 2024-04-12 | Apr 2024 | 10:30 | 10:35 | 0.0034722 | MRI MACHINE | PHILIPS | INGENIA ELITON S | THANISH | RADIOLOGY | GRADIENT COIL ERROR | SKY LIGHT ISSUE RECTIFIED.WORKING FINE. | NO | NO | 2024-04-12 | April 2024 | 13:15 | 2024-04-12 10:30 | 2024-04-12 13:15 | 2:45 | COMPLETED | CRITICAL
1498 | 2024-04-12 | Apr 2024 | 7:40 | 7:45 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SANTHOSH | OT | CALIBRATION DUE | O2 CELL CALIBRATION DONE | NO | NO | 2024-04-12 | April 2024 | 7:50 | 2024-04-12 07:40 | 2024-04-12 07:50 | 0:10 | COMPLETED | CRITICAL
1499 | 2024-04-12 | Apr 2024 | 10:30 | 10:35 | 0.0034722 | MRI MACHINE | PHILIPS | INGENIA ELITON S | THANISH | RADIOLOGY | GRADIENT COIL ERROR | SKY LIGHT ISSUE RECTIFIED.WORKING FINE. | NO | NO | 2024-04-12 | April 2024 | 13:10 | 2024-04-12 10:30 | 2024-04-12 13:10 | 2:40 | COMPLETED | CRITICAL
1500 | 2024-04-13 | Apr 2024 | 10:00 | 10:05 | 0.0034722 | TREADMILL TEST | WIPRO GE HEALTH CARE | T1000ST | SANTHOSH | OPD | PRINTER ISSUE | ISSUE RECTIFIED | NO | NO | 2024-04-13 | April 2024 | 10:10 | 2024-04-13 10:00 | 2024-04-13 10:10 | 0:10 | COMPLETED | NON-CRITICAL
1501 | 2024-04-13 | Apr 2024 | 15:10 | 15:20 | 0.0069444 | BMI MACHINE | WELCH ALLYN | TANITA | RAMKIE | EMERGENCY | NOT SWITCHING ON | CHECKED AND NEED KEYPAD REPLACEMENT | NO | KEYPAD | 2024-04-13 | April 2024 | 15:40 | 2024-04-13 15:10 | 2024-04-13 15:40 | 0:30 | PENDING | NON-CRITICAL
1502 | 2024-04-13 | Apr 2024 | 3:00 | 3:05 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | ICU | REPEATED OCCLUSION ALARM | CHANGED THE OCCCLUSION SETTINGS AND RECTIFIED THE ISSUE. | NO | NO | 2024-04-13 | April 2024 | 3:20 | 2024-04-13 03:00 | 2024-04-13 03:20 | 0:20 | COMPLETED | NON-CRITICAL
1503 | 2024-04-13 | Apr 2024 | 23:50 | 23:55 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | ICU | FLOW RATE ISSUE | REPLACED THE DAMAGED DOOR AND RECTIFIED THE ISSUE | NO | NO | 2024-04-14 | April 2024 | 0:10 | 2024-04-13 23:50 | 2024-04-14 00:10 | 0:20 | COMPLETED | NON-CRITICAL
1504 | 2024-04-14 | Apr 2024 | 15:00 | 15:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | FOUND MOISTURE ACCUMULATED ON TOP OF SODA LIME CHAMBER. REMOVED THE MOISTURE AND CLEANED THE CHAMBER DRY. CHECKOUT DONE AND ISSUE RECTIFIED. | NO | NO | 2024-04-14 | April 2024 | 15:20 | 2024-04-14 15:00 | 2024-04-14 15:20 | 0:20 | COMPLETED | CRITICAL
1505 | 2024-04-14 | Apr 2024 | 22:30 | 22:32 | 0.0013889 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | JERLIN | OT | DISPLAY BROKEN | DISPLAY HAS BEEN PHYSICALLY DAMAGED AND NEEDS TO BE REPLACED | NO | DISPLAY | 2024-04-14 | April 2024 | 22:46 | 2024-04-14 22:30 | 2024-04-14 22:46 | 0:16 | PENDING | SEMI-CRITICAL
1506 | 2024-04-15 | Apr 2024 | 14:30 | 14:34 | 0.0027778 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | CVP TEST DUE | RUNNED CVP TEST. NOW MACHINE READY TO USE. | NO | NO | 2024-04-15 | April 2024 | 14:45 | 2024-04-15 14:30 | 2024-04-15 14:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1507 | 2024-04-15 | Apr 2024 | 10:00 | 10:05 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | CATRIDGE EJECTED | REPLACED ANOTHER CATRIDGE | CARTRIDGE | NO | 2024-04-15 | April 2024 | 11:00 | 2024-04-15 10:00 | 2024-04-15 11:00 | 1:00 | COMPLETED | SEMI-CRITICAL
1508 | 2024-04-15 | Apr 2024 | 20:30 | 20:40 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | PROCESS CONTROL SOLUTION NOT DETECTED. | CLOT PATTERN OBSERVED. CHANGED THE CARTRIDGE AND RECTIFIED THE ISSUE | NO | NO | 2024-04-15 | April 2024 | 20:57 | 2024-04-15 20:30 | 2024-04-15 20:57 | 0:27 | COMPLETED | SEMI-CRITICAL
1509 | 2024-04-16 | Apr 2024 | 10:30 | 10:34 | 0.0027778 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | NEED TO REPLACE CARTRIDGE | CARTRIDGE REPLACED AND NOW MACHINE READY TO USE. | CARTRIDGE | NO | 2024-04-16 | April 2024 | 11:30 | 2024-04-16 10:30 | 2024-04-16 11:30 | 1:00 | COMPLETED | SEMI-CRITICAL
1510 | 2024-04-16 | Apr 2024 | 11:30 | 11:31 | 0.00069444 | PATIENT MONITOR | PHILIPS | MX 450 | HAFIZ | ICU | SPO2 PROBE NOT WORKING | CLEANED AND REFIXED THE PROBE. ISSUE RECTIFIED. | NO | NO | 2024-04-16 | April 2024 | 11:45 | 2024-04-16 11:30 | 2024-04-16 11:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1511 | 2024-04-16 | Apr 2024 | 20:00 | 20:10 | 0.0069444 | BLOOD GAS ANALYSER | GEM 3500 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | CATRIDGE NEED TO REPLACE | NEW CATRIDGE REPLACED | CARTRIDGE | NO | 2024-04-16 | April 2024 | 20:10 | 2024-04-16 20:00 | 2024-04-16 20:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1512 | 2024-04-16 | Apr 2024 | 15:20 | 15:30 | 0.0069444 | TREADMILL TEST | AFTON | AFTON | RAMKIE | PHYSIOTHERAPY | CONNECTION FAILED | CHECKED AND SAFETYKEY ADJUSTED | NO | NO | 2024-04-16 | April 2024 | 15:40 | 2024-04-16 15:20 | 2024-04-16 15:40 | 0:20 | COMPLETED | NON-CRITICAL
1513 | 2024-04-17 | Apr 2024 | 12:50 | 12:55 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SOORAJ | DIALYSIS | PRINTER NOT WORKING. | FOUND ROLLER TRAY BROKEN. REPLACED FROM ANOTHER DEFECTIVE DEFIB AND ISSUE RECTIFIED. | PRINTER ROLLER HEAD | NO | 2024-04-17 | April 2024 | 13:10 | 2024-04-17 12:50 | 2024-04-17 13:10 | 0:20 | COMPLETED | CRITICAL
1514 | 2024-04-17 | Apr 2024 | 12:05 | 12:10 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | ICU | PRINTER NOT WORKING | PROPERLY FIXED THE PAPERROLL AND ISSUE RECTIFIED. | NO | NO | 2024-04-17 | April 2024 | 12:25 | 2024-04-17 12:05 | 2024-04-17 12:25 | 0:20 | COMPLETED | CRITICAL
1515 | 2024-04-17 | Apr 2024 | 17:00 | 17:03 | 0.0020833 | TRANSPORT VENTILATOR | RESMED | ASTRAL 100 | HAFIZ | ICU | PRE-USE TEST FAILED | RERUNNED THE TEST WITH FILTER. NOW PASSED AND READY TO USE. | NO | NO | 2024-04-17 | April 2024 | 17:35 | 2024-04-17 17:00 | 2024-04-17 17:35 | 0:35 | COMPLETED | SEMI-CRITICAL
1516 | 2024-04-17 | Apr 2024 | 0:15 | 0:20 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID E95 | SANTHOSH | OPD | SUDDENLY TURN OF | ISSUE RECTIFIED | NO | NO | 2024-04-17 | April 2024 | 0:25 | 2024-04-17 00:15 | 2024-04-17 00:25 | 0:10 | COMPLETED | SEMI-CRITICAL
1517 | 2024-04-17 | Apr 2024 | 0:25 | 0:30 | 0.0034722 | URINE ANALYSER | COBAS U 411 | ROCHE | SANTHOSH | EMERGENCY | ERROR 350107 | MAINTANONCE RUN ISUE RECTIFIED | NO | NO | 2024-04-17 | April 2024 | 0:30 | 2024-04-17 00:25 | 2024-04-17 00:30 | 0:05 | COMPLETED | SEMI-CRITICAL
1518 | 2024-04-17 | Apr 2024 | 14:10 | 14:30 | 0.013889 | PLASMA BATH | TERUMO PENPOL | PB-100 | RAMKIE | BLOODBANK | NOT HEATING | CHECKED AND CLEARED THE BLOCK, NOT IT WORKING | NO | NO | 2024-04-17 | April 2024 | 14:40 | 2024-04-17 14:10 | 2024-04-17 14:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1519 | 2024-04-18 | Apr 2024 | 13:00 | 13:05 | 0.0034722 | CATHLAB - MONOPLANE | PHILIPS | AZURION 7M12 | HAFIZ | CATHLAB | PACS ISSUE | SETTINGS CHANGED AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-04-18 | April 2024 | 13:20 | 2024-04-18 13:00 | 2024-04-18 13:20 | 0:20 | COMPLETED | CRITICAL
1520 | 2024-04-18 | Apr 2024 | 15:00 | 15:05 | 0.0034722 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | HAFIZ | RADIOLOGY | DETECTOR BATTERY ISSUE | GIVEN ONE STANDBY DETECTOR BATTERY. | DETECTOR BATTERY | NO | 2024-04-18 | April 2024 | 15:15 | 2024-04-18 15:00 | 2024-04-18 15:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1521 | 2024-04-18 | Apr 2024 | 0:05 | 0:10 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SANTHOSH | OT | CIRCUIT LEAK TEST FAILED | ISSUE RECTIFIED | NO | NO | 2024-04-18 | April 2024 | 0:25 | 2024-04-18 00:05 | 2024-04-18 00:25 | 0:20 | COMPLETED | CRITICAL
1522 | 2024-04-19 | Apr 2024 | 0:20 | 0:25 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | SANTHOSH | ICU | CENTRAL MONITERING ISSUE | RESET DONE ,NOW WORKING | NO | NO | 2024-04-19 | April 2024 | 0:30 | 2024-04-19 00:20 | 2024-04-19 00:30 | 0:10 | COMPLETED | SEMI-CRITICAL
1523 | 2024-04-20 | Apr 2024 | 10:35 | 10:40 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | HAFIZ | WARD | ARTFACT ISSUE | SETTINGS CHANGED AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-04-20 | April 2024 | 10:50 | 2024-04-20 10:35 | 2024-04-20 10:50 | 0:15 | COMPLETED | SEMI-CRITICAL
1524 | 2024-04-20 | Apr 2024 | 11:35 | 11:40 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | HAFIZ | OPD | HANGING ISSUE | SPOOLER CLEARED AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-04-20 | April 2024 | 11:50 | 2024-04-20 11:35 | 2024-04-20 11:50 | 0:15 | COMPLETED | SEMI-CRITICAL
1525 | 2024-04-20 | Apr 2024 | 11:50 | 11:53 | 0.0020833 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | HAFIZ | OPD | ARTFACT ISSUE | SETTINGS CHANGED AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-04-20 | April 2024 | 12:00 | 2024-04-20 11:50 | 2024-04-20 12:00 | 0:10 | COMPLETED | SEMI-CRITICAL
1526 | 2024-04-20 | Apr 2024 | 3:00 | 3:10 | 0.0069444 | VENTILATOR, TRANSPORT | RESMED | ASTRAL 100 | SANTHOSH | WARD | CIRCUIT LEAK ISSUE | ISSUE RECTIFIED | NO | NO | 2024-04-20 | April 2024 | 3:10 | 2024-04-20 03:00 | 2024-04-20 03:10 | 0:10 | COMPLETED | CRITICAL
1527 | 2024-04-20 | Apr 2024 | 3:20 | 3:25 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | SANTHOSH | OPD | STORAGE FULL | ISSUE RECTIFIED | NO | NO | 2024-04-20 | April 2024 | 3:25 | 2024-04-20 03:20 | 2024-04-20 03:25 | 0:05 | COMPLETED | SEMI-CRITICAL
1528 | 2024-04-20 | Apr 2024 | 16:10 | 16:20 | 0.0069444 | EXAMINATION COUCH | MIDMARK JANOK | E122012 | RAMKIE | OPD | NOT HEATING | PROBLEM WITH ACTUATOR | NO | NO | 2024-04-20 | April 2024 | 16:40 | 2024-04-20 16:10 | 2024-04-20 16:40 | 0:30 | COMPLETED | NON-CRITICAL
1529 | 2024-04-21 | Apr 2024 | 6:35 | 6:40 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SANTHOSH | WARD | POWER SOCKET ISSUE | ISSUE RECTIFIED | NO | NO | 2024-04-21 | April 2024 | 6:40 | 2024-04-21 06:35 | 2024-04-21 06:40 | 0:05 | COMPLETED | NON-CRITICAL
1530 | 2024-04-21 | Apr 2024 | 6:40 | 6:45 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SANTHOSH | WARD | POWER SOCKET ISSUE | ISSUE RECTIFIED | NO | NO | 2024-04-21 | April 2024 | 6:50 | 2024-04-21 06:40 | 2024-04-21 06:50 | 0:10 | COMPLETED | NON-CRITICAL
1531 | 2024-04-21 | Apr 2024 | 9:30 | 9:35 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | OPD | OPERATIONAL CHECK FAILED | RERAN OP CHECK AND RECTIFIED THE ISSUE. | NO | NO | 2024-04-21 | April 2024 | 9:50 | 2024-04-21 09:30 | 2024-04-21 09:50 | 0:20 | COMPLETED | CRITICAL
1532 | 2024-04-22 | Apr 2024 | 16:00 | 16:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | SOORAJ | OT | BELLOW ISSUE | FOUND ISSUE WITH PATIENT CIRCUIT. REPLACED NEW CIRCUIT AND ISSUE RECTIFIED. | NO | NO | 2024-04-22 | April 2024 | 16:20 | 2024-04-22 16:00 | 2024-04-22 16:20 | 0:20 | COMPLETED | CRITICAL
1533 | 2024-04-22 | Apr 2024 | 20:00 | 20:02 | 0.0013889 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | NEED TO REPLACE CARTRIDGE | CARTRIDGE REPLACED AND NOW MACHINE READY TO USE. | CARTRIDGE | NO | 2024-04-22 | April 2024 | 21:00 | 2024-04-22 20:00 | 2024-04-22 21:00 | 1:00 | COMPLETED | SEMI-CRITICAL
1534 | 2024-04-22 | Apr 2024 | 0:10 | 0:20 | 0.0069444 | STEAM STERILIZER | GETINGE | GSS67H102E | SANTHOSH | CSSD | AIRVALVE ISSUE | AIRVALVE REPLACED FROM MACHINE 2 | NO | NO | 2024-04-22 | April 2024 | 0:25 | 2024-04-22 00:10 | 2024-04-22 00:25 | 0:15 | COMPLETED | SEMI-CRITICAL
1535 | 2024-04-22 | Apr 2024 | 16:30 | 16:40 | 0.0069444 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | JERLIN | RADIOLOGY | MACHINE NOT SWITCHING ON. | CHANGED THE EXTENSION BOX AND RECTIFIED THE ISSUE | NO | NO | 2024-04-22 | April 2024 | 16:55 | 2024-04-22 16:30 | 2024-04-22 16:55 | 0:25 | COMPLETED | SEMI-CRITICAL
1536 | 2024-04-23 | Apr 2024 | 6:00 | 6:05 | 0.0034722 | PLASMA STERILIZER | JOHNSON&JOHNSON | STERRAD 100NX | HAFIZ | CSSD | INSTRUMENT STUCK INSIDE | OPENED THE FRONT DOOR AND RETRIEVED THE INSTRUMENT. ISSUE RECTIFIED. | NO | NO | 2024-04-23 | April 2024 | 6:25 | 2024-04-23 06:00 | 2024-04-23 06:25 | 0:25 | COMPLETED | SEMI-CRITICAL
1537 | 2024-04-23 | Apr 2024 | 20:30 | 20:32 | 0.0013889 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | HAFIZ | CATHLAB | CIRCUIT LEAK TEST FAILED | FOUND ABSORBER CANISTER NOT LOCKED. RELOCKED IT. AND NOW TEST PASSED AND IS NOW READY TO USE. | NO | NO | 2024-04-23 | April 2024 | 20:40 | 2024-04-23 20:30 | 2024-04-23 20:40 | 0:10 | COMPLETED | CRITICAL
1538 | 2024-04-23 | Apr 2024 | 15:30 | 15:35 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | SANTHOSH | ICU | NOT WORKING | ISSUE RECTIFIED | NO | NO | 2024-04-23 | April 2024 | 15:45 | 2024-04-23 15:30 | 2024-04-23 15:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1539 | 2024-04-23 | Apr 2024 | 17:30 | 17:40 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | RAMKIE | WARD | NOT WORKING | CHECKED AND RESET DONE AND NOW IRT WORKING | NO | NO | 2024-04-23 | April 2024 | 17:50 | 2024-04-23 17:30 | 2024-04-23 17:50 | 0:20 | COMPLETED | NON-CRITICAL
1540 | 2024-04-23 | Apr 2024 | 14:10 | 14:30 | 0.013889 | EEG MACHINE | NOTUS NEUROLOGY | NICOLET EEG V32 | JERLIN | OPD | SOFTWARE ABOUT TO EXPIRE | UPDATED THE SOFTWARE AND RECTIFIED THE ISSUE | NO | NO | 2024-04-23 | April 2024 | 14:45 | 2024-04-23 14:10 | 2024-04-23 14:45 | 0:35 | COMPLETED | SEMI-CRITICAL
1541 | 2024-04-23 | Apr 2024 | 15:00 | 15:15 | 0.010417 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID E95 | JERLIN | OPD | MACHINE HANGS DURING THE PROCEDURE | RELOADED THE YSTEM SOFTWARE AND RECTIFIED THE ISSUE | NO | NO | 2024-04-23 | April 2024 | 15:35 | 2024-04-23 15:00 | 2024-04-23 15:35 | 0:35 | COMPLETED | SEMI-CRITICAL
1542 | 2024-04-24 | Apr 2024 | 14:00 | 14:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | SOORAJ | OT | BATTERY MALFUNCTION ALARM | FOUND BATTERY TO BE DEFECTIVE. AS OF NOW CONNECTED ON MAINS AND WORKING. | NO | NO | 2024-04-24 | April 2024 | 14:30 | 2024-04-24 14:00 | 2024-04-24 14:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1543 | 2024-04-24 | Apr 2024 | 1:45 | 1:50 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | HAFIZ | OT | GAS MODULE ISSUE | REPLACED THE GAS MODULE (S.NO: DE43803809): OLD MODULE S.NO: DE43805579. | GAS MODULE | NO | 2024-04-24 | April 2024 | 2:00 | 2024-04-24 01:45 | 2024-04-24 02:00 | 0:15 | COMPLETED | SEMI-CRITICAL
1544 | 2024-04-24 | Apr 2024 | 2:00 | 2:00 | 0 | PATIENT MONITOR | PHILIPS | MX 550 | HAFIZ | OT | GAS MODULE ISSUE | REPLACED THE GAS MODULE (S.NO: DE43812976); OLD MODULE S.NO: DE43805577. | GAS MODULE | NO | 2024-04-24 | April 2024 | 2:10 | 2024-04-24 02:00 | 2024-04-24 02:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1545 | 2024-04-24 | Apr 2024 | 22:00 | 22:10 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | FOUND ABSORBER CANISTER MOISTURE. CLEANED IT. AND NOW TEST PASSED AND IS NOW READY TO USE. | NO | NO | 2024-04-24 | April 2024 | 22:20 | 2024-04-24 22:00 | 2024-04-24 22:20 | 0:20 | COMPLETED | CRITICAL
1546 | 2024-04-24 | Apr 2024 | 18:30 | 18:40 | 0.0069444 | NITRIC OXIDE DELIVERY SYSTEM | WIPRO GE HEALTH CARE | SLE | RAMKIE | ICU | NOT DISPLAY DATA | CALIBRATION DONE AND DISPLAY CONTRAST ADJUTSED | NO | NO | 2024-04-24 | April 2024 | 18:50 | 2024-04-24 18:30 | 2024-04-24 18:50 | 0:20 | COMPLETED | CRITICAL
1547 | 2024-04-25 | Apr 2024 | 7:45 | 7:50 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | RESTARTED THE MACHINE AND CHECKOUT DONE AGAIN. ISSUE RECTIFIED. | NO | NO | 2024-04-25 | April 2024 | 7:55 | 2024-04-25 07:45 | 2024-04-25 07:55 | 0:10 | COMPLETED | CRITICAL
1548 | 2024-04-25 | Apr 2024 | 9:00 | 9:02 | 0.0013889 | STEAM STERILIZER | GETINGE | GSS67H102E | SOORAJ | CSSD | VACCUM ALARM | FOUND ISSUE WITH YV 300 CONTACTOR. REPALCED FROM ANOTHER MACHINE AND ISSUE RECTIFIED. | YV 300 MODULE | NO | 2024-04-25 | April 2024 | 12:15 | 2024-04-25 09:00 | 2024-04-25 12:15 | 3:15 | COMPLETED | SEMI-CRITICAL
1549 | 2024-04-25 | Apr 2024 | 1:00 | 1:06 | 0.0041667 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | ICU | HOLDER CLAMP ISSUE | REPLACED THE HOLDER CLAMP. ISSUE RECTIFIED. | HOLDER CLAMP | NO | 2024-04-25 | April 2024 | 1:20 | 2024-04-25 01:00 | 2024-04-25 01:20 | 0:20 | COMPLETED | NON-CRITICAL
1550 | 2024-04-25 | Apr 2024 | 4:20 | 4:24 | 0.0027778 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | THANISH | WARD | NIBP MEASUREMENT FAILED | AIR LEAK ARRESTED AND ISSUE RECTIFIED. | NO | NO | 2024-04-25 | April 2024 | 4:35 | 2024-04-25 04:20 | 2024-04-25 04:35 | 0:15 | COMPLETED | SEMI-CRITICAL
1551 | 2024-04-25 | Apr 2024 | 6:00 | 6:05 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | HAFIZ | ICU | NOT WORKING | FOUND THERMISTOR CABLE DETACHED. REFIXED IT AND ISSUE RECTIFIED. | NO | NO | 2024-04-25 | April 2024 | 6:15 | 2024-04-25 06:00 | 2024-04-25 06:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1552 | 2024-04-25 | Apr 2024 | 15:40 | 15:45 | 0.0034722 | NITRIC OXIDE DELIVERY SYSTEM | WIPRO GE HEALTH CARE | SLE | RAMKIE | ICU | NOT DISPLAY DATA | CALIBRATION DONE | NO | NO | 2024-04-25 | April 2024 | 17:40 | 2024-04-25 15:40 | 2024-04-25 17:40 | 2:00 | COMPLETED | CRITICAL
1553 | 2024-04-25 | Apr 2024 | 10:30 | 10:40 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 450 | JERLIN | EMERGENCY | BP CUFF LEAK ISSUE | REPLACED THE BP CUFFS AND RECTIFIED THE ISSUE | BP CUFF | NO | 2024-04-25 | April 2024 | 10:55 | 2024-04-25 10:30 | 2024-04-25 10:55 | 0:25 | COMPLETED | SEMI-CRITICAL
1554 | 2024-04-26 | Apr 2024 | 17:30 | 17:40 | 0.0069444 | BIOCHEMISTRY ANALYSER | COBAS C 501 | ROCHE | RAMKIE | LABORATORY | LEAK IN MACHINE | CHECKED AND TUBE RECONNECTED | NO | NO | 2024-04-26 | April 2024 | 18:10 | 2024-04-26 17:30 | 2024-04-26 18:10 | 0:40 | COMPLETED | SEMI-CRITICAL
1555 | 2024-04-27 | Apr 2024 | 11:20 | 11:30 | 0.0069444 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | JERLIN | CATHLAB | MACHINE NOT WORKING PROPERLY | REFIXED THE TEMP PROBE AND RECTIFIED THE ISSUE | NO | NO | 2024-04-27 | April 2024 | 11:45 | 2024-04-27 11:20 | 2024-04-27 11:45 | 0:25 | COMPLETED | SEMI-CRITICAL
1556 | 2024-04-27 | Apr 2024 | 14:55 | 15:00 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | JERLIN | RADIOLOGY | MACHINE RESTARTING AUTOMATICALLY | RELOADED THE SYSTEM SOFTWARE AND RECTIFIED THE ISSUE | NO | NO | 2024-04-27 | April 2024 | 16:05 | 2024-04-27 14:55 | 2024-04-27 16:05 | 1:10 | COMPLETED | SEMI-CRITICAL
1557 | 2024-04-28 | Apr 2024 | 7:00 | 7:02 | 0.0013889 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | ICU | DOOR BROKEN | REPLACED THE DOOR, OBSERVED AND ISSUE RECTIFIED. | NO | NO | 2024-04-28 | April 2024 | 7:10 | 2024-04-28 07:00 | 2024-04-28 07:10 | 0:10 | COMPLETED | NON-CRITICAL
1558 | 2024-04-28 | Apr 2024 | 21:00 | 21:03 | 0.0020833 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | OT | PRE-USE TEST FAILED | RERUNNED THE PRE-USE CHECK AND ISSUE RECTIFIED. | NO | NO | 2024-04-28 | April 2024 | 21:20 | 2024-04-28 21:00 | 2024-04-28 21:20 | 0:20 | COMPLETED | CRITICAL
1559 | 2024-04-28 | Apr 2024 | 22:30 | 22:34 | 0.0027778 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | HAFIZ | WARD | NOT WORKING | FOUND THERMISTOR CABLE DETACHED. REFIXED IT AND ISSUE RECTIFIED. | NO | NO | 2024-04-28 | April 2024 | 22:45 | 2024-04-28 22:30 | 2024-04-28 22:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1560 | 2024-04-28 | Apr 2024 | 7:30 | 7:35 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | JERLIN | RADIOLOGY | COLOR WAVE OPTION NOT ENOBLED | CHANGED THE SETTINGS AND RECTIFIED THE ISSUE | NO | NO | 2024-04-28 | April 2024 | 7:55 | 2024-04-28 07:30 | 2024-04-28 07:55 | 0:25 | COMPLETED | SEMI-CRITICAL
1561 | 2024-04-29 | Apr 2024 | 8:30 | 8:35 | 0.0034722 | STEAM STERILIZER | GETINGE | GSS67H102E | SANTHOSH | CSSD | VACCUM ISSUE | RESTARTED,ISSUE RECTIFIED | NO | NO | 2024-04-29 | April 2024 | 8:45 | 2024-04-29 08:30 | 2024-04-29 08:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1562 | 2024-04-29 | Apr 2024 | 18:20 | 18:30 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | RAMKIE | ICU | SCAN GUN NOT WORKING | CVP DONE, RECONNECTED AND RECTIFIED | NO | NO | 2024-04-29 | April 2024 | 18:50 | 2024-04-29 18:20 | 2024-04-29 18:50 | 0:30 | COMPLETED | SEMI-CRITICAL
1563 | 2024-04-30 | Apr 2024 | 14:00 | 14:05 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | ICU | NOT WORKING | ISSUE WITH ACTUATOR. NEED REPLACEMENT. | NO | ACTUATOR | 2024-04-30 | April 2024 | 14:15 | 2024-04-30 14:00 | 2024-04-30 14:15 | 0:15 | PENDING | NON-CRITICAL
1564 | 2024-04-30 | Apr 2024 | 18:30 | 18:35 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | HAFIZ | ICU | NEED TO REPLACE CARTRIDGE | INSTALLED NEW CARTRDIGE (S.NO: 301880575). USED CARTRIDGE (S.NO: 301872551). WARMING UP DONE. READY TO USE. | CARTRIDGE | NO | 2024-04-30 | April 2024 | 19:10 | 2024-04-30 18:30 | 2024-04-30 19:10 | 0:40 | COMPLETED | SEMI-CRITICAL
1565 | 2024-04-30 | Apr 2024 | 19:00 | 19:15 | 0.010417 | PATIENT MONITOR | PHILIPS | MX 450 | HAFIZ | ICU | NIBP MEASUREMENT FAILED | AIR LEAK ARRESTED AND ISSUE RECTIFIED. | NO | NO | 2024-04-30 | April 2024 | 19:35 | 2024-04-30 19:00 | 2024-04-30 19:35 | 0:35 | COMPLETED | SEMI-CRITICAL
1566 | 2024-05-01 | May 2024 | 19:00 | 19:05 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | ICU | BREAK LEVER TOO TIGHT | LUBRICATED THE BRAKE AND ISSUE RECTIFIED | NO | NO | 2024-05-01 | May 2024 | 19:30 | 2024-05-01 19:00 | 2024-05-01 19:30 | 0:30 | COMPLETED | NON-CRITICAL
1567 | 2024-05-01 | May 2024 | 19:45 | 19:50 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | SOORAJ | ICU | MACHINE NOT SWITCHING ON | FOUND MAIN POWER BATTERY SWITCH IN OFF POSITION. | NO | NO | 2024-05-01 | May 2024 | 20:00 | 2024-05-01 19:45 | 2024-05-01 20:00 | 0:15 | COMPLETED | SEMI-CRITICAL
1568 | 2024-05-01 | May 2024 | 20:00 | 20:00 | 0 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | SOORAJ | ICU | MACHINE NOT SWITCHING ON | USG SYSTEM NOT PLACED PROPERLY ON TROLLEY. FIXED PROPERLY AND CHECKED, WORKING FINE. | NO | NO | 2024-05-01 | May 2024 | 20:20 | 2024-05-01 20:00 | 2024-05-01 20:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1569 | 2024-05-01 | May 2024 | 21:00 | 21:05 | 0.0034722 | VENTILATOR, PEDIATRIC | MINDRAY MEDICAL INDIA | SV 300 | SOORAJ | EMERGENCY | O2 SENSOR TEST FAILED | CLEANED O2 SENSOR AND REFIXED, CALIBRATION PERFORMED BUT FAILED. O2 SENSOR GONE DEFECTIVE. | NO | O2 SENSOR | 2024-05-01 | May 2024 | 21:45 | 2024-05-01 21:00 | 2024-05-01 21:45 | 0:45 | PENDING | CRITICAL
1570 | 2024-05-01 | May 2024 | 22:00 | 22:05 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SOORAJ | ICU | PARAMETERS DISABLED | REPLACED CARTRIDGE | CARTRIDGE | NO | 2024-05-01 | May 2024 | 22:45 | 2024-05-01 22:00 | 2024-05-01 22:45 | 0:45 | COMPLETED | SEMI-CRITICAL
1571 | 2024-05-01 | May 2024 | 8:30 | 8:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | HAFIZ | EMERGENCY | NOT WORKING | FOUND POWERCORD DETACHED FROM CONTROLBOX, REFIXED IT. NOW WORKING GOOD. | NO | NO | 2024-05-01 | May 2024 | 8:45 | 2024-05-01 08:30 | 2024-05-01 08:45 | 0:15 | COMPLETED | NON-CRITICAL
1572 | 2024-05-01 | May 2024 | 9:00 | 9:02 | 0.0013889 | RADIANT HEAT WARMER CUM INCUBATOR | WIPRO GE HEALTH CARE | GIRAFFE | HAFIZ | ICU | AUTOMATICALLY SWITCHING OFF | ISSUE WITH THE POWERSOCKET. CHANGED IT AND ISSUE RECTIFIED. | NO | NO | 2024-05-01 | May 2024 | 9:15 | 2024-05-01 09:00 | 2024-05-01 09:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1573 | 2024-05-01 | May 2024 | 11:00 | 11:05 | 0.0034722 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | HAFIZ | ICU | HOLDER CLAMP ISSUE | REPLACED THE HOLDER CLAMP AND ISSUE RECTIFIED. | HOLDER CLAMP | NO | 2024-05-01 | May 2024 | 11:30 | 2024-05-01 11:00 | 2024-05-01 11:30 | 0:30 | COMPLETED | NON-CRITICAL
1574 | 2024-05-01 | May 2024 | 12:30 | 12:35 | 0.0034722 | MONITOR HIGH DEFINITION | KARL STORZ | TM220 | HAFIZ | OPD | RECORDING ISSUE | FOUND CABLE LOOSENED. TIGHTENED IT AND ISSUE RECTIFIED. | NO | NO | 2024-05-01 | May 2024 | 12:50 | 2024-05-01 12:30 | 2024-05-01 12:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1575 | 2024-05-01 | May 2024 | 13:00 | 13:05 | 0.0034722 | IABP MACHINE | MAQUET | CS300 | HAFIZ | ICU | HELIUM CYLINDER EMPTY | REPLACED THE HELIUM CYLINCDER AND ISSUE RECTIFIED. | HELIUM CYLINDER | NO | 2024-05-01 | May 2024 | 13:20 | 2024-05-01 13:00 | 2024-05-01 13:20 | 0:20 | COMPLETED | CRITICAL
1576 | 2024-05-01 | May 2024 | 13:30 | 13:34 | 0.0027778 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | ICU | PRE-USE CHECK FAILED | RE-RUNNED THE PRE-USE CHECK. ISSUE RECTIFIED. | NO | NO | 2024-05-01 | May 2024 | 13:50 | 2024-05-01 13:30 | 2024-05-01 13:50 | 0:20 | COMPLETED | CRITICAL
1577 | 2024-05-01 | May 2024 | 15:30 | 15:35 | 0.0034722 | BIPAP MACHINE | PHILIPS | A40 | HAFIZ | ICU | FILTER ISSUE | CLEANED AND REINSTALLED THE FILTER. ISSUE RECTIFIED. | NO | NO | 2024-05-01 | May 2024 | 15:50 | 2024-05-01 15:30 | 2024-05-01 15:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1578 | 2024-05-01 | May 2024 | 16:00 | 16:02 | 0.0013889 | ENTERAL FEEDING PUMP | FRESENIUS | AMIKA | HAFIZ | ICU | CALIBRATION DUE | CALIBRATION DONE AND ISSUE RECTIFIED. | NO | NO | 2024-05-01 | May 2024 | 16:30 | 2024-05-01 16:00 | 2024-05-01 16:30 | 0:30 | COMPLETED | NON-CRITICAL
1579 | 2024-05-02 | May 2024 | 0:15 | 0:20 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | SOORAJ | OT | O2 CELL ISSUE | FOUND O2 SENSOR CABLE BROKEN. CHANGED CABLE FROM DEFECTIVE MACHINE AND ISSUE RECTIFIED | O2 SENSOR CABLE | NO | 2024-05-02 | May 2024 | 1:00 | 2024-05-02 00:15 | 2024-05-02 01:00 | 0:45 | COMPLETED | CRITICAL
1580 | 2024-05-02 | May 2024 | 2:15 | 2:20 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | REDUCED SODALIME QTY AND CHANGED PATIENT CIRCUIT, CHECKOUT DONE AND PASSED. ISSUE RECTIFIED | NO | NO | 2024-05-02 | May 2024 | 2:35 | 2024-05-02 02:15 | 2024-05-02 02:35 | 0:20 | COMPLETED | CRITICAL
1581 | 2024-05-02 | May 2024 | 3:15 | 3:15 | 0 | NITRIC OXIDE DELIVERY SYSTEM | WIPRO GE HEALTH CARE | SLE | SOORAJ | OT | NO2 HIGH ALARM | FOUND ISSUE WITH PRESSURE GUAGE. ADJUSTED AND ISSUE RECTIFIED | NO | NO | 2024-05-02 | May 2024 | 4:00 | 2024-05-02 03:15 | 2024-05-02 04:00 | 0:45 | COMPLETED | CRITICAL
1582 | 2024-05-02 | May 2024 | 4:20 | 4:25 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SOORAJ | DIALYSIS | NOT ABLE TO PERFORM THERAPY | FOUND PREOP CHECK NOT PERFORMED. PERFORMED PREOP CHECK AND ISSUE RECTIFIED. | NO | NO | 2024-05-02 | May 2024 | 4:45 | 2024-05-02 04:20 | 2024-05-02 04:45 | 0:25 | COMPLETED | CRITICAL
1583 | 2024-05-02 | May 2024 | 6:00 | 6:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | SOORAJ | WARD | SPO2 VALUE NOT SHOWING | FOUND ISSUE WITH SPO2 SENSOR. ALTERNOTE SPO2 SENSOR ARRANGED FROM WARD SIDE MONITOR. ISSUE RECTIFIED | SPO2 PROBE | NO | 2024-05-02 | May 2024 | 6:30 | 2024-05-02 06:00 | 2024-05-02 06:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1584 | 2024-05-02 | May 2024 | 20:30 | 20:35 | 0.0034722 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | SOORAJ | RADIOLOGY | SHOCK OBSERVED WHILE TOUCHING THE MACHINE | FOUND EARTH WIRE DISCONNECTED INSIDE EXTENSION BOX. RECONNECTED EARTH WIRE AND ISSUE RECTIFIED. | NO | NO | 2024-05-02 | May 2024 | 21:30 | 2024-05-02 20:30 | 2024-05-02 21:30 | 1:00 | COMPLETED | SEMI-CRITICAL
1585 | 2024-05-02 | May 2024 | 9:10 | 9:15 | 0.0034722 | VERTICAL AUTOCLAVE | VANI SCIENTIFIC & EQUIPMENTS | NA | THANISH | LABORATORY | NOT SWITCHING ON | RECTIFIED THE PLUG TOP ISSUE | NO | NO | 2024-05-02 | May 2024 | 9:30 | 2024-05-02 09:10 | 2024-05-02 09:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1586 | 2024-05-02 | May 2024 | 10:15 | 10:20 | 0.0034722 | EEG MACHINE | NOTUS NEUROLOGY | NICOLET EEG V32 | THANISH | OPD | PATIENT REPORTING SYSTEM SOFTWARE FAILED | RENEWED THE SOFTWARE LICENCE KEY AND PATCH FILE UPLOADED FOR EEG SOFTWARE | NO | NO | 2024-05-02 | May 2024 | 10:45 | 2024-05-02 10:15 | 2024-05-02 10:45 | 0:30 | COMPLETED | SEMI-CRITICAL
1587 | 2024-05-02 | May 2024 | 13:05 | 13:10 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | THANISH | OT | TOUCH NOT WORKING | PERFORMED THE TOUCH CALIBRATION | NO | NO | 2024-05-02 | May 2024 | 13:40 | 2024-05-02 13:05 | 2024-05-02 13:40 | 0:35 | COMPLETED | SEMI-CRITICAL
1588 | 2024-05-02 | May 2024 | 9:40 | 9:50 | 0.0069444 | CATHLAB - BIPLANE | PHILIPS | AZURION 7B20 | RAMKIE | CATHLAB | OT LIGHT NOT WORKING (CEILING SPOT) | CABLE DISCONNECTION, RECONNECTED AND RECTIFIED | NO | NO | 2024-05-02 | May 2024 | 10:10 | 2024-05-02 09:40 | 2024-05-02 10:10 | 0:30 | COMPLETED | CRITICAL
1589 | 2024-05-03 | May 2024 | 21:30 | 21:35 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SOORAJ | ICU | CARTRIDGE EMPTY | REPLACED WITH NEW CARTRIDGE AND CVP PASSED. | CARTRIDGE | NO | 2024-05-03 | May 2024 | 22:45 | 2024-05-03 21:30 | 2024-05-03 22:45 | 1:15 | COMPLETED | SEMI-CRITICAL
1590 | 2024-05-03 | May 2024 | 14:45 | 14:50 | 0.0034722 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | HAFIZ | RADIOLOGY | BATTERY ISSUE | ISSUED ONE NEW DETECTOR BATTERY. | DETECTOR BATTERY | NO | 2024-05-03 | May 2024 | 15:10 | 2024-05-03 14:45 | 2024-05-03 15:10 | 0:25 | COMPLETED | SEMI-CRITICAL
1591 | 2024-05-03 | May 2024 | 18:00 | 18:02 | 0.0013889 | CENTRIFUGE TABLE TOP | REMI | NEYA 4 | HAFIZ | LABORATORY | NOT WORKING | SETTINGS CHANGED AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-05-03 | May 2024 | 18:20 | 2024-05-03 18:00 | 2024-05-03 18:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1592 | 2024-05-03 | May 2024 | 9:20 | 9:25 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | THANISH | OT | COT NOT WORKING | RESETTED THE CONTROL BOX.WORKING FINE | NO | NO | 2024-05-03 | May 2024 | 9:45 | 2024-05-03 09:20 | 2024-05-03 09:45 | 0:25 | COMPLETED | NON-CRITICAL
1593 | 2024-05-03 | May 2024 | 10:10 | 10:15 | 0.0034722 | VENTILATOR, NEONATAL | MAQUET | SERVO-I | THANISH | ICU | TIDAL VOLUME ISSUE | CALIBRATED THE FLOW SENSOR.WORKING FINE. | NO | NO | 2024-05-03 | May 2024 | 10:40 | 2024-05-03 10:10 | 2024-05-03 10:40 | 0:30 | COMPLETED | CRITICAL
1594 | 2024-05-03 | May 2024 | 10:50 | 10:55 | 0.0034722 | HEATER COOLER | GETINGE | HCU40 | THANISH | OT | BLANKET LEAKAGE | PATCHWORK DONE FOR HEMOTHERM BLANKET | NO | NO | 2024-05-03 | May 2024 | 11:15 | 2024-05-03 10:50 | 2024-05-03 11:15 | 0:25 | COMPLETED | CRITICAL
1595 | 2024-05-03 | May 2024 | 8:50 | 8:55 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | JERLIN | OT | CIRCUIT LEAK TEST FAILED | RERAN CHECKOUT AND RECTIFIED THE ISSUE | NO | NO | 2024-05-03 | May 2024 | 9:15 | 2024-05-03 08:50 | 2024-05-03 09:15 | 0:25 | COMPLETED | CRITICAL
1596 | 2024-05-03 | May 2024 | 10:10 | 10:20 | 0.0069444 | CENTRIFUGE TABLE TOP | REMI | NEYA 4 | RAMKIE | LABORATORY | NOT WORKING | REPLACED PCB | PCB | NO | 2024-05-03 | May 2024 | 10:40 | 2024-05-03 10:10 | 2024-05-03 10:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1597 | 2024-05-04 | May 2024 | 10:30 | 10:35 | 0.0034722 | CATHLAB - BIPLANE | PHILIPS | AZURION 7B20 | SANTHOSH | CATHLAB | BLURRED IMAGE | RESET DONE ,NOW WORKING | NO | NO | 2024-05-04 | May 2024 | 10:40 | 2024-05-04 10:30 | 2024-05-04 10:40 | 0:10 | COMPLETED | CRITICAL
1598 | 2024-05-04 | May 2024 | 10:35 | 10:45 | 0.0069444 | CATHLAB - BIPLANE | PHILIPS | AZURION 7B20 | SANTHOSH | CATHLAB | EXAMINOTION LED LIGHT NOT WORKING | ADJUSTED THE LED LIGHT CABLE | NO | NO | 2024-05-04 | May 2024 | 10:45 | 2024-05-04 10:35 | 2024-05-04 10:45 | 0:10 | COMPLETED | CRITICAL
1599 | 2024-05-04 | May 2024 | 12:00 | 12:05 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | SANTHOSH | OPD | PRINTER ISSUE | CLENED AND RESET DONE | NO | NO | 2024-05-04 | May 2024 | 12:10 | 2024-05-04 12:00 | 2024-05-04 12:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1600 | 2024-05-04 | May 2024 | 11:10 | 11:15 | 0.0034722 | PRESSURE INJECTOR,DUAL | LF OPTI VONTAGE | 846700A | THANISH | RADIOLOGY | NOT WORKING | CLEANED THE MAINBOARD,ADJUSTED THE DUAL SYRINGE KNOBS&WORKING FINE. | NO | NO | 2024-05-04 | May 2024 | 11:40 | 2024-05-04 11:10 | 2024-05-04 11:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1601 | 2024-05-04 | May 2024 | 12:05 | 12:10 | 0.0034722 | EXAMINATION COUCH | MIDMARK JANOK | E122012 | THANISH | OPD | NOT WORKING | REMOTE FAILED/REPLACED WITH ANOTHER REMOTE | DUAL FUNCTION REMOTE | NO | 2024-05-04 | May 2024 | 12:30 | 2024-05-04 12:05 | 2024-05-04 12:30 | 0:25 | COMPLETED | NON-CRITICAL
1602 | 2024-05-04 | May 2024 | 14:15 | 14:20 | 0.0034722 | NITRIC OXIDE DELIVERY SYSTEM | WIPRO GE HEALTH CARE | SLE | THANISH | ICU | NITROUS LEAKGE | REPAIRED THE NITROUS GAUGE.WORKING FINE. | WASHER | NO | 2024-05-04 | May 2024 | 14:50 | 2024-05-04 14:15 | 2024-05-04 14:50 | 0:35 | COMPLETED | CRITICAL
1603 | 2024-05-05 | May 2024 | 22:00 | 22:05 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SOORAJ | WARD | COT NOT WORKING | CHECKED AND FOUND NO ISSUES WITH COT | NO | NO | 2024-05-05 | May 2024 | 22:30 | 2024-05-05 22:00 | 2024-05-05 22:30 | 0:30 | COMPLETED | NON-CRITICAL
1604 | 2024-05-05 | May 2024 | 22:30 | 22:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SOORAJ | WARD | COT NOT WORKING | FOUND FUNCTIONS LOCKED. LOCK RELEASED AND ISSUE RECTIFIED. | NO | NO | 2024-05-05 | May 2024 | 22:50 | 2024-05-05 22:30 | 2024-05-05 22:50 | 0:20 | COMPLETED | NON-CRITICAL
1605 | 2024-05-05 | May 2024 | 22:50 | 22:50 | 0 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | WARD | COT NOT WORKING | FOUND FUNCTIONS LOCKED. LOCK RELEASED AND ISSUE RECTIFIED. | NO | NO | 2024-05-05 | May 2024 | 23:30 | 2024-05-05 22:50 | 2024-05-05 23:30 | 0:40 | COMPLETED | NON-CRITICAL
1606 | 2024-05-05 | May 2024 | 15:20 | 15:25 | 0.0034722 | USG MACHINE | PHILIPS | CX-50 | SANTHOSH | RADIOLOGY | SOFTWARE ISSUE | RESET DONE ,NOW WORKING | NO | NO | 2024-05-05 | May 2024 | 15:30 | 2024-05-05 15:20 | 2024-05-05 15:30 | 0:10 | COMPLETED | SEMI-CRITICAL
1607 | 2024-05-06 | May 2024 | 0:30 | 0:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | WARD | COT NOT WORKING | FOUND FUNCTIONS LOCKED. RESET DONE AND FUNCTIONS UNLOCKED. | NO | NO | 2024-05-06 | May 2024 | 1:30 | 2024-05-06 00:30 | 2024-05-06 01:30 | 1:00 | COMPLETED | NON-CRITICAL
1608 | 2024-05-06 | May 2024 | 17:05 | 17:10 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | HAFIZ | CATHLAB | ECG CABLE DEFECTIVE | 5 LEADS ECG CABLE GIVEN | 5-LEADS ECG CABLE | NO | 2024-05-06 | May 2024 | 17:25 | 2024-05-06 17:05 | 2024-05-06 17:25 | 0:20 | COMPLETED | SEMI-CRITICAL
1609 | 2024-05-06 | May 2024 | 17:35 | 17:37 | 0.0013889 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | ENDOSCOPY | NIBP MEASUREMENT FAILED | ISSUED NEW BP CONNECTOR, OBSERVED AND WORKING GOOD. | BP CONNECTOR | NO | 2024-05-06 | May 2024 | 17:45 | 2024-05-06 17:35 | 2024-05-06 17:45 | 0:10 | COMPLETED | SEMI-CRITICAL
1610 | 2024-05-06 | May 2024 | 17:45 | 17:45 | 0 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | ENDOSCOPY | NIBP MEASUREMENT FAILED | ISSUED NEW BP CONNECTOR, OBSERVED AND WORKING GOOD. | BP CONNECTOR | NO | 2024-05-06 | May 2024 | 18:00 | 2024-05-06 17:45 | 2024-05-06 18:00 | 0:15 | COMPLETED | SEMI-CRITICAL
1611 | 2024-05-06 | May 2024 | 23:00 | 23:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | JERLIN | WARD | PRINTER NOT WORKING | FIXED THE PAPER PROPERLY AD RECTIFIED THE ISSUE | NO | NO | 2024-05-06 | May 2024 | 23:25 | 2024-05-06 23:00 | 2024-05-06 23:25 | 0:25 | COMPLETED | CRITICAL
1612 | 2024-05-06 | May 2024 | 11:10 | 11:20 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | OT | NOT WORKING | RESET DONE AND UNLOCKED | NO | NO | 2024-05-06 | May 2024 | 11:40 | 2024-05-06 11:10 | 2024-05-06 11:40 | 0:30 | COMPLETED | NON-CRITICAL
1613 | 2024-05-07 | May 2024 | 13:20 | 13:25 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | SANTHOSH | ICU | O2 CELL FAILED | REPLACED NEW CELL | O2 CELL | NO | 2024-05-07 | May 2024 | 13:40 | 2024-05-07 13:20 | 2024-05-07 13:40 | 0:20 | COMPLETED | CRITICAL
1614 | 2024-05-07 | May 2024 | 21:00 | 21:10 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | CARTRIDGE EXPIRED | REPLACED THE CARTRIDGE AND RECTIFIED THE ISSUE | CARTRIDGE | NO | 2024-05-07 | May 2024 | 21:35 | 2024-05-07 21:00 | 2024-05-07 21:35 | 0:35 | COMPLETED | SEMI-CRITICAL
1615 | 2024-05-07 | May 2024 | 7:00 | 7:05 | 0.0034722 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | JERLIN | OPD | V3 LEAD CABLE BROKEN | CHANGED THE V3 LEAD CABLE AND RECTIFIED THE ISSUE | LEAD CABLE | NO | 2024-05-07 | May 2024 | 7:15 | 2024-05-07 07:00 | 2024-05-07 07:15 | 0:15 | COMPLETED | NON-CRITICAL
1616 | 2024-05-07 | May 2024 | 23:15 | 23:20 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | ICU | SET FLOW RATE NOT OBTAINED | REPLACED THE DOOR HINGES AND RECTIFIED THE ISSUE | DOOR ASSEMBLY | NO | 2024-05-07 | May 2024 | 23:45 | 2024-05-07 23:15 | 2024-05-07 23:45 | 0:30 | COMPLETED | NON-CRITICAL
1617 | 2024-05-07 | May 2024 | 23:35 | 23:40 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | ICU | OCCLUSION SENSOR NOT WORKING | REPLACED THE OCCLUSION SENSOR AND RECTIFIED THE ISSUE | PRESSURE SENSOR | NO | 2024-05-07 | May 2024 | 23:55 | 2024-05-07 23:35 | 2024-05-07 23:55 | 0:20 | COMPLETED | NON-CRITICAL
1618 | 2024-05-07 | May 2024 | 10:10 | 10:20 | 0.0069444 | VENTILATOR, ADULT | MAQUET | SERVO-I | RAMKIE | ICU | PRE TEST FAILED | EXPIRATORY DIAPHRAM CLEANED AND RECTIFIED | NO | NO | 2024-05-07 | May 2024 | 10:40 | 2024-05-07 10:10 | 2024-05-07 10:40 | 0:30 | COMPLETED | CRITICAL
1619 | 2024-05-08 | May 2024 | 18:00 | 18:02 | 0.0013889 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | WARD | NOT WORKING | SETTINGS CONFIGURED AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-05-08 | May 2024 | 18:10 | 2024-05-08 18:00 | 2024-05-08 18:10 | 0:10 | COMPLETED | NON-CRITICAL
1620 | 2024-05-08 | May 2024 | 2:00 | 2:15 | 0.010417 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | JERLIN | ICU | SPO2 SENSOR NOT WORKING | REPLACED THE SENSOR CABLE AND RECTIFIED THE ISSUE | NO | NO | 2024-05-08 | May 2024 | 14:55 | 2024-05-08 02:00 | 2024-05-08 14:55 | 12:55 | COMPLETED | SEMI-CRITICAL
1621 | 2024-05-08 | May 2024 | 12:10 | 12:20 | 0.0069444 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | RAMKIE | RADIOLOGY | DETECTOR NOT READY | CHECKED AND FOUND THE PROBLEM WITH DETECTOR BATTRY | NO | NO | 2024-05-08 | May 2024 | 12:40 | 2024-05-08 12:10 | 2024-05-08 12:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1622 | 2024-05-09 | May 2024 | 8:45 | 8:50 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | HAFIZ | CATHLAB | CIRCUIT LEAK TEST FAILED | FOUND ISSUE WITH BELLOWS, REPLACED IT AND OBSERVED. NOW RECTIFIED. | NO | NO | 2024-05-09 | May 2024 | 9:00 | 2024-05-09 08:45 | 2024-05-09 09:00 | 0:15 | COMPLETED | CRITICAL
1623 | 2024-05-09 | May 2024 | 15:00 | 15:06 | 0.0041667 | PATIENT MONITOR | PHILIPS | MX 450 | HAFIZ | ENDOSCOPY | GM OCCLUSION ALARM | FOUND FLUID INSIDE WATER TRAP. REPLACED AND ISSUE RECTIFIED. | WATER TRAP | NO | 2024-05-09 | May 2024 | 15:30 | 2024-05-09 15:00 | 2024-05-09 15:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1624 | 2024-05-09 | May 2024 | 18:00 | 18:02 | 0.0013889 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | HAFIZ | ICU | PROBE NOT DETECTING | PROPERLY FIXED AND ISSUE RECTIFIED. | NO | NO | 2024-05-09 | May 2024 | 18:10 | 2024-05-09 18:00 | 2024-05-09 18:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1625 | 2024-05-09 | May 2024 | 9:00 | 9:05 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | ICU | UNCALIBRATED ALARM ISSUE | SETTINGS CONFIGURED AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-05-10 | May 2024 | 9:15 | 2024-05-09 09:00 | 2024-05-10 09:15 | 24:15 | COMPLETED | NON-CRITICAL
1626 | 2024-05-09 | May 2024 | 9:15 | 9:17 | 0.0013889 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | HAFIZ | ICU | HOLDER CLAMP ISSUE | REPLACED THE HOLDER CLAMP AND ISSUE RECTIFIED. | HOLDER CLAMP | NO | 2024-05-10 | May 2024 | 9:30 | 2024-05-09 09:15 | 2024-05-10 09:30 | 24:15 | COMPLETED | NON-CRITICAL
1627 | 2024-05-09 | May 2024 | 20:30 | 20:40 | 0.0069444 | IMMUNOLOGY ANALYSER | COBAS E 601 | ROCHE | JERLIN | LABORATORY | REAGENT PROBE NOT WORKING | FOUND THAT PROBE CABLE WAS DISCONNECTED FROM THE CONNECTOR. FIXED IT AND RECTIFIED THE ISSUE | NO | NO | 2024-05-09 | May 2024 | 20:50 | 2024-05-09 20:30 | 2024-05-09 20:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1628 | 2024-05-09 | May 2024 | 23:00 | 23:15 | 0.010417 | PATIENT MONITOR | PHILIPS | MX 450 | JERLIN | ICU | NOT ABLE TO READ BLOOD PRESSURE | FOUND LEAK IN THE LARGE ADULT BP CUFF. FIXED IT AND RECTIFIED THE ISSUE | NO | NO | 2024-05-09 | May 2024 | 23:35 | 2024-05-09 23:00 | 2024-05-09 23:35 | 0:35 | COMPLETED | SEMI-CRITICAL
1629 | 2024-05-09 | May 2024 | 23:25 | 23:30 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | JERLIN | ICU | NOT ABLE TO READ BLOOD PRESSURE | FOUND LEAK IN THE LARGE ADULT BP CUFF. FIXED IT AND RECTIFIED THE ISSUE | NO | NO | 2024-05-09 | May 2024 | 23:50 | 2024-05-09 23:25 | 2024-05-09 23:50 | 0:25 | COMPLETED | SEMI-CRITICAL
1630 | 2024-05-09 | May 2024 | 22:30 | 22:35 | 0.0034722 | MOVEMENT THERAPHY UNIT, UPPER AND LOWER LIMBS | RECK | MOTOMED LETTO2 | JERLIN | PHYSIOTHERAPY | BRAKE NOT WORKING PROPERLY | TIGHTENED THE BRAKE SCREWS AND RECTIFIED THE ISSUE | NO | NO | 2024-05-09 | May 2024 | 22:55 | 2024-05-09 22:30 | 2024-05-09 22:55 | 0:25 | COMPLETED | SEMI-CRITICAL
1631 | 2024-05-09 | May 2024 | 12:20 | 12:30 | 0.0069444 | BLOOD COLLECTION MONITOR | FRESENIUS | COMPOGUARD | RAMKIE | BLOODBANK | ERROR | CHECKED AND CHANGED THE PROGRAM | NO | NO | 2024-05-09 | May 2024 | 12:40 | 2024-05-09 12:20 | 2024-05-09 12:40 | 0:20 | COMPLETED | SEMI-CRITICAL
1632 | 2024-05-10 | May 2024 | 18:00 | 18:05 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SOORAJ | WARD | COT UP FUNCTIONS NOT WORKING | COT FUNCTINS UNLOCKED AND ISSUE RECTIFIED. | NO | NO | 2024-05-10 | May 2024 | 18:30 | 2024-05-10 18:00 | 2024-05-10 18:30 | 0:30 | COMPLETED | NON-CRITICAL
1633 | 2024-05-10 | May 2024 | 12:00 | 12:03 | 0.0020833 | PATIENT MONITOR | PHILIPS | MX 450 | HAFIZ | ICU | NIBP MEASUREMENT FAILED | AIR LEAK ARRESTED AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-05-10 | May 2024 | 12:30 | 2024-05-10 12:00 | 2024-05-10 12:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1634 | 2024-05-10 | May 2024 | 10:05 | 10:10 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | THANISH | OT | NOT SWITCHING ON | MOISTURE FORMATION ISSUE IN PMB BOARD AND CLEANED THE PMB BOARD.WORKING FINE. | NO | NO | 2024-05-10 | May 2024 | 10:30 | 2024-05-10 10:05 | 2024-05-10 10:30 | 0:25 | COMPLETED | CRITICAL
1635 | 2024-05-11 | May 2024 | 13:00 | 13:04 | 0.0027778 | LABOUR BED | ARJO HUNTLEIGH | HCB-04LY/BXA | HAFIZ | ICU | NOT WORKING | RESET DONE AND ISSUE RECTIFIED. | NO | NO | 2024-05-11 | May 2024 | 13:10 | 2024-05-11 13:00 | 2024-05-11 13:10 | 0:10 | COMPLETED | NON-CRITICAL
1636 | 2024-05-11 | May 2024 | 19:00 | 19:05 | 0.0034722 | WASHER DISINFECTOR | GETINGE | 46-5-403 | HAFIZ | CSSD | DOOR LOCKING ISSUE | SETTINGS CONFIGURED AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-05-11 | May 2024 | 19:20 | 2024-05-11 19:00 | 2024-05-11 19:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1637 | 2024-05-11 | May 2024 | 19:20 | 19:22 | 0.0013889 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | HAFIZ | ICU | HOLDER CLAMP ISSUE | REPLACED THE HOLDER CLAMP AND ISSUE RECTIFIED. | HOLDER CLAMP | NO | 2024-05-11 | May 2024 | 19:35 | 2024-05-11 19:20 | 2024-05-11 19:35 | 0:15 | COMPLETED | NON-CRITICAL
1638 | 2024-05-11 | May 2024 | 21:15 | 21:30 | 0.010417 | CAMERA SYSTEM | KARL STORZ | TC200 | JERLIN | OT | NOISE APPEARED IN THE DISPLAY | CLEANED THE CAMERA CABLE AND RECTIFIED THE ISSUE | NO | NO | 2024-05-11 | May 2024 | 22:00 | 2024-05-11 21:15 | 2024-05-11 22:00 | 0:45 | COMPLETED | SEMI-CRITICAL
1639 | 2024-05-12 | May 2024 | 23:00 | 23:15 | 0.010417 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | JERLIN | WARD | SPO2 SENSOR NOT WORKING | FOUND THAT SPO2 SENSOR CABLE CONNECTOR WAS DISCONNECTED FROM THE MAIN BOARD, FIXED IT AND ISSUE RECTIFIED | NO | NO | 2024-05-12 | May 2024 | 23:35 | 2024-05-12 23:00 | 2024-05-12 23:35 | 0:35 | COMPLETED | SEMI-CRITICAL
1640 | 2024-05-12 | May 2024 | 19:30 | 19:35 | 0.0034722 | BLOOD GAS ANALYSER | GEM 3500 | INSTRUMENTATION LABORATORY | JERLIN | ICU | CARTRIDGE EXPIRED | REPLACED THE CARTRIDGE AND RECTIFIED THE ISSUE | CARTRIDGE | NO | 2024-05-12 | May 2024 | 19:55 | 2024-05-12 19:30 | 2024-05-12 19:55 | 0:25 | COMPLETED | SEMI-CRITICAL
1641 | 2024-05-13 | May 2024 | 12:15 | 12:20 | 0.0034722 | VENTILATOR, PEDIATRIC | MINDRAY MEDICAL INDIA | SV 300 | SOORAJ | ICU | CHECKOUT FAILED | FOUND O2 SENSOR TEST FAILED. O2 SENSOR CALIBRATION PERFORMED AND FAILED. O2 SENSOR AND BATTERY GONE DEFECTIVE | NO | O2 SENSOR | 2024-05-13 | May 2024 | 13:00 | 2024-05-13 12:15 | 2024-05-13 13:00 | 0:45 | PENDING | CRITICAL
1642 | 2024-05-13 | May 2024 | 13:00 | 13:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | SOORAJ | ICU | FRONT PANEL FOUND OPEN | FOUND INTERNAL DAMAGE ON THE BASE OF FRONT PANEL WHERE SCREW IS SECURED. TEMPORARILY RECTIFIED | NO | NO | 2024-05-13 | May 2024 | 14:00 | 2024-05-13 13:00 | 2024-05-13 14:00 | 1:00 | COMPLETED | SEMI-CRITICAL
1643 | 2024-05-13 | May 2024 | 8:45 | 8:50 | 0.0034722 | BMD MACHINE | HOLOGIC | ASY-05119 | HAFIZ | RADIOLOGY | DAILY QC TEST FAILED | DONE BASIC TROUBLESHOOTING, POSITIONED THE PHANTOM AND OBSERVED. NOW ISSUE RECTIFIED. | NO | NO | 2024-05-13 | May 2024 | 9:05 | 2024-05-13 08:45 | 2024-05-13 09:05 | 0:20 | COMPLETED | SEMI-CRITICAL
1644 | 2024-05-13 | May 2024 | 9:20 | 9:20 | 0 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | CLEANED THE ABSORBER CANISTER AND ISSUE RECTIFIED. | NO | NO | 2024-05-13 | May 2024 | 9:40 | 2024-05-13 09:20 | 2024-05-13 09:40 | 0:20 | COMPLETED | CRITICAL
1645 | 2024-05-13 | May 2024 | 12:00 | 12:02 | 0.0013889 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX | HAFIZ | OT | NOT WORKING | RESET DONE AND ISSUE RECTIFIED. | NO | NO | 2024-05-13 | May 2024 | 12:10 | 2024-05-13 12:00 | 2024-05-13 12:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1646 | 2024-05-13 | May 2024 | 13:00 | 13:02 | 0.0013889 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | HAFIZ | OT | NOT WORKING | LEAK ARRESTED AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-05-13 | May 2024 | 13:25 | 2024-05-13 13:00 | 2024-05-13 13:25 | 0:25 | COMPLETED | SEMI-CRITICAL
1647 | 2024-05-13 | May 2024 | 3:00 | 3:15 | 0.010417 | VENTILATOR, ADULT | MAQUET | SERVO-I | JERLIN | ICU | AIR FILTER BROKEN | REPLACED AIR FILTER FROM ANOTHER VENTILATOR AND RECTIFIED THE ISSUE | NO | NO | 2024-05-13 | May 2024 | 3:20 | 2024-05-13 03:00 | 2024-05-13 03:20 | 0:20 | COMPLETED | CRITICAL
1648 | 2024-05-13 | May 2024 | 12:30 | 12:40 | 0.0069444 | BINOCULAR MICROSCOPE | ZEISS | PRIMO STAR | RAMKIE | LABORATORY | VISION NOT CLEAR | CHECKED AND CLEANED THE LENS AND RECTIFIED | NO | NO | 2024-05-13 | May 2024 | 12:40 | 2024-05-13 12:30 | 2024-05-13 12:40 | 0:10 | COMPLETED | SEMI-CRITICAL
1649 | 2024-05-14 | May 2024 | 17:00 | 17:10 | 0.0069444 | FREEZER,DEEP FREEZING -40 | THERMO FISHER SCIENTIFIC | FORMA 7000 SERIES | SOORAJ | BLOODBANK | HIGH TEMPERATURE ALARM | INFORMED TECH TO DEFROST AND OBSERVE ONCE, AS OF NOW MACHINE IS WORKING. UNDER OBSERVATION | NO | NO | 2024-05-14 | May 2024 | 17:45 | 2024-05-14 17:00 | 2024-05-14 17:45 | 0:45 | COMPLETED | SEMI-CRITICAL
1650 | 2024-05-14 | May 2024 | 7:10 | 7:12 | 0.0013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | HAFIZ | WARD | NOT WORKING | BUTTONS UNLOCKED AND ISSUE RECTIFIED. | NO | NO | 2024-05-14 | May 2024 | 7:20 | 2024-05-14 07:10 | 2024-05-14 07:20 | 0:10 | COMPLETED | NON-CRITICAL
1651 | 2024-05-14 | May 2024 | 8:30 | 8:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | WARD | NOT WORKING | BUTTONS UNLOCKED AND ISSUE RECTIFIED. | NO | NO | 2024-05-14 | May 2024 | 8:45 | 2024-05-14 08:30 | 2024-05-14 08:45 | 0:15 | COMPLETED | NON-CRITICAL
1652 | 2024-05-14 | May 2024 | 9:00 | 9:00 | 0 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | CLEANED THE ABSORBER CANISTER AND ISSUE RECTIFIED. | NO | NO | 2024-05-14 | May 2024 | 9:15 | 2024-05-14 09:00 | 2024-05-14 09:15 | 0:15 | COMPLETED | CRITICAL
1653 | 2024-05-14 | May 2024 | 1:05 | 1:10 | 0.0034722 | CT MACHINE | PHILIPS | INGENUITY | THANISH | RADIOLOGY | GANTRY INITIALIZE ISSUE | REPLACED THE COUPLER IN THE ENCODER ASSEMBLY.WORKINGFINE. | COUPLER | NO | 2024-05-15 | May 2024 | 1:53 | 2024-05-14 01:05 | 2024-05-15 01:53 | 24:48 | COMPLETED | CRITICAL
1654 | 2024-05-14 | May 2024 | 19:40 | 19:50 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | CARTRIDGE EXPIRED | REPLACED THE CARTRIDGE AND RECTIFIED THE ISSUE | CARTRIDGE | NO | 2024-05-14 | May 2024 | 20:05 | 2024-05-14 19:40 | 2024-05-14 20:05 | 0:25 | COMPLETED | SEMI-CRITICAL
1655 | 2024-05-14 | May 2024 | 11:00 | 11:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | RAMKIE | ICU | OP CHECK TEST FAILED | CHECKED AND OP CHECK PASSED | NO | NO | 2024-05-14 | May 2024 | 11:30 | 2024-05-14 11:00 | 2024-05-14 11:30 | 0:30 | COMPLETED | CRITICAL
1656 | 2024-05-15 | May 2024 | 12:00 | 12:05 | 0.0034722 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX 8 | SOORAJ | OT | E232 ERROR | CHECKED AND CONFIRMED CONTROL BOARD DEFECTIVE | NO | CONTROL BOARD | 2024-05-15 | May 2024 | 13:00 | 2024-05-15 12:00 | 2024-05-15 13:00 | 1:00 | PENDING | SEMI-CRITICAL
1657 | 2024-05-15 | May 2024 | 17:15 | 17:20 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SOORAJ | WARD | NO POWER INDICATIONS | FOUND POWER SOCKET NOT WORKING. CONNECTED TO ANOTHER SOCKET AND ISSUE RECTIFIED. | NO | NO | 2024-05-15 | May 2024 | 17:45 | 2024-05-15 17:15 | 2024-05-15 17:45 | 0:30 | COMPLETED | NON-CRITICAL
1658 | 2024-05-15 | May 2024 | 17:35 | 17:35 | 0 | DEFIBRILLATOR MONITOR WITH AED | PHILIPS | EFFICIA DFM 100 | JERLIN | EMERGENCY | THERAPY TEST DISABLED | CHANGED THE MAIN BOARD AND RAN OPERATIONAL CHECK , IT'S WORKING GOOD. | NO | NO | 2024-05-15 | May 2024 | 17:55 | 2024-05-15 17:35 | 2024-05-15 17:55 | 0:20 | COMPLETED | CRITICAL
1659 | 2024-05-16 | May 2024 | 11:35 | 11:40 | 0.0034722 | BMI MACHINE | AKAS MEDICAL | QVS-100 | SOORAJ | OPD | MACHINE NOT SWITCHING ON | FOUND POWER SOCKET NOT WORKING. CHANGED TO ANOTHER SOCKET AND ISSUE RECTIFIED. | NO | NO | 2024-05-16 | May 2024 | 11:50 | 2024-05-16 11:35 | 2024-05-16 11:50 | 0:15 | COMPLETED | NON-CRITICAL
1660 | 2024-05-16 | May 2024 | 11:35 | 11:45 | 0.0069444 | BMD MACHINE | HOLOGIC | ASY-05119 | SOORAJ | RADIOLOGY | NOT ABLE TO PERFORM TEST | FOUND TANK ISSUE. TEMPORARILY KEPT ON SERVICE MODE AND PERFORMED CALIBRATIO 10 TIMES, PASSED AND SCAN DONE ON TWO PATIENTS | NO | NO | 2024-05-16 | May 2024 | 12:30 | 2024-05-16 11:35 | 2024-05-16 12:30 | 0:55 | COMPLETED | SEMI-CRITICAL
1661 | 2024-05-16 | May 2024 | 7:30 | 7:33 | 0.0020833 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | HAFIZ | OT | CALIBRATION DUE | O2 CALIBRATION DONE AND IS NOW READY TO USE. | NO | NO | 2024-05-16 | May 2024 | 7:45 | 2024-05-16 07:30 | 2024-05-16 07:45 | 0:15 | COMPLETED | CRITICAL
1662 | 2024-05-16 | May 2024 | 11:10 | 11:15 | 0.0034722 | BMD MACHINE | HOLOGIC | ASY-05119 | THANISH | RADIOLOGY | NOT WORKING | REPLACED THE TUBE UNIT AND PERFORMED TUBE CALIBRATION.WORKING FINE. | TUBE KIT | NO | 2024-05-16 | May 2024 | 12:00 | 2024-05-16 11:10 | 2024-05-16 12:00 | 0:50 | COMPLETED | SEMI-CRITICAL
1663 | 2024-05-16 | May 2024 | 15:40 | 15:50 | 0.0069444 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | JERLIN | ICU | NOT ABLE TO STORE EXAM DATA | CLEARED THE UNWANTED DATA AND RECTIFIED THE ISSUE | NO | NO | 2024-05-16 | May 2024 | 16:05 | 2024-05-16 15:40 | 2024-05-16 16:05 | 0:25 | COMPLETED | SEMI-CRITICAL
1664 | 2024-05-17 | May 2024 | 7:40 | 7:45 | 0.0034722 | TRANSPORT VENTILATOR | RESMED | ASTRAL 100 | HAFIZ | ICU | NOT WORKING | UNLOCKED THE BIPAP MACHINE, SETTINGS CONFIGURED AND IS NOW READY TO USE. | NO | NO | 2024-05-17 | May 2024 | 8:00 | 2024-05-17 07:40 | 2024-05-17 08:00 | 0:20 | COMPLETED | SEMI-CRITICAL
1665 | 2024-05-17 | May 2024 | 15:04 | 15:10 | 0.0041667 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | WARD | DOOR AND HOLDER CLAMP ISSUE | REPLACED THE HOLDER CLAMP AND DOOR. ISSUE RECTIFIED. | DOOR, HOLDER CLAMP | NO | 2024-05-17 | May 2024 | 15:25 | 2024-05-17 15:04 | 2024-05-17 15:25 | 0:21 | COMPLETED | NON-CRITICAL
1666 | 2024-05-17 | May 2024 | 13:30 | 13:35 | 0.0034722 | PULSE OXIMETER,PORTABLE | COVIDIEN | NELLCOR | JERLIN | ICU | ROTATORY SWITCH KNOB NOT WORKING | FOUND THAT THE DISPLAY PCB WAS DISCONNECTED FROM THE CASING. FIXED IT PROPERLY AND RECTIFIED THE IISUE. | NO | NO | 2024-05-17 | May 2024 | 13:55 | 2024-05-17 13:30 | 2024-05-17 13:55 | 0:25 | COMPLETED | SEMI-CRITICAL
1667 | 2024-05-17 | May 2024 | 17:50 | 18:00 | 0.0069444 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | JERLIN | WARD | NOT ABLE TO READ SPO2 VALUE | FOUND THAT PINS WERE BENT. FIXED IT AND RECTIFIED THE ISSUE. | NO | NO | 2024-05-17 | May 2024 | 18:20 | 2024-05-17 17:50 | 2024-05-17 18:20 | 0:30 | COMPLETED | SEMI-CRITICAL
1668 | 2024-05-18 | May 2024 | 12:10 | 12:15 | 0.0034722 | DIGITAL PLATFORM SCALE | SECA | 676 | SOORAJ | DIALYSIS | ERROR CODE 212 | FOUND CHARGER NOT CONNECTED PROPERLY. RECONNECTED AND ISSUE RECTIFIED. | NO | NO | 2024-05-18 | May 2024 | 12:30 | 2024-05-18 12:10 | 2024-05-18 12:30 | 0:20 | COMPLETED | NON-CRITICAL
1669 | 2024-05-18 | May 2024 | 8:00 | 8:00 | 0 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | ICU | AIR BUBBLE DETECTOR ISSUE | REPLACED THE AIR BUBBLE DETECTOR AND IS NOW READY TO USE. | AIR BUBBLE DETECTOR | NO | 2024-05-18 | May 2024 | 8:30 | 2024-05-18 08:00 | 2024-05-18 08:30 | 0:30 | COMPLETED | NON-CRITICAL
1670 | 2024-05-18 | May 2024 | 8:30 | 8:33 | 0.0020833 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | CLEANED THE ABSORBER CANISTER AND ISSUE RECTIFIED. | NO | NO | 2024-05-18 | May 2024 | 8:45 | 2024-05-18 08:30 | 2024-05-18 08:45 | 0:15 | COMPLETED | CRITICAL
1671 | 2024-05-19 | May 2024 | 14:00 | 14:02 | 0.0013889 | STEAM STERILIZER | GETINGE | GSS67H102E | HAFIZ | CSSD | VACCUUM TIMEOUT ERROR | SETTINGS CONFIGURED AS PER COMPANY PERSON'S GUIDELINES. OBSERVED AND ISSUE RECTIFIED. | NO | NO | 2024-05-19 | May 2024 | 17:00 | 2024-05-19 14:00 | 2024-05-19 17:00 | 3:00 | COMPLETED | SEMI-CRITICAL
1672 | 2024-05-19 | May 2024 | 10:00 | 10:02 | 0.0013889 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | CATHLAB | ECG EQUIPMENT MALFUNCTION | CLEANED THE PROBE AND OBSERVED. NOW ISSUE RECTIFIED. | NO | NO | 2024-05-19 | May 2024 | 10:20 | 2024-05-19 10:00 | 2024-05-19 10:20 | 0:20 | COMPLETED | CRITICAL
1673 | 2024-05-19 | May 2024 | 10:20 | 10:22 | 0.0013889 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | EMERGENCY | ECG EQUIPMENT MALFUNCTION | CLEANED THE PROBE AND OBSERVED. NOW ISSUE RECTIFIED. | NO | NO | 2024-05-19 | May 2024 | 10:35 | 2024-05-19 10:20 | 2024-05-19 10:35 | 0:15 | COMPLETED | CRITICAL
1674 | 2024-05-19 | May 2024 | 16:00 | 16:05 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | WARD | HOLDER CLAMP ISSUE | REPLACED THE HOLDER CLAMP AND ISSUE RECTIFIED. | HOLDER CLAMP | NO | 2024-05-19 | May 2024 | 16:30 | 2024-05-19 16:00 | 2024-05-19 16:30 | 0:30 | COMPLETED | NON-CRITICAL
1675 | 2024-05-20 | May 2024 | 10:10 | 10:15 | 0.0034722 | OT TABLE | MAQUET | ALPHA CLASSIC PRO | SOORAJ | OT | PATIENT SLANTING POSTION WORKING INTERMITTENTLY | APPLIED LIBRICATION AND ISSUE RECTIFIED | NO | NO | 2024-05-20 | May 2024 | 10:50 | 2024-05-20 10:10 | 2024-05-20 10:50 | 0:40 | COMPLETED | NON-CRITICAL
1676 | 2024-05-20 | May 2024 | 19:45 | 19:50 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | HAFIZ | ICU | NEED TO REPLACE A CARTRIDGE | SAME CARTRIDGE LOADED WITH DIFFERENT BARCODE. WARMING UP DONE AND ISSUE RECTIFIED. | CARTRIDGE | NO | 2024-05-20 | May 2024 | 20:15 | 2024-05-20 19:45 | 2024-05-20 20:15 | 0:30 | COMPLETED | SEMI-CRITICAL
1677 | 2024-05-20 | May 2024 | 9:30 | 9:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | ICU | REMOTE NOT WORKING | FOUND THAT REMOTE CABLE WAS DISCONNECTED FROM THE CONTROL BOX. FIXED IT AND ISSUE RECTIFIED | NO | NO | 2024-05-20 | May 2024 | 9:55 | 2024-05-20 09:30 | 2024-05-20 09:55 | 0:25 | COMPLETED | NON-CRITICAL
1678 | 2024-05-20 | May 2024 | 9:30 | 9:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | ICU | REMOTE NOT WORKING | REMOTE WAS NOT CONNECTED TO THE CONTROL BOX. CONNECTED AND RECTIFIED THE ISSUE | NO | NO | 2024-05-20 | May 2024 | 9:55 | 2024-05-20 09:30 | 2024-05-20 09:55 | 0:25 | COMPLETED | NON-CRITICAL
1679 | 2024-05-20 | May 2024 | 9:50 | 10:10 | 0.013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | WARD | NOT WORKING | RESET DONE AND ISSUE RECTIFIED | NO | NO | 2024-05-20 | May 2024 | 10:10 | 2024-05-20 09:50 | 2024-05-20 10:10 | 0:20 | COMPLETED | NON-CRITICAL
1680 | 2024-05-20 | May 2024 | 10:45 | 10:55 | 0.0069444 | RADIOGRAPHIC UNIT | CARESTREAM | DRF ASCEND ANOLOG HF | RAMKIE | OPD | HANDLE LOOSEN | CHECKED AND CONNECTED IN POSITION | NO | NO | 2024-05-20 | May 2024 | 12:40 | 2024-05-20 10:45 | 2024-05-20 12:40 | 1:55 | COMPLETED | SEMI-CRITICAL
1681 | 2024-05-22 | May 2024 | 6:00 | 6:05 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | HAFIZ | WARD | NOT WORKING | BUTTONS UNLOCKED AND ISSUE RECTIFIED. | NO | NO | 2024-05-22 | May 2024 | 6:15 | 2024-05-22 06:00 | 2024-05-22 06:15 | 0:15 | COMPLETED | NON-CRITICAL
1682 | 2024-05-22 | May 2024 | 10:00 | 10:18 | 0.0125 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | RAMKIE | ICU | CATRIDGE FAILURE | CHECKED AND FOUND THE PROBLEM WITH CATRIDGE | NO | NO | 2024-05-22 | May 2024 | 12:40 | 2024-05-22 10:00 | 2024-05-22 12:40 | 2:40 | COMPLETED | SEMI-CRITICAL
1683 | 2024-05-23 | May 2024 | 10:15 | 10:20 | 0.0034722 | DEFIBRILLATOR MONITOR WITH AED | PHILIPS | EFFICIA DFM 100 | SOORAJ | ICU | PRINTER NOT WORKING. | PRINTER ROLLER CLEANED AND CHECKED, ISSUE RECTIFIED. | NO | NO | 2024-05-23 | May 2024 | 10:40 | 2024-05-23 10:15 | 2024-05-23 10:40 | 0:25 | COMPLETED | CRITICAL
1684 | 2024-05-23 | May 2024 | 20:30 | 20:35 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | PRINTER NOT WORKING | CLEANED THE PRINTER HEAD, AND CLEARED THE QUEUE. ISSUE RECTIFIED. | NO | NO | 2024-05-23 | May 2024 | 20:50 | 2024-05-23 20:30 | 2024-05-23 20:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1685 | 2024-05-23 | May 2024 | 20:55 | 21:00 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | HAFIZ | ICU | NOT WORKING | REFIXED THE THERMISTOR CABLE AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-05-23 | May 2024 | 21:10 | 2024-05-23 20:55 | 2024-05-23 21:10 | 0:15 | COMPLETED | SEMI-CRITICAL
1686 | 2024-05-23 | May 2024 | 21:20 | 21:22 | 0.0013889 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | WARD | NOT WORKING | SETTINGS CONFIGURED AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-05-23 | May 2024 | 21:30 | 2024-05-23 21:20 | 2024-05-23 21:30 | 0:10 | COMPLETED | NON-CRITICAL
1687 | 2024-05-24 | May 2024 | 11:30 | 11:35 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | SOORAJ | OT | TEE PROBE ANGULATION STUCK DURING CASE | RESTARTED THE SYSTEM AND ISSUE RECTIFIED. | NO | NO | 2024-05-24 | May 2024 | 11:45 | 2024-05-24 11:30 | 2024-05-24 11:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1688 | 2024-05-24 | May 2024 | 0:45 | 0:50 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | HAFIZ | OT | NOT WORKING | RESET DONE AND ISSUE RECTIFIED. | NO | NO | 2024-05-24 | May 2024 | 1:00 | 2024-05-24 00:45 | 2024-05-24 01:00 | 0:15 | COMPLETED | SEMI-CRITICAL
1689 | 2024-05-24 | May 2024 | 6:00 | 6:02 | 0.0013889 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | HAFIZ | EMERGENCY | HOLDER CLAMP ISSUE | REPLACED THE HOLDER CLAMP AND ISSUE RECTIFIED. | HOLDER CLAMP | NO | 2024-05-24 | May 2024 | 6:30 | 2024-05-24 06:00 | 2024-05-24 06:30 | 0:30 | COMPLETED | NON-CRITICAL
1690 | 2024-05-24 | May 2024 | 6:00 | 6:02 | 0.0013889 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | HAFIZ | EMERGENCY | HOLDER CLAMP ISSUE | REPLACED THE HOLDER CLAMP AND ISSUE RECTIFIED. | HOLDER CLAMP | NO | 2024-05-24 | May 2024 | 6:30 | 2024-05-24 06:00 | 2024-05-24 06:30 | 0:30 | COMPLETED | NON-CRITICAL
1691 | 2024-05-24 | May 2024 | 11:15 | 11:30 | 0.010417 | NEBULIZER | PARI | PARI BOY SX | JERLIN | WARD | NOT ABLE TO FIX CONNECTOR IN THE AIR LUFT | FIXED IT AND RECTIFIED THE ISSUE | NO | NO | 2024-05-24 | May 2024 | 11:50 | 2024-05-24 11:15 | 2024-05-24 11:50 | 0:35 | COMPLETED | SEMI-CRITICAL
1692 | 2024-05-24 | May 2024 | 13:00 | 13:15 | 0.010417 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | WARD | NOT WORKING | POWER CORD WAS REMOVED FROM THE CONTROL BOX , FIXED IT AND RECTIFIED THE ISSUE | NO | NO | 2024-05-24 | May 2024 | 13:35 | 2024-05-24 13:00 | 2024-05-24 13:35 | 0:35 | COMPLETED | NON-CRITICAL
1693 | 2024-05-24 | May 2024 | 15:00 | 15:10 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | CARTRIDGE EXPIRED | REPLACED THE CARTRIDGE AND RECTIFIED THE ISSUE | CARTRIDGE | NO | 2024-05-24 | May 2024 | 15:30 | 2024-05-24 15:00 | 2024-05-24 15:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1694 | 2024-05-24 | May 2024 | 16:30 | 16:40 | 0.0069444 | BLOOD GAS ANALYSER | GEM 3500 | INSTRUMENTATION LABORATORY | JERLIN | ICU | CARTRIDGE EXPIRED | REPLACED THE CARTRIDGE AND RECTIFIED THE ISSUE | CARTRIDGE | NO | 2024-05-24 | May 2024 | 16:55 | 2024-05-24 16:30 | 2024-05-24 16:55 | 0:25 | COMPLETED | SEMI-CRITICAL
1695 | 2024-05-24 | May 2024 | 10:10 | 10:20 | 0.0069444 | DEFIBRILLATOR MONITOR WITH AED | PHILIPS | EFFICIA DFM 100 | RAMKIE | EMERGENCY | TEST FAILED | CHECKED AND OP CHECK PASSED | NO | NO | 2024-05-24 | May 2024 | 10:20 | 2024-05-24 10:10 | 2024-05-24 10:20 | 0:10 | COMPLETED | CRITICAL
1696 | 2024-05-25 | May 2024 | 8:15 | 8:20 | 0.0034722 | HEATER COOLER | GETINGE | HCU 40 | SOORAJ | ICU | WATER FLOW TOO LOW ALARM | FOUND ISSUE IN SETTINGS. CONFIGURED AND ISSUE RECTIFIED. | NO | NO | 2024-05-25 | May 2024 | 8:35 | 2024-05-25 08:15 | 2024-05-25 08:35 | 0:20 | COMPLETED | SEMI-CRITICAL
1697 | 2024-05-25 | May 2024 | 21:30 | 21:35 | 0.0034722 | HEAD LIGHT SOURCE, LED | SUNOPTIC | SSL-5500 | HAFIZ | ICU | POWER ADAPTOR ISSUE | REPLACED A NEW POWER ADAPTOR AND OBSERVED. ISSUE RECTIFIED. | POWER ADAPTOR | NO | 2024-05-25 | May 2024 | 21:45 | 2024-05-25 21:30 | 2024-05-25 21:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1698 | 2024-05-25 | May 2024 | 22:30 | 22:35 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | HAFIZ | ICU | NOT WORKING | REFIXED THE THERMISTOR CABLE AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-05-25 | May 2024 | 22:40 | 2024-05-25 22:30 | 2024-05-25 22:40 | 0:10 | COMPLETED | SEMI-CRITICAL
1699 | 2024-05-25 | May 2024 | 10:10 | 10:15 | 0.0034722 | RADIOGRAPHIC UNIT | CARESTREAM | DR ASCEND CMD | THANISH | RADIOLOGY | COLLIMATOR LAMP NOT WORKING | REPLACED THE COLLIMATOR LAMP | COLLIMATOR LAMP | NO | 2024-05-25 | May 2024 | 10:40 | 2024-05-25 10:10 | 2024-05-25 10:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1700 | 2024-05-25 | May 2024 | 11:20 | 11:40 | 0.013889 | VENTILATOR, UNIVERSAL | MINDRAY MEDICAL INDIA | SV 300 | JERLIN | EMERGENCY | INTERNOL LEAKAGE ABOVE 1000 ML | CHANGED THE INSPIRATORY VALVE AND RECTIFIED THE ISSUE | INSPIRATORY VALVE | NO | 2024-05-25 | May 2024 | 12:00 | 2024-05-25 11:20 | 2024-05-25 12:00 | 0:40 | COMPLETED | CRITICAL
1701 | 2024-05-26 | May 2024 | 20:00 | 20:02 | 0.0013889 | NEBULIZER | PARI | PARI BOY SX | HAFIZ | ICU | NOT WORKING | CONNECTIONS REFIXED IN POWER SOCKET AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-05-26 | May 2024 | 20:15 | 2024-05-26 20:00 | 2024-05-26 20:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1702 | 2024-05-26 | May 2024 | 22:30 | 22:32 | 0.0013889 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | HAFIZ | ICU | K+ FAILED | RERUNNED THE IQM PROCESS. OBSERVED AND ISSUE RECTIFIED. | NO | NO | 2024-05-26 | May 2024 | 22:40 | 2024-05-26 22:30 | 2024-05-26 22:40 | 0:10 | COMPLETED | SEMI-CRITICAL
1703 | 2024-05-26 | May 2024 | 22:20 | 22:21 | 0.00069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | HAFIZ | WARD | NOT WORKING | RESET DONE AND ISSUE RECTIFIED. | NO | NO | 2024-05-26 | May 2024 | 22:30 | 2024-05-26 22:20 | 2024-05-26 22:30 | 0:10 | COMPLETED | NON-CRITICAL
1704 | 2024-05-27 | May 2024 | 19:35 | 19:40 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | SANTHOSH | ICU | NOT GET PRINT | RESET DONE ,NOW WORKING | NO | NO | 2024-05-27 | May 2024 | 19:50 | 2024-05-27 19:35 | 2024-05-27 19:50 | 0:15 | COMPLETED | SEMI-CRITICAL
1705 | 2024-05-27 | May 2024 | 7:50 | 7:55 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | WARD | SIDE RAIL NOT WORKING | LUBRICATED THE HINGES AND RECTIFIED THE ISSUE | NO | NO | 2024-05-27 | May 2024 | 8:15 | 2024-05-27 07:50 | 2024-05-27 08:15 | 0:25 | COMPLETED | NON-CRITICAL
1706 | 2024-05-27 | May 2024 | 14:00 | 14:10 | 0.0069444 | LITHOCLAST | SURGIMEDIK HEALTHCARE | DIGI NOVA | JERLIN | OT | PROBE BROKEN | CHANGED THE PROBE AND RECTIFIED THE ISSUE | LITHOCLAST PROBE | NO | 2024-05-27 | May 2024 | 14:25 | 2024-05-27 14:00 | 2024-05-27 14:25 | 0:25 | COMPLETED | SEMI-CRITICAL
1707 | 2024-05-27 | May 2024 | 10:10 | 10:15 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | JERLIN | OT | CIRCUIT LEAK TEST FAILED | CHANGED THE CIRCUIT AND CHECKED . ISSUE RECCTIFIED | CIRCUIT | NO | 2024-05-27 | May 2024 | 10:35 | 2024-05-27 10:10 | 2024-05-27 10:35 | 0:25 | COMPLETED | CRITICAL
1708 | 2024-05-27 | May 2024 | 14:10 | 14:15 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | CARTRIDGE EXPIRED | REPLACED THE CARTRIDGE AND RECTIFIED THE ISSUE | CARTRIDGE | NO | 2024-05-27 | May 2024 | 14:35 | 2024-05-27 14:10 | 2024-05-27 14:35 | 0:25 | COMPLETED | SEMI-CRITICAL
1709 | 2024-05-28 | May 2024 | 23:30 | 23:35 | 0.0034722 | BIPAP MACHINE | PHILIPS | A40 | SOORAJ | ICU | RESPIRATE RATE TOO HIGH ALARM | FOUND ISSUE WITH SETTINGS. CONFIGURED AND ISSUE RECTIFIED | NO | NO | 2024-05-28 | May 2024 | 23:55 | 2024-05-28 23:30 | 2024-05-28 23:55 | 0:25 | COMPLETED | SEMI-CRITICAL
1710 | 2024-05-28 | May 2024 | 10:30 | 10:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 9000 | SANTHOSH | DIALYSIS | KEYBOARD LOCKED | RESET DONE ,NOW WORKING | NO | NO | 2024-05-28 | May 2024 | 10:45 | 2024-05-28 10:30 | 2024-05-28 10:45 | 0:15 | COMPLETED | NON-CRITICAL
1711 | 2024-05-29 | May 2024 | 2:00 | 2:05 | 0.0034722 | TRANSPORT VENTILATOR | RESMED | ASTRAL 100 | SOORAJ | ICU | CIRCUIT ALARM | FOUND ISSUE WITH CIRCUIT. CHANGED CIRCUIT AND ISSUE RECTIFIED. | NO | NO | 2024-05-29 | May 2024 | 2:20 | 2024-05-29 02:00 | 2024-05-29 02:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1712 | 2024-05-29 | May 2024 | 8:15 | 8:20 | 0.0034722 | EXAMINATION COUCH | MIDMARK JANOK | E122012 | SANTHOSH | OPD | NOT WORKING | ISSUE RECTIFIED | NO | NO | 2024-05-29 | May 2024 | 8:25 | 2024-05-29 08:15 | 2024-05-29 08:25 | 0:10 | COMPLETED | NON-CRITICAL
1713 | 2024-05-29 | May 2024 | 10:15 | 10:20 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | SANTHOSH | CATHLAB | NOT IMAGE SHOWING | CHANGED SETTINGS | NO | NO | 2024-05-29 | May 2024 | 10:30 | 2024-05-29 10:15 | 2024-05-29 10:30 | 0:15 | COMPLETED | SEMI-CRITICAL
1714 | 2024-05-29 | May 2024 | 10:45 | 10:50 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | SANTHOSH | OT | ETCO2 NOT SHOWING | ISSUE RECTIFIED | NO | NO | 2024-05-29 | May 2024 | 11:00 | 2024-05-29 10:45 | 2024-05-29 11:00 | 0:15 | COMPLETED | SEMI-CRITICAL
1715 | 2024-05-30 | May 2024 | 1:30 | 1:35 | 0.0034722 | BIPAP MACHINE | PHILIPS | A40 | SOORAJ | WARD | CIRCUIT LEAK ALARM | CIRCUIT CONNECTED PROPERLY AND ISSUE RECTIFIED | NO | NO | 2024-05-30 | May 2024 | 2:05 | 2024-05-30 01:30 | 2024-05-30 02:05 | 0:35 | COMPLETED | SEMI-CRITICAL
1716 | 2024-05-30 | May 2024 | 6:00 | 6:05 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | SOORAJ | ICU | PATIENT WARMER NOT WORKING | FOUND ONE OF THE HOSE WIRE DISCONNECTED. SOLDERED WIRE AND ISSUE RECTIFED | NO | NO | 2024-05-30 | May 2024 | 7:15 | 2024-05-30 06:00 | 2024-05-30 07:15 | 1:15 | COMPLETED | SEMI-CRITICAL
1717 | 2024-05-30 | May 2024 | 21:00 | 21:05 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SOORAJ | WARD | COT FUNCTIONS LOCKED. | RESET DONE AND ISSUE RECTIFIED | NO | NO | 2024-05-30 | May 2024 | 21:20 | 2024-05-30 21:00 | 2024-05-30 21:20 | 0:20 | COMPLETED | NON-CRITICAL
1718 | 2024-05-30 | May 2024 | 21:30 | 21:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | WARD | COT SIDE RAIL NOT WORKING. | FOUND RUSTING, LUBRICATION DONE AND ISSUE RECTIFIED | NO | NO | 2024-05-30 | May 2024 | 21:50 | 2024-05-30 21:30 | 2024-05-30 21:50 | 0:20 | COMPLETED | NON-CRITICAL
1719 | 2024-05-30 | May 2024 | 21:40 | 21:45 | 0.0034722 | VENTILATOR, ADULT | MAQUET | SERVO-I | SOORAJ | ICU | REGULATION PRESSURE LIMIT ALAMR | ISSUE WITH SETTINGS. SETTINGS CONFIGURED AND ISSUE RECTIFIED. | NO | NO | 2024-05-30 | May 2024 | 22:05 | 2024-05-30 21:40 | 2024-05-30 22:05 | 0:25 | COMPLETED | CRITICAL
1720 | 2024-05-30 | May 2024 | 14:40 | 14:45 | 0.0034722 | THROMBOELASTOMETER | FRAMER HEMOLOGIC | 5000 | SANTHOSH | ICU | NEEDLE BROKEN STUCKED INSIDE | TOOK OUT PROPERLY | NO | NO | 2024-05-30 | May 2024 | 15:10 | 2024-05-30 14:40 | 2024-05-30 15:10 | 0:30 | COMPLETED | SEMI-CRITICAL
1721 | 2024-05-30 | May 2024 | 14:30 | 14:45 | 0.010417 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | JERLIN | RADIOLOGY | NOT ABLE TO FIX THE TAB PROPERLY | FOUND THAT THE TAB STAND WAS BROKEN , CHANGED IT AND ISSUE RECTIFIED | TAB STAND | NO | 2024-05-30 | May 2024 | 15:05 | 2024-05-30 14:30 | 2024-05-30 15:05 | 0:35 | COMPLETED | SEMI-CRITICAL
1722 | 2024-05-30 | May 2024 | 15:55 | 16:00 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | WARD | DOOR BROKEN | CHANGED THE DOOR ASSEMBLY AND RECTIFIED THE ISSUE | DOOR ASSEMBLY | NO | 2024-05-30 | May 2024 | 16:18 | 2024-05-30 15:55 | 2024-05-30 16:18 | 0:23 | COMPLETED | NON-CRITICAL
1723 | 2024-05-31 | May 2024 | 6:30 | 6:35 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SOORAJ | EMERGENCY | ECG EQUIPMENT MALFUNCTION | ISSUE WITH ECG LEAD SET. REPLACED LEAD SET AND ISSUE RECTIFIED | ECG CABLE | NO | 2024-05-31 | May 2024 | 7:00 | 2024-05-31 06:30 | 2024-05-31 07:00 | 0:30 | COMPLETED | CRITICAL
1724 | 2024-05-31 | May 2024 | 21:40 | 21:45 | 0.0034722 | PLASMA STERILIZER | JOHNSON&JOHNSON | STERRAD 100NX | SOORAJ | CSSD | CYCLE CANCELLED | FOUND INTERMITTENT POWER FAILURE. RESTARTED THE SYSTEM AND STARTED CYCLE, ISSUE RECTIFIED. | NO | NO | 2024-05-31 | May 2024 | 22:15 | 2024-05-31 21:40 | 2024-05-31 22:15 | 0:35 | COMPLETED | SEMI-CRITICAL
1725 | 2024-05-31 | May 2024 | 22:30 | 22:35 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | SOORAJ | ICU | POWER INDICATIO LIGHT NOT GLOWING | FOUND LED INDICATORS TO BE DEFECTIVE. OBSERVED BATTERY CHARGE REDUCING DRASTICALLY. BATTERY NEEDS REPLACEMENT BUT AS OF NOW WORKING FINE. | NO | BATTERY | 2024-05-31 | May 2024 | 23:00 | 2024-05-31 22:30 | 2024-05-31 23:00 | 0:30 | PENDING | SEMI-CRITICAL
1726 | 2024-05-31 | May 2024 | 23:15 | 23:20 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | WARD | COT FUNCTIONS LOCKED. | RESET DONE AND ISSUE RECTIFIED | NO | NO | 2024-05-31 | May 2024 | 23:50 | 2024-05-31 23:15 | 2024-05-31 23:50 | 0:35 | COMPLETED | NON-CRITICAL
1727 | 2024-05-31 | May 2024 | 13:20 | 13:25 | 0.0034722 | DEFIBRILLATOR MONITOR WITH AED | PHILIPS | EFFICIA DFM 100 | SANTHOSH | EMERGENCY | CANNOT ANOLYSE ECG | OP CHECK DONE | NO | NO | 2024-05-31 | May 2024 | 13:30 | 2024-05-31 13:20 | 2024-05-31 13:30 | 0:10 | COMPLETED | CRITICAL
1728 | 2024-05-31 | May 2024 | 11:15 | 11:20 | 0.0034722 | BMD MACHINE | HOLOGIC | ASY-05119 | THANISH | RADIOLOGY | DUAL HIP PROCEDURE NOT WORKING | PERFORMED THE TABLE CALIRATION/WORKING FINE. | NO | NO | 2024-05-31 | May 2024 | 11:55 | 2024-05-31 11:15 | 2024-05-31 11:55 | 0:40 | COMPLETED | SEMI-CRITICAL
1729 | 2024-05-31 | May 2024 | 14:05 | 14:10 | 0.0034722 | VENTILATOR, PEDIATRIC | MINDRAY MEDICAL INDIA | SV 300 | THANISH | EMERGENCY | TIDAL VOLUME ISSUE | CLEANED THE EXPIRATORY VALVE AND PERFORMED THE FLOW SENSOR CALIBRATION.WORKING FINE. | NO | NO | 2024-05-31 | May 2024 | 14:30 | 2024-05-31 14:05 | 2024-05-31 14:30 | 0:25 | COMPLETED | CRITICAL
1730 | 2024-05-31 | May 2024 | 16:10 | 16:15 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | THANISH | ICU | DISPLAY FLICKERING ISSUE | DISPLAY ISSUE RECTIFIED | NO | NO | 2024-05-31 | May 2024 | 16:50 | 2024-05-31 16:10 | 2024-05-31 16:50 | 0:40 | COMPLETED | SEMI-CRITICAL
1731 | 2024-05-31 | May 2024 | 10:00 | 10:10 | 0.0069444 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | ICU | AIR BUBBLE DETECTOR BROKEN | REPLACED THE AIR BUBBLE DETECTOR AND RECTIFIED THE ISSUE | AIR BUBBLE DETECTOR | NO | 2024-05-31 | May 2024 | 10:35 | 2024-05-31 10:00 | 2024-05-31 10:35 | 0:35 | COMPLETED | NON-CRITICAL
1732 | 2024-05-31 | May 2024 | 10:20 | 10:25 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | ICU | OCCLUSION SENSOR NOT WORKING | CHANGED THE PRESSURE SENSOR AND RECTIFIED THE ISSUE | PRESSURE SENSOR | NO | 2024-05-31 | May 2024 | 10:40 | 2024-05-31 10:20 | 2024-05-31 10:40 | 0:20 | COMPLETED | NON-CRITICAL
1733 | 2024-05-31 | May 2024 | 11:00 | 11:05 | 0.0034722 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | JERLIN | ICU | SYRINGE HOLDER BROKEN | CHANGED THE SYRINGE HOLDER AND RECTIFIED THE ISSUE | HOLDER CLAMP | NO | 2024-05-31 | May 2024 | 11:30 | 2024-05-31 11:00 | 2024-05-31 11:30 | 0:30 | COMPLETED | NON-CRITICAL
1734 | 2024-05-31 | May 2024 | 11:30 | 11:35 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | ICU | DOOR BROKEN | CHANGED THE DOOR ASSEMBLY AND RECTIFIED THE ISSUE | DOOR ASSEMBLY | NO | 2024-05-31 | May 2024 | 11:55 | 2024-05-31 11:30 | 2024-05-31 11:55 | 0:25 | COMPLETED | NON-CRITICAL
1735 | 2024-05-31 | May 2024 | 13:00 | 13:10 | 0.0069444 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | ICU | AIR BUBBLE DETECTOR BROKEN | REPLACED THE AIR BUBBLE DETECTOR AND RECTIFIED THE ISSUE | AIR BUBBLE DETECTOR | NO | 2024-05-31 | May 2024 | 13:23 | 2024-05-31 13:00 | 2024-05-31 13:23 | 0:23 | COMPLETED | NON-CRITICAL
1736 | 2024-05-31 | May 2024 | 13:30 | 13:35 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | WARD | AIR BUBBLE DETECTOR BROKEN | REPLACED THE AIR BUBBLE DETECTOR AND RECTIFIED THE ISSUE | AIR BUBBLE DETECTOR | NO | 2024-05-31 | May 2024 | 14:03 | 2024-05-31 13:30 | 2024-05-31 14:03 | 0:33 | COMPLETED | NON-CRITICAL
1737 | 2024-06-01 | Jun 2024 | 3:00 | 3:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | SOORAJ | ICU | BATTERY DEPLETED ALARM | FOUND BATTERY TO BE DEFECTIVE. NEEDS REPLACEMENT. AS OF NOW WORKING ON MAINS | NO | BATTERY | 2024-06-01 | June 2024 | 3:10 | 2024-06-01 03:00 | 2024-06-01 03:10 | 0:10 | PENDING | SEMI-CRITICAL
1738 | 2024-06-01 | Jun 2024 | 21:30 | 21:35 | 0.0034722 | BIPAP MACHINE | PHILIPS | A40 | SOORAJ | WARD | CIRCUIT LEAK ALARM | FOUND PATIENT CIRCUIT TO BE LOOSE. TIGHTENED PROPERLY, ISSUE RECTIFIED. | NO | NO | 2024-06-01 | June 2024 | 21:45 | 2024-06-01 21:30 | 2024-06-01 21:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1739 | 2024-06-01 | Jun 2024 | 15:30 | 15:40 | 0.0069444 | COBLATOR | ARTHO CARE CORPORATION | COBLATOR II | JERLIN | OT | NOT ABLE TO CONNECT TO THE IRRIGATION SYSTEM | FOUND ISSUE WITH THE CONNECTING CABLE. FIXED IT AND RECTIFIED THE ISSUE | NO | NO | 2024-06-01 | June 2024 | 15:50 | 2024-06-01 15:30 | 2024-06-01 15:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1740 | 2024-06-02 | Jun 2024 | 2:30 | 2:35 | 0.0034722 | PLASMA STERILIZER | JOHNSON&JOHNSON | STERRAD 100NX | SOORAJ | CSSD | FOUND ISSUE WITH ONE OF THE BATTERIES | CONNECTED ON UPS AND ISSUE RECTIFIED TEMPORARILY | NO | NO | 2024-06-02 | June 2024 | 3:00 | 2024-06-02 02:30 | 2024-06-02 03:00 | 0:30 | COMPLETED | SEMI-CRITICAL
1741 | 2024-06-02 | Jun 2024 | 15:00 | 15:15 | 0.010417 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | JERLIN | ICU | NOT WORKING | FIXED THE WARMER HOSE AND RECTIFIED THE ISSUE | NO | NO | 2024-06-02 | June 2024 | 15:27 | 2024-06-02 15:00 | 2024-06-02 15:27 | 0:27 | COMPLETED | SEMI-CRITICAL
1742 | 2024-06-02 | Jun 2024 | 16:00 | 16:10 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 450 | JERLIN | ICU | ETCO2 SENSOR NOT WORKING | CLEANED THE ETCO2 SENSOR AND RECTIFIED THE ISSUE | NO | NO | 2024-06-02 | June 2024 | 16:30 | 2024-06-02 16:00 | 2024-06-02 16:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1743 | 2024-06-03 | Jun 2024 | 9:40 | 9:50 | 0.0069444 | LAMINAR AIR FLOW | MICROFLOW DEVICES INDIA PVT LTD | MFD V 1200 | RAMKIE | LABORATORY | DOOR JAM | CHECKED AND FOUND ROPE SLEEVE PEELED OFF, ADJUSTED AMD TEMP RECTIFIED | NO | NO | 2024-06-03 | June 2024 | 10:10 | 2024-06-03 09:40 | 2024-06-03 10:10 | 0:30 | COMPLETED | NON-CRITICAL
1744 | 2024-06-03 | Jun 2024 | 10:00 | 10:10 | 0.0069444 | BMI MACHINE | WELCH ALLYN | TANITA | JERLIN | OPD | NOT WORKING | CHANGED THE POWER ADAPTOR AND RECTIFIED THE ISSUE | POWER ADAPTOR | NO | 2024-06-03 | June 2024 | 10:25 | 2024-06-03 10:00 | 2024-06-03 10:25 | 0:25 | COMPLETED | NON-CRITICAL
1745 | 2024-06-03 | Jun 2024 | 15:00 | 15:15 | 0.010417 | LABOUR BED | ARJO HUNTLEIGH | HCB-04LY/BXA | JERLIN | WARD | NOT ABLE TO ADJUST THE HEIGHT | FIXED THE REMOTE PROPERLY AND RECTIFIED THE ISSUE | NO | NO | 2024-06-03 | June 2024 | 15:35 | 2024-06-03 15:00 | 2024-06-03 15:35 | 0:35 | COMPLETED | NON-CRITICAL
1746 | 2024-06-03 | Jun 2024 | 7:10 | 7:25 | 0.010417 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | CLEANED THE ABSORBER CANISTER AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-06-03 | June 2024 | 7:35 | 2024-06-03 07:10 | 2024-06-03 07:35 | 0:25 | COMPLETED | CRITICAL
1747 | 2024-06-03 | Jun 2024 | 8:45 | 8:50 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | HAFIZ | OT | NOT WORKING | CHECKED AND FOUND THERMISTOR RECEPTABLE DETACHED. REFIXED IT. NOW WORKING GOOD. | NO | NO | 2024-06-03 | June 2024 | 9:00 | 2024-06-03 08:45 | 2024-06-03 09:00 | 0:15 | COMPLETED | SEMI-CRITICAL
1748 | 2024-06-03 | Jun 2024 | 9:00 | 9:02 | 0.0013889 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | HAFIZ | OT | NOT WORKING | CHECKED AND FOUND THERMISTOR RECEPTABLE DETACHED. REFIXED IT. NOW WORKING GOOD. | NO | NO | 2024-06-03 | June 2024 | 9:20 | 2024-06-03 09:00 | 2024-06-03 09:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1749 | 2024-06-03 | Jun 2024 | 9:20 | 9:22 | 0.0013889 | LITHOCLAST | SURGIMEDIK HEALTHCARE | DIGI NOVA | HAFIZ | OT | ETCO2 NOT WORKING | REFIXED THE ETCO2 CABLE AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-06-03 | June 2024 | 9:30 | 2024-06-03 09:20 | 2024-06-03 09:30 | 0:10 | COMPLETED | SEMI-CRITICAL
1750 | 2024-06-03 | Jun 2024 | 15:00 | 15:10 | 0.0069444 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | HAFIZ | OT | NOT WORKING | CHECKED AND FOUND THERMISTOR RECEPTABLE DETACHED. REFIXED IT. NOW WORKING GOOD. | NO | NO | 2024-06-03 | June 2024 | 15:20 | 2024-06-03 15:00 | 2024-06-03 15:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1751 | 2024-06-03 | Jun 2024 | 8:05 | 8:10 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | THANISH | ICU | REMOTE NOT WORKING | RESETTED THE COT REMOTE.WORKING FINE. | NO | NO | 2024-06-03 | June 2024 | 8:15 | 2024-06-03 08:05 | 2024-06-03 08:15 | 0:10 | COMPLETED | NON-CRITICAL
1752 | 2024-06-03 | Jun 2024 | 9:30 | 9:35 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | RAMKIE | EMERGENCY | ECG EQUIPMENT MALFUNCTION | ISSUE WITH ECG LEAD SET. REPLACED LEAD SET AND ISSUE RECTIFIED | ECG CABLE | NO | 2024-06-03 | June 2024 | 9:50 | 2024-06-03 09:30 | 2024-06-03 09:50 | 0:20 | COMPLETED | CRITICAL
1753 | 2024-06-03 | Jun 2024 | 13:20 | 13:25 | 0.0034722 | DEFIBRILLATOR MONITOR WITH AED | PHILIPS | EFFICIA DFM 100 | RAMKIE | EMERGENCY | CANNOT ANOLYSE ECG | OP CHECK DONE | NO | NO | 2024-06-03 | June 2024 | 13:40 | 2024-06-03 13:20 | 2024-06-03 13:40 | 0:20 | COMPLETED | CRITICAL
1754 | 2024-06-04 | Jun 2024 | 10:10 | 10:20 | 0.0069444 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | RAMKIE | OPD | LEAD 3 NOT AVAILABLE | CHECKED AND ELETRRODE BULB THRED ISSUED, LUBRICATED AND RECTIFIED | NO | NO | 2024-06-04 | June 2024 | 10:40 | 2024-06-04 10:10 | 2024-06-04 10:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1755 | 2024-06-04 | Jun 2024 | 7:30 | 7:35 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SANTHOSH | OT | CIRCUIT LEAK TEST FAILED | FULL TEST DONE ISSUE RECTIFIED | NO | NO | 2024-06-04 | June 2024 | 7:40 | 2024-06-04 07:30 | 2024-06-04 07:40 | 0:10 | COMPLETED | CRITICAL
1756 | 2024-06-04 | Jun 2024 | 7:40 | 7:45 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SANTHOSH | OT | CIRCUIT LEAK TEST FAILED | ISSUE RECTIFIED | NO | NO | 2024-06-04 | June 2024 | 7:50 | 2024-06-04 07:40 | 2024-06-04 07:50 | 0:10 | COMPLETED | CRITICAL
1757 | 2024-06-05 | Jun 2024 | 11:10 | 11:20 | 0.0069444 | CRYO CUFF | CHATTANOOGA | C2 | RAMKIE | PHYSIOTHERAPY | LEAK DURING THERAPY | CHECKED AND FOUND THE WASHER DAMAGE, SO REPLACED AND RECTIFIED | WASHER | NO | 2024-06-05 | June 2024 | 11:40 | 2024-06-05 11:10 | 2024-06-05 11:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1758 | 2024-06-05 | Jun 2024 | 14:30 | 14:35 | 0.0034722 | NEBULIZER | JK MEDICALS | ULTIMA | JERLIN | WARD | TUBING | FOUND THE CONNECTR WAS BROKEN . FIXED IT AND RECTIFIED THE ISSUE | NO | NO | 2024-06-05 | June 2024 | 14:55 | 2024-06-05 14:30 | 2024-06-05 14:55 | 0:25 | COMPLETED | SEMI-CRITICAL
1759 | 2024-06-05 | Jun 2024 | 13:45 | 13:55 | 0.0069444 | HUMIDIFIER AIR | FISHER AND PAYKEL | MR 810 | JERLIN | ICU | NOT SWITCHING ON | FOUND ISSUE IN THE POWER SOCKET , FIXED IT AND RECTIFIED THE ISSUE | NO | NO | 2024-06-05 | June 2024 | 14:07 | 2024-06-05 13:45 | 2024-06-05 14:07 | 0:22 | COMPLETED | SEMI-CRITICAL
1760 | 2024-06-05 | Jun 2024 | 10:10 | 10:20 | 0.0069444 | PRESSURE INJECTOR | LF OPTI VONTAGE | ILLUMENO NEO | HAFIZ | CATHLAB | NOT SWITCHING ON | CHECKED AND ISSUE WITH POWERSOCKET. FIXED IT AND RESTARTED THE MACHINE. NOW READY TO USE. | NO | NO | 2024-06-05 | June 2024 | 10:30 | 2024-06-05 10:10 | 2024-06-05 10:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1761 | 2024-06-05 | Jun 2024 | 12:00 | 12:04 | 0.0027778 | USG MACHINE | MINDRAY MEDICAL INDIA | MX7 | HAFIZ | OT | PROBE NOT WORKING | PROPERLY FIXED IT AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-06-05 | June 2024 | 12:10 | 2024-06-05 12:00 | 2024-06-05 12:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1762 | 2024-06-05 | Jun 2024 | 12:20 | 12:22 | 0.0013889 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | OT | NIBP AND ECG NOT WORKING | BP CUFF AND ECG CABLE PASTED. NOW WORKING GOOD. | NO | NO | 2024-06-05 | June 2024 | 12:40 | 2024-06-05 12:20 | 2024-06-05 12:40 | 0:20 | COMPLETED | SEMI-CRITICAL
1763 | 2024-06-05 | Jun 2024 | 10:00 | 10:05 | 0.0034722 | MORECELLATOR | KARL STORZ | UNIDRIVE S III | SANTHOSH | OT | UNWANTED SOUND APPEARS | LUBED AND ISSUE RECTIFIED | NO | NO | 2024-06-05 | June 2024 | 10:10 | 2024-06-05 10:00 | 2024-06-05 10:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1764 | 2024-06-05 | Jun 2024 | 15:40 | 15:42 | 0.0013889 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | JERLIN | OT | CIRCUIT LEAK TEST FAILED | FIXED THE CANISTER PROPERLY AND NOW IT'S WORKING GOOD | NO | NO | 2024-06-05 | June 2024 | 16:00 | 2024-06-05 15:40 | 2024-06-05 16:00 | 0:20 | COMPLETED | CRITICAL
1765 | 2024-06-06 | Jun 2024 | 16:30 | 16:40 | 0.0069444 | BLOOD GAS ANALYSER | GEM 3500 | INSTRUMENTATION LABORATORY | JERLIN | ICU | CARTRIDGE EXPIRED | REPLACED THE CARTRIDGE AND RECTIFIED THE ISSUE | CARTRIDGE | NO | 2024-06-06 | June 2024 | 17:00 | 2024-06-06 16:30 | 2024-06-06 17:00 | 0:30 | COMPLETED | SEMI-CRITICAL
1766 | 2024-06-06 | Jun 2024 | 9:10 | 9:15 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | THANISH | OT | TIDAL VOLUME ABNORMAL ISSUE | CLEANED THE FLOW SENSOR AND PERFORMED THE FLOW SENSOR CALIBRATION.WORKING FINE. | NO | NO | 2024-06-06 | June 2024 | 9:30 | 2024-06-06 09:10 | 2024-06-06 09:30 | 0:20 | COMPLETED | CRITICAL
1767 | 2024-06-07 | Jun 2024 | 14:10 | 14:30 | 0.013889 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | RAMKIE | OT | OCCLUSION ERROR DURING PURGE | CHECKED AND CHANGED SETTING | NO | NO | 2024-06-07 | June 2024 | 14:55 | 2024-06-07 14:10 | 2024-06-07 14:55 | 0:45 | COMPLETED | NON-CRITICAL
1768 | 2024-06-07 | Jun 2024 | 9:35 | 9:40 | 0.0034722 | PLASMA STERILIZER | JOHNSON&JOHNSON | STERRAD 100NX | THANISH | CSSD | H2O2 CYCLE ISSUE | RE-STARTED THE MACHINE.WORKING FINE. | NO | NO | 2024-06-07 | June 2024 | 9:55 | 2024-06-07 09:35 | 2024-06-07 09:55 | 0:20 | COMPLETED | SEMI-CRITICAL
1769 | 2024-06-08 | Jun 2024 | 15:30 | 15:40 | 0.0069444 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | RAMKIE | OT | NOT WORKING | CHECKED AND SESNOR CABLE DAMAGE, RECTIFIED AFTER SOLDERING | NO | NO | 2024-06-08 | June 2024 | 15:55 | 2024-06-08 15:30 | 2024-06-08 15:55 | 0:25 | COMPLETED | SEMI-CRITICAL
1770 | 2024-06-08 | Jun 2024 | 18:35 | 18:40 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | NEED TO REPLACE A CARTRIDGE | REPLACED A NEW CARTRIDGE AND WARMING UP DONE. NOW READY TO USE. | CARTRIDGE | NO | 2024-06-08 | June 2024 | 18:50 | 2024-06-08 18:35 | 2024-06-08 18:50 | 0:15 | COMPLETED | SEMI-CRITICAL
1771 | 2024-06-09 | Jun 2024 | 10:10 | 10:15 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | THANISH | ICU | COT BRAKE PEDAL NOT WORKING | ISSUE RECTIFIED | NO | NO | 2024-06-09 | June 2024 | 10:40 | 2024-06-09 10:10 | 2024-06-09 10:40 | 0:30 | COMPLETED | NON-CRITICAL
1772 | 2024-06-09 | Jun 2024 | 15:30 | 15:40 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | THANISH | OT | VOLUME CONTROL MODE ISSUE | CLEANED THE FLOW SENSOR AND SODA LIME JAR&INSPIRATORY,EXPIRATORY VALVES.WORKING FINE. | NO | NO | 2024-06-09 | June 2024 | 15:50 | 2024-06-09 15:30 | 2024-06-09 15:50 | 0:20 | COMPLETED | CRITICAL
1773 | 2024-06-10 | Jun 2024 | 10:20 | 10:30 | 0.0069444 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | RAMKIE | OT | DOOR BROKEN | REPLACED FROM ANOTHER MACHINE | DOOR ASSEMBLY | NO | 2024-06-10 | June 2024 | 10:55 | 2024-06-10 10:20 | 2024-06-10 10:55 | 0:35 | COMPLETED | NON-CRITICAL
1774 | 2024-06-10 | Jun 2024 | 8:15 | 8:20 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | CANISTER CHANGED AND OBSERVED. ISSUE RECTIFIED | CANISTER | NO | 2024-06-10 | June 2024 | 8:45 | 2024-06-10 08:15 | 2024-06-10 08:45 | 0:30 | COMPLETED | CRITICAL
1775 | 2024-06-10 | Jun 2024 | 9:00 | 9:00 | 0 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | CLEANED CANISTER, SODALIME QUANTITY REDUCED. WORKING GOOD. | NO | NO | 2024-06-10 | June 2024 | 9:30 | 2024-06-10 09:00 | 2024-06-10 09:30 | 0:30 | COMPLETED | CRITICAL
1776 | 2024-06-11 | Jun 2024 | 10:40 | 10:50 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | OT | NOT WORKING | CHECKED AND POWER CORD CONNECTION RECTIFIED | NO | NO | 2024-06-11 | June 2024 | 11:40 | 2024-06-11 10:40 | 2024-06-11 11:40 | 1:00 | COMPLETED | NON-CRITICAL
1777 | 2024-06-12 | Jun 2024 | 11:40 | 11:45 | 0.0034722 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | ePM 12c | SOORAJ | WARD | SPO2 VALUE NOT SHOWING. | FOUND SPO2 SESOR PIN BENT. REMOVED BENT AND ISSUE RECTIFIED. | NO | NO | 2024-06-12 | June 2024 | 12:30 | 2024-06-12 11:40 | 2024-06-12 12:30 | 0:50 | COMPLETED | SEMI-CRITICAL
1778 | 2024-06-12 | Jun 2024 | 8:50 | 8:55 | 0.0034722 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX | JERLIN | OT | NOT ABLE TO ADJUST BIPOLAR INTENSITY | FOUND ISSUE WITH THE SWITCH CONTROL BOARD. FIXED AND RECTIFIED THE ISSUE | NO | NO | 2024-06-12 | June 2024 | 9:15 | 2024-06-12 08:50 | 2024-06-12 09:15 | 0:25 | COMPLETED | SEMI-CRITICAL
1779 | 2024-06-12 | Jun 2024 | 15:00 | 15:10 | 0.0069444 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | JERLIN | WARD | NOT SWITCHING ON | FIXED AND RECTIFIED THE ISSUE | NO | NO | 2024-06-12 | June 2024 | 16:00 | 2024-06-12 15:00 | 2024-06-12 16:00 | 1:00 | COMPLETED | SEMI-CRITICAL
1780 | 2024-06-12 | Jun 2024 | 6:45 | 6:50 | 0.0034722 | PLASMA STERILIZER | JOHNSON&JOHNSON | STERRAD 100NX | HAFIZ | CSSD | INSTRUMENT STRUCK INSIDE | REMOVED THE FRONT DOOR AND INSTRUMENT RETRIEVED. NOW WORKING GOOD. | NO | NO | 2024-06-12 | June 2024 | 7:00 | 2024-06-12 06:45 | 2024-06-12 07:00 | 0:15 | COMPLETED | SEMI-CRITICAL
1781 | 2024-06-12 | Jun 2024 | 7:00 | 7:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | CLEANED THE ABSORBER CANISTER AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-06-12 | June 2024 | 7:20 | 2024-06-12 07:00 | 2024-06-12 07:20 | 0:20 | COMPLETED | CRITICAL
1782 | 2024-06-12 | Jun 2024 | 15:20 | 15:25 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SANTHOSH | WARD | PRINTER ISSUE | ISSUE RECTIFIED | NO | NO | 2024-06-12 | June 2024 | 15:30 | 2024-06-12 15:20 | 2024-06-12 15:30 | 0:10 | COMPLETED | CRITICAL
1783 | 2024-06-12 | Jun 2024 | 10:00 | 10:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SANTHOSH | DIALYSIS | ECG FAILIURE | OP CHECK DONE | NO | NO | 2024-06-12 | June 2024 | 10:05 | 2024-06-12 10:00 | 2024-06-12 10:05 | 0:05 | COMPLETED | CRITICAL
1784 | 2024-06-12 | Jun 2024 | 11:00 | 11:10 | 0.0069444 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SANTHOSH | ICU | THERAPHY KNOB FAILURE | OP CHECK DONE | NO | NO | 2024-06-12 | June 2024 | 11:05 | 2024-06-12 11:00 | 2024-06-12 11:05 | 0:05 | COMPLETED | CRITICAL
1785 | 2024-06-13 | Jun 2024 | 12:20 | 12:30 | 0.0069444 | DIALYSIS MACHINE | FRESENIUS | 4008 S NG | RAMKIE | DIALYSIS | NOT SWITCHING ON | CHECKED AND RECONNECTED CONNECTED AND MACHINE WORKING FONE | NO | NO | 2024-06-13 | June 2024 | 13:00 | 2024-06-13 12:20 | 2024-06-13 13:00 | 0:40 | COMPLETED | CRITICAL
1786 | 2024-06-13 | Jun 2024 | 12:00 | 12:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | JERLIN | ENDOSCOPY | WATER TRAP MALFUNCTION | REPLACED THE WATER TRAP AND RECTIFIED THE ISSUE | WATER TRAP | NO | 2024-06-13 | June 2024 | 12:25 | 2024-06-13 12:00 | 2024-06-13 12:25 | 0:25 | COMPLETED | SEMI-CRITICAL
1787 | 2024-06-13 | Jun 2024 | 14:00 | 14:13 | 0.0090278 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | WARD | SIDE RAIL NOT WORKING | FIXED AND RECTIFIED THE ISSUE | NO | NO | 2024-06-13 | June 2024 | 14:23 | 2024-06-13 14:00 | 2024-06-13 14:23 | 0:23 | COMPLETED | NON-CRITICAL
1788 | 2024-06-13 | Jun 2024 | 7:45 | 8:00 | 0.010417 | TRANSCRANIAL DOPPLER | VIASONIC | DOLPHIN 4D | JERLIN | ICU | REMOTE SENSOR NOT WORKING | FIXED IT AND RECTIFIED THE ISSUE | NO | NO | 2024-06-13 | June 2024 | 8:25 | 2024-06-13 07:45 | 2024-06-13 08:25 | 0:40 | COMPLETED | SEMI-CRITICAL
1789 | 2024-06-13 | Jun 2024 | 22:00 | 22:04 | 0.0027778 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | WARD | KEY STRUCK ALARM | REMOVED, CLEANED AND RECONNECTED THE CONNECTOR CABLE. NOW WORKING GOOD. | NO | NO | 2024-06-13 | June 2024 | 22:30 | 2024-06-13 22:00 | 2024-06-13 22:30 | 0:30 | COMPLETED | NON-CRITICAL
1790 | 2024-06-13 | Jun 2024 | 21:35 | 21:40 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | HAFIZ | ICU | NEED TO REPLACE A CARTRIDGE | REPLACED A NEW CARTRIDGE AND WARMING UP DONE. NOW READY TO USE. | CARTRIDGE | NO | 2024-06-13 | June 2024 | 22:10 | 2024-06-13 21:35 | 2024-06-13 22:10 | 0:35 | COMPLETED | SEMI-CRITICAL
1791 | 2024-06-13 | Jun 2024 | 14:20 | 14:25 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SANTHOSH | CATHLAB | NO BATTERY BACKUP ISSUE | BATTERY REPLACED AND WORKING. | BATTERY | NO | 2024-06-13 | June 2024 | 15:30 | 2024-06-13 14:20 | 2024-06-13 15:30 | 1:10 | COMPLETED | CRITICAL
1792 | 2024-06-13 | Jun 2024 | 15:00 | 15:15 | 0.010417 | EXAMINATION COUCH | MIDMARK JANOK | E122012 | THANISH | OPD | HEADREST BROKEN | REPLACED WITH ANOTHER HEAD REST.WORKING FINE. | HEADREST | NO | 2024-06-13 | June 2024 | 15:27 | 2024-06-13 15:00 | 2024-06-13 15:27 | 0:27 | COMPLETED | NON-CRITICAL
1793 | 2024-06-14 | Jun 2024 | 19:30 | 19:35 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | SOORAJ | ICU | PATIENT DATA NOT SHOWING ON LIST | FOUND SPOOLER TO BE FULL. CLEARED SPOOLER AND ISSUE RECTIFIED. | NO | NO | 2024-06-14 | June 2024 | 20:10 | 2024-06-14 19:30 | 2024-06-14 20:10 | 0:40 | COMPLETED | SEMI-CRITICAL
1794 | 2024-06-14 | Jun 2024 | 13:20 | 13:30 | 0.0069444 | BMI MACHINE | WELCH ALLYN | TANITA | RAMKIE | WARD | HEIGHT LENGTH NOT IN POSITION | CHECKED AND FIXED THE SCREW AND RECTIFIED | NO | NO | 2024-06-14 | June 2024 | 13:50 | 2024-06-14 13:20 | 2024-06-14 13:50 | 0:30 | COMPLETED | NON-CRITICAL
1795 | 2024-06-14 | Jun 2024 | 9:10 | 9:15 | 0.0034722 | IABP MACHINE | MAQUET | CS300 | JERLIN | OT | HELIUM GAS LEAK ISSUE | FIXED A WASHER AND RECTIFIED THE ISSUE | WASHER | NO | 2024-06-14 | June 2024 | 9:30 | 2024-06-14 09:10 | 2024-06-14 09:30 | 0:20 | COMPLETED | CRITICAL
1796 | 2024-06-14 | Jun 2024 | 3:00 | 3:03 | 0.0020833 | PLASMA STERILIZER | JOHNSON&JOHNSON | STERRAD 100NX | HAFIZ | CSSD | H2O2 ADJUSTMENT TAKING LONGER TIME AND UV PATH BLOCKED ALARM. | CLEANED THE UV WINDOW SHIELD AND RESTARTED THE AMCHINE. NOW READY TO USE. | NO | NO | 2024-06-14 | June 2024 | 4:30 | 2024-06-14 03:00 | 2024-06-14 04:30 | 1:30 | COMPLETED | SEMI-CRITICAL
1797 | 2024-06-14 | Jun 2024 | 22:00 | 22:03 | 0.0020833 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | ePM 12c | HAFIZ | WARD | TOUCH NOT WORKING | REMOVED, CLEANED AND REFIXED THE DISPLAY CABLE. NOW WORKING GOOD. | NO | NO | 2024-06-14 | June 2024 | 22:30 | 2024-06-14 22:00 | 2024-06-14 22:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1798 | 2024-06-14 | Jun 2024 | 17:35 | 17:40 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | SANTHOSH | ICU | STORAGE FULL | DELETED OLD DATA | NO | NO | 2024-06-14 | June 2024 | 17:50 | 2024-06-14 17:35 | 2024-06-14 17:50 | 0:15 | COMPLETED | SEMI-CRITICAL
1799 | 2024-06-15 | Jun 2024 | 11:20 | 11:30 | 0.0069444 | FREEZER,DEEP FREEZING -40 | THERMO FISHER SCIENTIFIC | FORMA 7000 SERIES | RAMKIE | BLOODBANK | TEMP NOT REACHING | CHECKED AND GAS REFILLING DONE BY VENDOR | NO | NO | 2024-06-15 | June 2024 | 11:50 | 2024-06-15 11:20 | 2024-06-15 11:50 | 0:30 | COMPLETED | SEMI-CRITICAL
1800 | 2024-06-15 | Jun 2024 | 10:05 | 10:15 | 0.0069444 | EXAMINATION COUCH | MIDMARK JANOK | E122012 | THANISH | OPD | NOT WORKING | REFIXED THE CABLE CONNECTIONS IN CONTROL BOX.WORKING FINE. | NO | NO | 2024-06-15 | June 2024 | 10:40 | 2024-06-15 10:05 | 2024-06-15 10:40 | 0:35 | COMPLETED | NON-CRITICAL
1801 | 2024-06-17 | Jun 2024 | 20:00 | 20:15 | 0.010417 | DIGITAL PLATFORM SCALE | SECA | 676 | JERLIN | DIALYSIS | WEIGHING SCALE NOT WORKING | FIXED THE POWER ADAPTOR AND RECTIFIED THE ISSUE | POWER ADAPTOR | NO | 2024-06-17 | June 2024 | 20:35 | 2024-06-17 20:00 | 2024-06-17 20:35 | 0:35 | COMPLETED | NON-CRITICAL
1802 | 2024-06-17 | Jun 2024 | 22:30 | 22:45 | 0.010417 | FETAL HEART MONITOR | WIPRO GE HEALTH CARE | COROMETRICS 259 | JERLIN | WARD | FETAL MOVEMENT MARKER NOT WORKING | FIXED THE PROBE AND RECTIFIED THE ISSUE | NO | NO | 2024-06-17 | June 2024 | 23:00 | 2024-06-17 22:30 | 2024-06-17 23:00 | 0:30 | COMPLETED | SEMI-CRITICAL
1803 | 2024-06-17 | Jun 2024 | 3:00 | 3:05 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | WARD | NOT WORKING | FOUND THAT THE REMOTE WAS LOCKED . UNLOCKED THE REMOTE AND RECTIFIED THE ISSUE. | NO | NO | 2024-06-17 | June 2024 | 3:20 | 2024-06-17 03:00 | 2024-06-17 03:20 | 0:20 | COMPLETED | NON-CRITICAL
1804 | 2024-06-17 | Jun 2024 | 10:50 | 10:55 | 0.0034722 | BMI MACHINE | AKAS MEDICAL | QVS-100 | THANISH | OPD | NIBP NOT WORKING | REPLACED THE NIBP HOSE.WORKING FINE. | NIBP HOSE | NO | 2024-06-17 | June 2024 | 11:20 | 2024-06-17 10:50 | 2024-06-17 11:20 | 0:30 | COMPLETED | NON-CRITICAL
1805 | 2024-06-18 | Jun 2024 | 12:20 | 12:30 | 0.0069444 | HOLTER,RECORDER | WIPRO GE HEALTH CARE | SEER 1000 | RAMKIE | OPD | LEAD 2 NOT AVAILABLE | CHECKED AND FOUND THE PROBLEM WITH HOLTER CABLE, REPLACED NEW ONE AND RECTIFIED | LEAD CABLE | NO | 2024-06-18 | June 2024 | 13:20 | 2024-06-18 12:20 | 2024-06-18 13:20 | 1:00 | COMPLETED | SEMI-CRITICAL
1806 | 2024-06-18 | Jun 2024 | 15:10 | 15:12 | 0.0013889 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | WARD | HOLDER CLAMP BROKEN | REPLACED HOLDER CLAMP FROM DEFECTIVE PUMP. NOW ISSUE RECTIFIED. | HOLDER CLAMP | NO | 2024-06-18 | June 2024 | 15:40 | 2024-06-18 15:10 | 2024-06-18 15:40 | 0:30 | COMPLETED | NON-CRITICAL
1807 | 2024-06-18 | Jun 2024 | 18:50 | 18:52 | 0.0013889 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | HAFIZ | WARD | NOT WORKING | CHECKED AND FOUND THERMISTOR RECEPTABLE DETACHED. REFIXED IT. NOW WORKING GOOD. | NO | NO | 2024-06-18 | June 2024 | 19:00 | 2024-06-18 18:50 | 2024-06-18 19:00 | 0:10 | COMPLETED | SEMI-CRITICAL
1808 | 2024-06-18 | Jun 2024 | 16:30 | 16:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SANTHOSH | WARD | CONTROL BOARD ISSUE | REPLACED CONTROLLED BOARD | CONTROL BOARD | NO | 2024-06-18 | June 2024 | 16:45 | 2024-06-18 16:30 | 2024-06-18 16:45 | 0:15 | COMPLETED | NON-CRITICAL
1809 | 2024-06-18 | Jun 2024 | 10:45 | 10:50 | 0.0034722 | STEAM STERILIZER | GETINGE | GSS67H102E | SANTHOSH | CSSD | VACCUM PRESSURE ISSUE | PILOT VALVE PROBLEM | NO | NO | 2024-06-18 | June 2024 | 15:50 | 2024-06-18 10:45 | 2024-06-18 15:50 | 5:05 | COMPLETED | SEMI-CRITICAL
1810 | 2024-06-19 | Jun 2024 | 10:10 | 10:30 | 0.013889 | HOLTER,RECORDER | WIPRO GE HEALTH CARE | SEER 1000 | RAMKIE | OPD | NO RECORDED ECG | CHECKED AND FOUND THE PROBLEM WITH HOLTER CABLE, REPLACED NEW ONE AND RECTIFIED | LEAD CABLE | NO | 2024-06-19 | June 2024 | 11:10 | 2024-06-19 10:10 | 2024-06-19 11:10 | 1:00 | COMPLETED | SEMI-CRITICAL
1811 | 2024-06-19 | Jun 2024 | 9:30 | 9:40 | 0.0069444 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | JERLIN | ICU | TRANSDUCER TEST FAILED IN PREUSE CHECK | RERAN PREUSE CHECK AND RECTIFIED THE ISSUE | NO | NO | 2024-06-19 | June 2024 | 9:55 | 2024-06-19 09:30 | 2024-06-19 09:55 | 0:25 | COMPLETED | CRITICAL
1812 | 2024-06-19 | Jun 2024 | 18:40 | 18:45 | 0.0034722 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | NEED TO REPLACE A CARTRIDGE | REPLACED A NEW CARTRIDGE AND WARMING UP DONE. NOW READY TO USE. | CARTRIDGE | NO | 2024-06-19 | June 2024 | 19:35 | 2024-06-19 18:40 | 2024-06-19 19:35 | 0:55 | COMPLETED | SEMI-CRITICAL
1813 | 2024-06-19 | Jun 2024 | 10:45 | 10:50 | 0.0034722 | FILM ARRAY | BIOMERIUX | BIOFIRE RR | SANTHOSH | LABORATORY | NOT WORKING | ISSUE RECTIFIED | NO | NO | 2024-06-19 | June 2024 | 12:00 | 2024-06-19 10:45 | 2024-06-19 12:00 | 1:15 | COMPLETED | SEMI-CRITICAL
1814 | 2024-06-20 | Jun 2024 | 9:10 | 9:15 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | SOORAJ | OPD | PATIENT DATA NOT ABLE TO PULL FROM PACS. | FOUND ETHERNET DISCONNECTED. RECONNECTED LAN CABLE, ISSUE RECTIFIED. | NO | NO | 2024-06-20 | June 2024 | 9:30 | 2024-06-20 09:10 | 2024-06-20 09:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1815 | 2024-06-20 | Jun 2024 | 15:30 | 15:35 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | L1CW | HAFIZ | OT | NOT WORKING | CLEANED THE AIR FILTER AND OBSERVED. NOW READY TO USE. | NO | NO | 2024-06-20 | June 2024 | 16:00 | 2024-06-20 15:30 | 2024-06-20 16:00 | 0:30 | COMPLETED | SEMI-CRITICAL
1816 | 2024-06-20 | Jun 2024 | 11:15 | 11:20 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | SANTHOSH | OT | ECG EARTHING ISSUE | ISSUE RECTIFIED | NO | NO | 2024-06-20 | June 2024 | 11:30 | 2024-06-20 11:15 | 2024-06-20 11:30 | 0:15 | COMPLETED | SEMI-CRITICAL
1817 | 2024-06-20 | Jun 2024 | 11:10 | 11:15 | 0.0034722 | BIPAP MACHINE | PHILIPS | A40 | THANISH | WARD | SETTINGS ISSUE | RECTIFIED THE SETTINGS ISSUE.WORKING FINE. | NO | NO | 2024-06-20 | June 2024 | 11:30 | 2024-06-20 11:10 | 2024-06-20 11:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1818 | 2024-06-21 | Jun 2024 | 11:10 | 11:30 | 0.013889 | CRRT MACHINE | BAXTER | PRISMA FLEX | RAMKIE | DIALYSIS | BLD ALARM | CHECKED CLEANED AND CALIBRATED, PROBLEM RECTIFIED | NO | NO | 2024-06-21 | June 2024 | 11:50 | 2024-06-21 11:10 | 2024-06-21 11:50 | 0:40 | COMPLETED | CRITICAL
1819 | 2024-06-22 | Jun 2024 | 12:10 | 12:30 | 0.013889 | DIALYSIS MACHINE | FRESENIUS | 4008 S NG | RAMKIE | DIALYSIS | V99 | CLEANED FILTER AND RECTIFIED | NO | NO | 2024-06-22 | June 2024 | 12:50 | 2024-06-22 12:10 | 2024-06-22 12:50 | 0:40 | COMPLETED | CRITICAL
1820 | 2024-06-22 | Jun 2024 | 10:30 | 10:35 | 0.0034722 | EXAMINATION COUCH | MIDMARK JANOK | E122012 | SANTHOSH | OPD | NOT WORKING | GAS SPRING AND REMOTE REPLACED | GAS SPRING AND REMOTE | NO | 2024-06-22 | June 2024 | 10:40 | 2024-06-22 10:30 | 2024-06-22 10:40 | 0:10 | COMPLETED | NON-CRITICAL
1821 | 2024-06-22 | Jun 2024 | 17:40 | 17:45 | 0.0034722 | BLOOD GAS ANALYSER | GEM 3500 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | CADRIDGE FINISHED | REPLACED CARTRIDGE | CARTRIDGE | NO | 2024-06-22 | June 2024 | 17:50 | 2024-06-22 17:40 | 2024-06-22 17:50 | 0:10 | COMPLETED | SEMI-CRITICAL
1822 | 2024-06-23 | Jun 2024 | 12:00 | 12:05 | 0.0034722 | DIALYSIS MACHINE | FRESENIUS | 4008 S | SANTHOSH | DIALYSIS | V99 ERROR | ISSUE RECTIFIED | NO | NO | 2024-06-23 | June 2024 | 12:10 | 2024-06-23 12:00 | 2024-06-23 12:10 | 0:10 | COMPLETED | CRITICAL
1823 | 2024-06-23 | Jun 2024 | 15:00 | 15:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | SANTHOSH | OT | FLASH CARD PROBLEM | REPLACED FROM ANOTHER MACHINE | NO | NO | 2024-06-23 | June 2024 | 15:20 | 2024-06-23 15:00 | 2024-06-23 15:20 | 0:20 | COMPLETED | CRITICAL
1824 | 2024-06-23 | Jun 2024 | 15:10 | 15:15 | 0.0034722 | VENTILATOR, PEDIATRIC | MINDRAY MEDICAL INDIA | SV 300 | SANTHOSH | ICU | O2 CELL FAIL | O2 CELL DEFECTIVE AND NEED REPLACEMENT | NO | O2 CELL | 2024-06-23 | June 2024 | 15:30 | 2024-06-23 15:10 | 2024-06-23 15:30 | 0:20 | PENDING | CRITICAL
1825 | 2024-06-23 | Jun 2024 | 10:05 | 10:10 | 0.0034722 | HEATER COOLER | GETINGE | HCU40 | THANISH | OT | HEMOTHERM BLANKET WATER LEAKAGE | LEAK ARRESTED IN HEMOTHERM BLANKET. | NO | NO | 2024-06-23 | June 2024 | 10:30 | 2024-06-23 10:05 | 2024-06-23 10:30 | 0:25 | COMPLETED | CRITICAL
1826 | 2024-06-23 | Jun 2024 | 9:10 | 9:10 | 0 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | THANISH | RADIOLOGY | MACHINE NOT SWITCHING ON | PLUG TOP BROKEN/REPLACED WITH NEW 16AMPS PLUG TOP.WORKING FINE. | PLUG TOP | NO | 2024-06-23 | June 2024 | 9:30 | 2024-06-23 09:10 | 2024-06-23 09:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1827 | 2024-06-24 | Jun 2024 | 7:30 | 7:35 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | SOORAJ | OT | SLAVE MONITOR NOT SHOWING | DVI CABLE LOOSE CONTACT. RECONNECTED, ISSUE RECTIFIED. | NO | NO | 2024-06-24 | June 2024 | 7:50 | 2024-06-24 07:30 | 2024-06-24 07:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1828 | 2024-06-24 | Jun 2024 | 8:15 | 8:20 | 0.0034722 | HEATER COOLER | GETINGE | HCU40 | SOORAJ | OT | LEAK IN BLANKET | LEAK ARRESTED AND ISSUE RECTIFIED. | NO | NO | 2024-06-24 | June 2024 | 9:00 | 2024-06-24 08:15 | 2024-06-24 09:00 | 0:45 | COMPLETED | CRITICAL
1829 | 2024-06-24 | Jun 2024 | 14:10 | 14:30 | 0.013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | RAMKIE | WARD | NOT WORKING | CHECKED AND RESET DONE AND RECTIFIED | NO | NO | 2024-06-24 | June 2024 | 14:50 | 2024-06-24 14:10 | 2024-06-24 14:50 | 0:40 | COMPLETED | NON-CRITICAL
1830 | 2024-06-25 | Jun 2024 | 9:30 | 9:35 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | SOORAJ | OT | WARMER NOT WORKING | FOUND THERMISTOR CABLE CUT. CABLE SOLDERED AND ISSUE RECTIFIED. | NO | NO | 2024-06-25 | June 2024 | 10:30 | 2024-06-25 09:30 | 2024-06-25 10:30 | 1:00 | COMPLETED | SEMI-CRITICAL
1831 | 2024-06-25 | Jun 2024 | 11:00 | 11:35 | 0.024306 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | SODA LIME CANNISTER FIXED PROPERLY AND LEAK ARRESTED. | NO | NO | 2024-06-25 | June 2024 | 12:00 | 2024-06-25 11:00 | 2024-06-25 12:00 | 1:00 | COMPLETED | CRITICAL
1832 | 2024-06-25 | Jun 2024 | 10:10 | 10:30 | 0.013889 | MICROSCOPE, LABORATORY,TRINOCULAR, | ZEISS | PRIMO STAR | RAMKIE | LABORATORY | LIGHT NOT SWITCHING ON | CHECKED AND REPLACED BULB AND SERVICED HOLDER | HALOGEN BULB | NO | 2024-06-25 | June 2024 | 14:50 | 2024-06-25 10:10 | 2024-06-25 14:50 | 4:40 | COMPLETED | SEMI-CRITICAL
1833 | 2024-06-25 | Jun 2024 | 15:00 | 15:10 | 0.0069444 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID E95 | JERLIN | OPD | NOT ABLE TO SENT TEST IN PACS | CONNECTED THELAN PORT AND RECTIFIED THE ISSUE | NO | NO | 2024-06-25 | June 2024 | 15:30 | 2024-06-25 15:00 | 2024-06-25 15:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1834 | 2024-06-25 | Jun 2024 | 12:15 | 12:20 | 0.0034722 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | SANTHOSH | OPD | SOFTWARE ISSUE | ISSUE RECTIFIED | NO | NO | 2024-06-25 | June 2024 | 12:25 | 2024-06-25 12:15 | 2024-06-25 12:25 | 0:10 | COMPLETED | SEMI-CRITICAL
1835 | 2024-06-25 | Jun 2024 | 9:35 | 9:40 | 0.0034722 | TREATMENT UNIT | SKANRAY DENTAL | SKANLECTICA OPAL | THANISH | OPD | MOBILE LIGHT ISSUE | CLEANED THE MOISTURE INSIDE THE MOBILE LIGHT.WORKING FINE. | NO | NO | 2024-06-25 | June 2024 | 10:10 | 2024-06-25 09:35 | 2024-06-25 10:10 | 0:35 | COMPLETED | SEMI-CRITICAL
1836 | 2024-06-25 | Jun 2024 | 10:20 | 10:25 | 0.0034722 | LIGHT OPERATING,TRIPLE DOME | MAQUET | ARDVCS209006A VCS64SF | THANISH | OT | OT LIGHT ALIGNMENT ISSUE | ADJUSTED THE OT LIGHT PENDANT ALIGNMENT.WORKING FINE. | NO | NO | 2024-06-25 | June 2024 | 10:45 | 2024-06-25 10:20 | 2024-06-25 10:45 | 0:25 | COMPLETED | NON-CRITICAL
1837 | 2024-06-26 | Jun 2024 | 11:10 | 11:30 | 0.013889 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | RAMKIE | RADIOLOGY | PACS CONFIGURATION ERROR | CHECKED AND CONFIGURED SETTING AND RECTIFIED | NO | NO | 2024-06-26 | June 2024 | 12:20 | 2024-06-26 11:10 | 2024-06-26 12:20 | 1:10 | COMPLETED | SEMI-CRITICAL
1838 | 2024-06-26 | Jun 2024 | 16:00 | 16:10 | 0.0069444 | EEG MACHINE | NOTUS NEUROLOGY | NICOLET EEG V32 | JERLIN | OPD | CAMERA NOT WORKING WHILE DOING THE PROCEDURE | FIXED THE CONNECTOR CABLE AND FOUND THAT IT'S WORKING GOOD | NO | NO | 2024-06-26 | June 2024 | 16:35 | 2024-06-26 16:00 | 2024-06-26 16:35 | 0:35 | COMPLETED | SEMI-CRITICAL
1839 | 2024-06-26 | Jun 2024 | 18:25 | 18:30 | 0.0034722 | VIDEO ELECTRO ENCEPHALOGRAPHY | NOTUS NEUROLOGY | VEEG | SANTHOSH | OPD | VIDEO NOT SHOWING | ISSUE RECTIFIED | NO | NO | 2024-06-26 | June 2024 | 18:30 | 2024-06-26 18:25 | 2024-06-26 18:30 | 0:05 | COMPLETED | SEMI-CRITICAL
1840 | 2024-06-26 | Jun 2024 | 11:10 | 11:15 | 0.0034722 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | THANISH | ICU | NOT WORKING | SYRINGE CLAMP HOLDER ISSUE RECTIFIED | HOLDER CLAMP | NO | 2024-06-26 | June 2024 | 11:30 | 2024-06-26 11:10 | 2024-06-26 11:30 | 0:20 | COMPLETED | NON-CRITICAL
1841 | 2024-06-26 | Jun 2024 | 12:05 | 12:10 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQP9 | THANISH | OPD | FREEZE BUTTON ISSUE | REPLACED THE FREEZE BUTTON CONTROLLER BOARD. | FREEZE KEY ASSEMBLY | NO | 2024-06-26 | June 2024 | 12:50 | 2024-06-26 12:05 | 2024-06-26 12:50 | 0:45 | COMPLETED | SEMI-CRITICAL
1842 | 2024-06-27 | Jun 2024 | 9:30 | 9:35 | 0.0034722 | HEAD LIGHT SOURCE, XENON | SUNOPTIC | TITAN X450 | SOORAJ | OT | LAMP NOT GLOWING | FOUND LAMP DOOR LOOSE CONTACT. TEMPORARILY SECURED DOOR WITH TAPE AND ISSUE RECTIFIED | NO | NO | 2024-06-27 | June 2024 | 10:00 | 2024-06-27 09:30 | 2024-06-27 10:00 | 0:30 | COMPLETED | SEMI-CRITICAL
1843 | 2024-06-27 | Jun 2024 | 14:30 | 14:35 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SOORAJ | OT | O2 LEAK ALARM | FOUND O2 CYCLINDER WASHER MISSING. REPLACED WITH NEW ONE AND ISSUE RECTIFIED | WASHER | NO | 2024-06-27 | June 2024 | 15:30 | 2024-06-27 14:30 | 2024-06-27 15:30 | 1:00 | COMPLETED | CRITICAL
1844 | 2024-06-27 | Jun 2024 | 14:00 | 14:10 | 0.0069444 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | JERLIN | RADIOLOGY | CRITICAL FAILURE OCCURRED | RECONNECTED THE USB WITH THE TAB AND RECTIFIED THE ISSUE | NO | NO | 2024-06-27 | June 2024 | 14:30 | 2024-06-27 14:00 | 2024-06-27 14:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1845 | 2024-06-27 | Jun 2024 | 9:30 | 9:32 | 0.0013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | ICU | NOT WORKING | ISSUE WITH POWER SOCKET. REPLACED AND OBSERVED. WORKING GOOD. | NO | NO | 2024-06-27 | June 2024 | 9:40 | 2024-06-27 09:30 | 2024-06-27 09:40 | 0:10 | COMPLETED | NON-CRITICAL
1846 | 2024-06-27 | Jun 2024 | 9:43 | 9:45 | 0.0013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | DIALYSIS | NOT WORKING | POWERCORD GOT DETACHED. REFIXED IT AND READY TO USE. | NO | NO | 2024-06-27 | June 2024 | 9:55 | 2024-06-27 09:43 | 2024-06-27 09:55 | 0:12 | COMPLETED | NON-CRITICAL
1847 | 2024-06-27 | Jun 2024 | 9:55 | 9:56 | 0.00069444 | PATIENT MONITOR | PHILIPS | MX 450 | HAFIZ | ICU | UNOBLE TO CHANGED DATE AND TIME | REMOVED THE CENTRAL LINE CONNECTION (LAN CABLE), SETTINGS DONE AND REFIXED. NOW GOOD. | NO | NO | 2024-06-27 | June 2024 | 10:05 | 2024-06-27 09:55 | 2024-06-27 10:05 | 0:10 | COMPLETED | SEMI-CRITICAL
1848 | 2024-06-27 | Jun 2024 | 11:50 | 11:55 | 0.0034722 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | HAFIZ | OPD | UNOBLE TO CHANGED DATE AND TIME | SETTINGS DONE AS REQUIRED. | NO | NO | 2024-06-27 | June 2024 | 12:05 | 2024-06-27 11:50 | 2024-06-27 12:05 | 0:15 | COMPLETED | NON-CRITICAL
1849 | 2024-06-27 | Jun 2024 | 0:30 | 0:35 | 0.0034722 | VENTILATOR, ADULT | MAQUET | SERVO-I | HAFIZ | ICU | PRE-USE CHECK FAILED | OBSERVED ISSUE WITH AIR PENDANT. PROPERLY FIXED AND NOW READY TO USE. | NO | NO | 2024-06-27 | June 2024 | 1:00 | 2024-06-27 00:30 | 2024-06-27 01:00 | 0:30 | COMPLETED | CRITICAL
1850 | 2024-06-27 | Jun 2024 | 10:55 | 11:00 | 0.0034722 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | HAFIZ | OT | NOT WORKING | CONTROL BOARD REPLACED AND OBSERVED. ISSUE RECTIFIED. | CONTROL BOARD | NO | 2024-06-27 | June 2024 | 11:30 | 2024-06-27 10:55 | 2024-06-27 11:30 | 0:35 | COMPLETED | NON-CRITICAL
1851 | 2024-06-28 | Jun 2024 | 8:30 | 8:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | ICU | COT LOCKED. | RESET DONE AND ISSUE RECTIFIED. | NO | NO | 2024-06-28 | June 2024 | 9:05 | 2024-06-28 08:30 | 2024-06-28 09:05 | 0:35 | COMPLETED | NON-CRITICAL
1852 | 2024-06-28 | Jun 2024 | 15:10 | 15:30 | 0.013889 | PATIENT LIFTER | ARJO HUNTLEIGH | MAXIMOVE | RAMKIE | PHYSIOTHERAPY | NOT WORKING | CHECKED AND FOUND THE PROBLEM WITH BATTER, NEED REPLACEMENT | NO | BATTERY | 2024-06-28 | June 2024 | 15:40 | 2024-06-28 15:10 | 2024-06-28 15:40 | 0:30 | PENDING | SEMI-CRITICAL
1853 | 2024-06-28 | Jun 2024 | 15:35 | 15:40 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | HAFIZ | OT | TIDAL VOLUME ABNORMAL ISSUE | REPLACED STANDBY CSB BOARD AND OBSERVED. NOW WORKING GOOD. | CONTROL SAMPLE BOARD | NO | 2024-06-28 | June 2024 | 16:35 | 2024-06-28 15:35 | 2024-06-28 16:35 | 1:00 | COMPLETED | CRITICAL
1854 | 2024-06-28 | Jun 2024 | 17:05 | 17:10 | 0.0034722 | FREEZER,DEEP FREEZING -40 | REMI | RQVD-200 PLUS | SANTHOSH | BLOODBANK | NOT GETS CHILLING | CHILLING GAS FILLED ISSUE RECTIFIED | NO | NO | 2024-06-28 | June 2024 | 17:20 | 2024-06-28 17:05 | 2024-06-28 17:20 | 0:15 | COMPLETED | SEMI-CRITICAL
1855 | 2024-06-28 | Jun 2024 | 9:05 | 9:10 | 0.0034722 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | THANISH | OPD | TMT NOT WORKING | UNLOCKED THE EMERGENCY SWITCH.WORKING FINE. | NO | NO | 2024-06-28 | June 2024 | 9:15 | 2024-06-28 09:05 | 2024-06-28 09:15 | 0:10 | COMPLETED | NON-CRITICAL
1856 | 2024-06-28 | Jun 2024 | 14:00 | 14:13 | 0.0090278 | BMI MACHINE | AKAS MEDICAL | QVS-100 | THANISH | OPD | NOT SWITCHING ON | OLD POWER CORD FAILED/REPLCAED WITH NEW POWER CORD. | POWER CORD | NO | 2024-06-28 | June 2024 | 14:23 | 2024-06-28 14:00 | 2024-06-28 14:23 | 0:23 | COMPLETED | NON-CRITICAL
1857 | 2024-06-28 | Jun 2024 | 7:45 | 8:00 | 0.010417 | PATIENT MONITOR | PHILIPS | MX 450 | THANISH | ICU | ECG NOT WORKING | REPAIRED THE 5-LEAD ECG CABLE.WORKING FINE. | NO | NO | 2024-06-28 | June 2024 | 8:25 | 2024-06-28 07:45 | 2024-06-28 08:25 | 0:40 | COMPLETED | SEMI-CRITICAL
1858 | 2024-06-29 | Jun 2024 | 16:10 | 16:15 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | SOORAJ | OT | GM OCCLUSION ALARM | FOUND ISSUE WITH WATER TRAP. DRIED IN ROOM TEMPERATURE AND REFIXED, WORKING FINE. | NO | NO | 2024-06-29 | June 2024 | 17:00 | 2024-06-29 16:10 | 2024-06-29 17:00 | 0:50 | COMPLETED | SEMI-CRITICAL
1859 | 2024-06-29 | Jun 2024 | 9:10 | 9:15 | 0.0034722 | VENTILATOR, UNIVERSAL | MAQUET | SERVO-I | THANISH | ICU | EXPIRATORY CASETTE NOT WORKING | REPLACED THE MEMEBRANE FOR EXPIRATORY CASETTE.WORKING FINE. | MEMBRANE | NO | 2024-06-29 | June 2024 | 9:30 | 2024-06-29 09:10 | 2024-06-29 09:30 | 0:20 | COMPLETED | CRITICAL
1860 | 2024-06-30 | Jun 2024 | 20:00 | 20:15 | 0.010417 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | THANISH | ICU | NEED TO INSTALL NEW ABG CATRIDGE | INSTALLED THE NEW ABG CATRIDGE AND CVP TEST PERFORMED.WORKING FINE. | CARTRIDGE | NO | 2024-06-30 | June 2024 | 20:35 | 2024-06-30 20:00 | 2024-06-30 20:35 | 0:35 | COMPLETED | SEMI-CRITICAL
1861 | 2024-07-01 | Jul 2024 | 7:25 | 7:25 | 0 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | CLEANED AND REFIXED THE ABSORBER CANISTER PROPERLY. NOW READY TO USE. | NO | NO | 2024-07-01 | July 2024 | 7:45 | 2024-07-01 07:25 | 2024-07-01 07:45 | 0:20 | COMPLETED | CRITICAL
1862 | 2024-07-01 | Jul 2024 | 8:30 | 8:35 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S8 | HAFIZ | OPD | UNOBLE TO PUSH IMAGES IN PACS. | ISSUE WITH THE LAN CABLE, REFIXED IT. NOW WORKING GOOD. | NO | NO | 2024-07-01 | July 2024 | 8:45 | 2024-07-01 08:30 | 2024-07-01 08:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1863 | 2024-07-01 | Jul 2024 | 10:30 | 10:34 | 0.0027778 | PATIENT MONITOR | PHILIPS | MX 550 | HAFIZ | OT | NOT WORKING. | ISSUE WITH THE BATTERY, REPLACED IT. NOW READY TO USE. | BATTERY | NO | 2024-07-01 | July 2024 | 10:45 | 2024-07-01 10:30 | 2024-07-01 10:45 | 0:15 | COMPLETED | SEMI-CRITICAL
1864 | 2024-07-01 | Jul 2024 | 9:40 | 9:50 | 0.0069444 | DIALYSIS MACHINE | FRESENIUS | 4008 S NG | RAMKIE | DIALYSIS | V 99 FAILURE | CHECKED AND CLEANED FILTERS, PROBLEM RECTIFIED | NO | NO | 2024-07-01 | July 2024 | 10:30 | 2024-07-01 09:40 | 2024-07-01 10:30 | 0:50 | COMPLETED | CRITICAL
1865 | 2024-07-01 | Jul 2024 | 7:25 | 7:30 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | RAMKIE | OT | CIRCUIT LEAK TEST FAILED | CLEANED AND REFIXED THE ABSORBER CANISTER PROPERLY. NOW READY TO USE. | NO | NO | 2024-07-01 | July 2024 | 7:55 | 2024-07-01 07:25 | 2024-07-01 07:55 | 0:30 | COMPLETED | CRITICAL
1866 | 2024-07-01 | Jul 2024 | 22:45 | 22:45 | 0 | NEBULIZER | JK MEDICALS | ULTIMA | THANISH | WARD | KNOB BROKEN | FIXED THE KNOB AND RECTIFIED THE ISSUE | NO | NO | 2024-07-01 | July 2024 | 23:05 | 2024-07-01 22:45 | 2024-07-01 23:05 | 0:20 | COMPLETED | SEMI-CRITICAL
1867 | 2024-07-01 | Jul 2024 | 23:10 | 23:10 | 0 | NEBULIZER | JK MEDICALS | ULTIMA | THANISH | WARD | KNOB BROKEN | FIXED THE KNOB AND RECTIFIED THE ISSUE | NO | NO | 2024-07-01 | July 2024 | 23:15 | 2024-07-01 23:10 | 2024-07-01 23:15 | 0:05 | COMPLETED | SEMI-CRITICAL
1868 | 2024-07-02 | Jul 2024 | 15:00 | 15:12 | 0.0083333 | USG MACHINE | WIPRO GE HEALTH CARE | VOLUSON E10 | HAFIZ | OPD | TV PROBE NOT WORKING. | INSIDE CRYSTAL DAMAGE, SO REPLACED UNDER CONTRACT. NOW WORKING GOOD. | R1C-5-D PROBE | NO | 2024-07-02 | July 2024 | 15:45 | 2024-07-02 15:00 | 2024-07-02 15:45 | 0:45 | COMPLETED | SEMI-CRITICAL
1869 | 2024-07-02 | Jul 2024 | 9:10 | 9:20 | 0.0069444 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | JERLIN | OPD | ACQUISITION MODULE CABLE BROKEN | FIXED THE CABLE AND RECTIFIED THE ISSUE | NO | NO | 2024-07-02 | July 2024 | 9:45 | 2024-07-02 09:10 | 2024-07-02 09:45 | 0:35 | COMPLETED | NON-CRITICAL
1870 | 2024-07-02 | Jul 2024 | 11:40 | 12:30 | 0.034722 | DIALYSIS MACHINE | FRESENIUS | 4008 S | RAMKIE | DIALYSIS | V 99 FAILURE | CHECKED AND CLEANED FILTERS, PROBLEM RECTIFIED | NO | NO | 2024-07-02 | July 2024 | 12:10 | 2024-07-02 11:40 | 2024-07-02 12:10 | 0:30 | COMPLETED | CRITICAL
1871 | 2024-07-02 | Jul 2024 | 7:30 | 7:35 | 0.0034722 | LIGHT OPERATING,TRIPLE DOME | MAQUET | ARDVCS209006A VCS64SF | SANTHOSH | OT | ARM LOOSENED | TIGHTENED | NO | NO | 2024-07-02 | July 2024 | 7:40 | 2024-07-02 07:30 | 2024-07-02 07:40 | 0:10 | COMPLETED | SEMI-CRITICAL
1872 | 2024-07-02 | Jul 2024 | 12:40 | 12:45 | 0.0034722 | MICRO DRILL,DEBRIDING SYSTEM | MEDTRONIC | IPC 1898001 | SANTHOSH | OT | FOOT PEDAL NOT WORKING | ISSUE RECTIFIED | NO | NO | 2024-07-02 | July 2024 | 12:50 | 2024-07-02 12:40 | 2024-07-02 12:50 | 0:10 | COMPLETED | SEMI-CRITICAL
1873 | 2024-07-03 | Jul 2024 | 7:10 | 7:15 | 0.0034722 | BIPAP MACHINE | PHILIPS | A40 | SOORAJ | ICU | VENTILATOR INOPERATIVE | BOARD ISSUE, REPLACED A NEW BIPAP MACHINE | BIPAP MACHINE | NO | 2024-07-03 | July 2024 | 7:20 | 2024-07-03 07:10 | 2024-07-03 07:20 | 0:10 | COMPLETED | SEMI-CRITICAL
1874 | 2024-07-03 | Jul 2024 | 15:00 | 15:05 | 0.0034722 | CATHLAB - BIPLANE | PHILIPS | AZURION 7B20 | THANISH | CATHLAB | COLLISION SENSOR PROBLEM | CHECKED AND IT'S WORKING GOOD AFTER MAINTAINING TEMP. | NO | NO | 2024-07-03 | July 2024 | 15:20 | 2024-07-03 15:00 | 2024-07-03 15:20 | 0:20 | COMPLETED | CRITICAL
1875 | 2024-07-03 | Jul 2024 | 12:30 | 12:35 | 0.0034722 | MONITOR HIGH DEFINITION | OLYMPUS | OEV262H | THANISH | OPD | NEED TO CHANGE PIP SETTINGS | CHANGED THE SETTINGS AND RECTIFIED THE ISSUE. | NO | NO | 2024-07-03 | July 2024 | 12:50 | 2024-07-03 12:30 | 2024-07-03 12:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1876 | 2024-07-03 | Jul 2024 | 10:30 | 10:40 | 0.0069444 | MOVEMENT THERAPHY UNIT, UPPER AND LOWER LIMBS | RECK | MOTOMED LETTO2 | THANISH | PHYSIOTHERAPY | NOT WORKING | CHECKED AND FOUND IT'S WORKING GOOD | NO | NO | 2024-07-03 | July 2024 | 10:55 | 2024-07-03 10:30 | 2024-07-03 10:55 | 0:25 | COMPLETED | SEMI-CRITICAL
1877 | 2024-07-03 | Jul 2024 | 9:30 | 9:40 | 0.0069444 | BMI MACHINE | WELCH ALLYN | TANITA | THANISH | OPD | NOT WORKING | CHECKED AND IT'S WORKING GOOD. | NO | NO | 2024-07-03 | July 2024 | 9:54 | 2024-07-03 09:30 | 2024-07-03 09:54 | 0:24 | COMPLETED | NON-CRITICAL
1878 | 2024-07-03 | Jul 2024 | 12:15 | 12:20 | 0.0034722 | MONITOR HIGH DEFINITION | OLYMPUS | OEV262H | THANISH | OPD | NOT ABLE TO RECORD THE PROCEDURE | CONNECTED THE VIDEO OUT CABLE AND RECORDED THE PROCEDURE | NO | NO | 2024-07-03 | July 2024 | 12:30 | 2024-07-03 12:15 | 2024-07-03 12:30 | 0:15 | COMPLETED | SEMI-CRITICAL
1879 | 2024-07-04 | Jul 2024 | 22:00 | 22:10 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SOORAJ | ICU | CARTRIDGE EXPIRED | NEW CARTRIDGE REPLACED AND CVP DONE. | CARTRIDGE | NO | 2024-07-04 | July 2024 | 22:50 | 2024-07-04 22:00 | 2024-07-04 22:50 | 0:50 | COMPLETED | SEMI-CRITICAL
1880 | 2024-07-04 | Jul 2024 | 11:00 | 11:10 | 0.0069444 | DIALYSIS MACHINE | FRESENIUS | 4008S NG | RAMKIE | DIALYSIS | V 99 FAILURE | CHECKED AND CLEANED FILTERS, PROBLEM RECTIFIED | NO | NO | 2024-07-04 | July 2024 | 12:00 | 2024-07-04 11:00 | 2024-07-04 12:00 | 1:00 | COMPLETED | CRITICAL
1881 | 2024-07-05 | Jul 2024 | 20:20 | 20:25 | 0.0034722 | RADIOGRAPHIC UNIT | CARESTREAM | DR ASCEND CMD | SOORAJ | RADIOLOGY | DETECTOR PLATE NOT LOCKING IN HORIZONTAL POSITION. | SCREWS TIGHTENED AND ISSUE RECTIFIED. | NO | NO | 2024-07-05 | July 2024 | 21:00 | 2024-07-05 20:20 | 2024-07-05 21:00 | 0:40 | COMPLETED | SEMI-CRITICAL
1882 | 2024-07-05 | Jul 2024 | 21:00 | 21:05 | 0.0034722 | RADIOGRAPHIC UNIT | CARESTREAM | DR ASCEND CMD | SOORAJ | RADIOLOGY | COLLIMATION LIGHT NOT WORKING | CHECKED AND REFIXED LIGHT POSITION. ISSUE RECTIFIED. | NO | NO | 2024-07-05 | July 2024 | 21:10 | 2024-07-05 21:00 | 2024-07-05 21:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1883 | 2024-07-05 | Jul 2024 | 23:30 | 23:35 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SOORAJ | WARD | COT HEIGHT NOT ADJUSTING | FOUND COT LOCKED, RESET DONE AND ISSUE RECTIFIED. | NO | NO | 2024-07-05 | July 2024 | 23:55 | 2024-07-05 23:30 | 2024-07-05 23:55 | 0:25 | COMPLETED | NON-CRITICAL
1884 | 2024-07-05 | Jul 2024 | 10:00 | 10:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | SANTHOSH | OT | NOT SWITCHING ON | INTERFACE BOARD REPLACED | INTERFACE BOARD | NO | 2024-07-05 | July 2024 | 10:25 | 2024-07-05 10:00 | 2024-07-05 10:25 | 0:25 | COMPLETED | CRITICAL
1885 | 2024-07-06 | Jul 2024 | 14:00 | 14:10 | 0.0069444 | ANKLE MOTION DEVICE | ANKLE MOTION PRIVATE LIMITED | AM-03 | JERLIN | PHYSIOTHERAPY | POWER ADAPTOR NOT WORKING | FIXED THE POWER ADAPTOR . CHECKED AND IT'S WORKING GOOD. | NO | NO | 2024-07-06 | July 2024 | 14:30 | 2024-07-06 14:00 | 2024-07-06 14:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1886 | 2024-07-06 | Jul 2024 | 21:25 | 21:30 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | ICU | COT FUNCTIONS LOCKED | RESET DONE AND ISSUE RECTIFIED | NO | NO | 2024-07-06 | July 2024 | 21:45 | 2024-07-06 21:25 | 2024-07-06 21:45 | 0:20 | COMPLETED | NON-CRITICAL
1887 | 2024-07-06 | Jul 2024 | 10:10 | 10:30 | 0.013889 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | RAMKIE | OT | OCCLUSION ERROR | CHECKED AND RESTARTED AND CHANGED MODE | NO | NO | 2024-07-06 | July 2024 | 11:00 | 2024-07-06 10:10 | 2024-07-06 11:00 | 0:50 | COMPLETED | NON-CRITICAL
1888 | 2024-07-06 | Jul 2024 | 9:20 | 9:25 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | SANTHOSH | ICU | NOT IMAGE STORING | ISSUE RECTIFIED | NO | NO | 2024-07-06 | July 2024 | 9:30 | 2024-07-06 09:20 | 2024-07-06 09:30 | 0:10 | COMPLETED | SEMI-CRITICAL
1889 | 2024-07-06 | Jul 2024 | 14:20 | 14:25 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SANTHOSH | CATHLAB | NOT SWITCHING ON | ISSUE RECTIFIED | NO | NO | 2024-07-06 | July 2024 | 14:40 | 2024-07-06 14:20 | 2024-07-06 14:40 | 0:20 | COMPLETED | CRITICAL
1890 | 2024-07-06 | Jul 2024 | 12:30 | 12:35 | 0.0034722 | DIGITAL PLATFORM SCALE | SECA | 676 | THANISH | DIALYSIS | POWER CORD NOT WORKING | CHANGED THE CHARGER AND IT'S WORKING GOOD | POWER ADAPTOR | NO | 2024-07-06 | July 2024 | 12:50 | 2024-07-06 12:30 | 2024-07-06 12:50 | 0:20 | COMPLETED | NON-CRITICAL
1891 | 2024-07-06 | Jul 2024 | 10:00 | 10:10 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 550 | THANISH | CATHLAB | WATER TRAP ISSUE | CHANGED THE WATER TRAP AND RECTIFIED THE ISSUE | WATER TRAP | NO | 2024-07-06 | July 2024 | 10:15 | 2024-07-06 10:00 | 2024-07-06 10:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1892 | 2024-07-07 | Jul 2024 | 17:30 | 17:35 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | CARTRIDGE EXPIRED | REPLACED ANEW CARTRIDGE AND CVP DONE | CARTRIDGE | NO | 2024-07-07 | July 2024 | 18:15 | 2024-07-07 17:30 | 2024-07-07 18:15 | 0:45 | COMPLETED | SEMI-CRITICAL
1893 | 2024-07-07 | Jul 2024 | 15:00 | 15:15 | 0.010417 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | WARD | REMOTE NOT WORKING | FIXED THE SIDE RAIL REMOTE CABLE AND RECTIFIED THE ISSUE | NO | NO | 2024-07-07 | July 2024 | 15:35 | 2024-07-07 15:00 | 2024-07-07 15:35 | 0:35 | COMPLETED | NON-CRITICAL
1894 | 2024-07-07 | Jul 2024 | 10:00 | 10:14 | 0.0097222 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | JERLIN | OT | NOT WORKING | FIXED THE CONNECTING HOSE AND CHECKED IT'S WORKING GOOD. | NO | NO | 2024-07-07 | July 2024 | 10:25 | 2024-07-07 10:00 | 2024-07-07 10:25 | 0:25 | COMPLETED | SEMI-CRITICAL
1895 | 2024-07-07 | Jul 2024 | 1:10 | 1:15 | 0.0034722 | ECHO MACHINE | GE HEALTHCARE | VIVID IQ | SOORAJ | ICU | MACHINE NOT SWITCHING ON | FOUND MACHINE NOT SWITHCED OFF PROPERLY, HENCE DID NOT BOOT UP. FORCE RESTART DONE AND ISSUE RECTIFIED. | NO | NO | 2024-07-07 | July 2024 | 1:45 | 2024-07-07 01:10 | 2024-07-07 01:45 | 0:35 | COMPLETED | SEMI-CRITICAL
1896 | 2024-07-07 | Jul 2024 | 17:35 | 17:40 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | SANTHOSH | ICU | STORAGE FULL | DELETED OLD DATA | NO | NO | 2024-07-07 | July 2024 | 17:50 | 2024-07-07 17:35 | 2024-07-07 17:50 | 0:15 | COMPLETED | SEMI-CRITICAL
1897 | 2024-07-08 | Jul 2024 | 9:00 | 9:02 | 0.0013889 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | HAFIZ | OT | NOT WORKING. | ISSUE WITH THE FILTER, CLEANED IT AND REFIXED. NOW WORKING GOOD. | NO | NO | 2024-07-08 | July 2024 | 9:40 | 2024-07-08 09:00 | 2024-07-08 09:40 | 0:40 | COMPLETED | SEMI-CRITICAL
1898 | 2024-07-08 | Jul 2024 | 9:35 | 9:40 | 0.0034722 | STEAM STERILIZER | GETINGE | GSS67H102E | HAFIZ | CSSD | LOW PRESSURE CHAMBER | REFILLED THE WATER AND OBSERVED. WORKING FINE. | NO | NO | 2024-07-08 | July 2024 | 10:00 | 2024-07-08 09:35 | 2024-07-08 10:00 | 0:25 | COMPLETED | SEMI-CRITICAL
1899 | 2024-07-08 | Jul 2024 | 11:00 | 11:02 | 0.0013889 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | HAFIZ | OT | INTERMITTENTLY WORKING | ISSUE WITH THE THERMISTOR WIRE CONNECTION. RESOLDERED AND RECTIFIED. | NO | NO | 2024-07-08 | July 2024 | 11:20 | 2024-07-08 11:00 | 2024-07-08 11:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1900 | 2024-07-08 | Jul 2024 | 15:00 | 15:02 | 0.0013889 | HEATER COOLER | GETINGE | HCU40 | HAFIZ | OT | WATER LEAKAGE IN BLANKET | LEAK ARRESTED AND OBSERVED. NOW READY TO USE. | NO | NO | 2024-07-08 | July 2024 | 15:30 | 2024-07-08 15:00 | 2024-07-08 15:30 | 0:30 | COMPLETED | CRITICAL
1901 | 2024-07-08 | Jul 2024 | 15:30 | 15:35 | 0.0034722 | HEATER COOLER | GETINGE | HCU 40 | HAFIZ | OT | WATER LEAKAGE IN BLANKET | LEAK ARRESTED AND OBSERVED. NOW READY TO USE. | NO | NO | 2024-07-08 | July 2024 | 16:00 | 2024-07-08 15:30 | 2024-07-08 16:00 | 0:30 | COMPLETED | CRITICAL
1902 | 2024-07-08 | Jul 2024 | 4:10 | 4:15 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | JERLIN | ICU | PRINTER NOT WORKING PROPERLY | CLEANED THE PRINTER HEAD AND TEST DONE . ISSUE RECTIFIED | NO | NO | 2024-07-08 | July 2024 | 4:35 | 2024-07-08 04:10 | 2024-07-08 04:35 | 0:25 | COMPLETED | SEMI-CRITICAL
1903 | 2024-07-08 | Jul 2024 | 0:05 | 0:10 | 0.0034722 | HUMIDIFIER AIR | FISHER AND PAYKEL | MR 850 | SOORAJ | ICU | TEMPERATURE NOT MAINTAINING | OBSERVED FOR 30 MINS AND FOUND NO ISSUES. | NO | NO | 2024-07-08 | July 2024 | 0:50 | 2024-07-08 00:05 | 2024-07-08 00:50 | 0:45 | COMPLETED | SEMI-CRITICAL
1904 | 2024-07-08 | Jul 2024 | 5:30 | 5:35 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | SOORAJ | EMERGENCY | REPORTED FLOW DEVIATION | CHECKED FLOW FOR 1 HR AND FOUND NO DEVIATION. RETURNED TO ER AND DID NOT GET ANY COMPLAINS FURTHER | NO | NO | 2024-07-08 | July 2024 | 7:00 | 2024-07-08 05:30 | 2024-07-08 07:00 | 1:30 | COMPLETED | NON-CRITICAL
1905 | 2024-07-08 | Jul 2024 | 11:45 | 12:05 | 0.013889 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | RAMKIE | OT | OCCLUSION ERROR | CHECKED AND RESTARTED AND CHANGED MODE | NO | NO | 2024-07-08 | July 2024 | 12:30 | 2024-07-08 11:45 | 2024-07-08 12:30 | 0:45 | COMPLETED | NON-CRITICAL
1906 | 2024-07-08 | Jul 2024 | 20:30 | 20:35 | 0.0034722 | VENTILATOR, NEONATAL | MAQUET | SERVO-I | SANTHOSH | ICU | WAVEFORM ISSUE | ISSUE RECTIFIED | NO | NO | 2024-07-08 | July 2024 | 20:45 | 2024-07-08 20:30 | 2024-07-08 20:45 | 0:15 | COMPLETED | CRITICAL
1907 | 2024-07-08 | Jul 2024 | 10:45 | 10:50 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SANTHOSH | WARD | VACCUM PRESSURE ISSUE | PILOT VALVE PROBLEM | NO | NO | 2024-07-08 | July 2024 | 11:50 | 2024-07-08 10:45 | 2024-07-08 11:50 | 1:05 | COMPLETED | NON-CRITICAL
1908 | 2024-07-09 | Jul 2024 | 7:30 | 7:35 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 550 | HAFIZ | OT | GM MALFUNCTION. | SWAPPED THE GAS MODULE FROM MX450 MONITOR (NOT NEEDED) FROM ICU. OT MONITOR GIVEN BACK TO OT. | GAS MODULE | NO | 2024-07-09 | July 2024 | 8:30 | 2024-07-09 07:30 | 2024-07-09 08:30 | 1:00 | COMPLETED | SEMI-CRITICAL
1909 | 2024-07-09 | Jul 2024 | 15:30 | 15:40 | 0.0069444 | HUMIDIFIER AIR | FISHER AND PAYKEL | MR 810 | JERLIN | ICU | TEMP NOT ACHIVED | RESET DONE AND ISSUE RECTIFIED | NO | NO | 2024-07-09 | July 2024 | 15:55 | 2024-07-09 15:30 | 2024-07-09 15:55 | 0:25 | COMPLETED | SEMI-CRITICAL
1910 | 2024-07-09 | Jul 2024 | 13:00 | 13:10 | 0.0069444 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | JERLIN | ICU | BARREL HOLDER NOT WORKING | REPLACED THE HOLDER AND RECTIFIED THE ISSUE. | HOLDER CLAMP | NO | 2024-07-09 | July 2024 | 13:30 | 2024-07-09 13:00 | 2024-07-09 13:30 | 0:30 | COMPLETED | NON-CRITICAL
1911 | 2024-07-09 | Jul 2024 | 12:50 | 13:05 | 0.010417 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | OT | NOT WORKING | CHECKED AND RESET DONE, PROBLEM RECTIFED | NO | NO | 2024-07-09 | July 2024 | 13:40 | 2024-07-09 12:50 | 2024-07-09 13:40 | 0:50 | COMPLETED | NON-CRITICAL
1912 | 2024-07-09 | Jul 2024 | 10:45 | 10:50 | 0.0034722 | STEAM STERILIZER | GETINGE | GSS67H102E | SANTHOSH | CSSD | NOT WORKING | ISSUE RECTIFIED | NO | NO | 2024-07-09 | July 2024 | 12:00 | 2024-07-09 10:45 | 2024-07-09 12:00 | 1:15 | COMPLETED | SEMI-CRITICAL
1913 | 2024-07-10 | Jul 2024 | 8:20 | 8:22 | 0.0013889 | PATIENT MONITOR | PHILIPS | MX 550 | HAFIZ | OT | GM NOT WORKING. | REFIXED IN ANOTHER SLOT AND OBSERVED WORKING FINE. | NO | NO | 2024-07-10 | July 2024 | 8:30 | 2024-07-10 08:20 | 2024-07-10 08:30 | 0:10 | COMPLETED | SEMI-CRITICAL
1914 | 2024-07-10 | Jul 2024 | 11:35 | 11:37 | 0.0013889 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | HAFIZ | OT | NOT WORKING. | ISSUE WITH THE THERMISTOR WIRE CONNECTION. RECONNECTED AND RECTIFIED. | NO | NO | 2024-07-10 | July 2024 | 11:55 | 2024-07-10 11:35 | 2024-07-10 11:55 | 0:20 | COMPLETED | SEMI-CRITICAL
1915 | 2024-07-10 | Jul 2024 | 15:30 | 15:40 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | WARD | HEAD END ACTUATOR NOT WORKING | REPLACED THE ACTUATOR AND NOW IT'S WORKING GOOD. | ACTUATOR | NO | 2024-07-10 | July 2024 | 15:55 | 2024-07-10 15:30 | 2024-07-10 15:55 | 0:25 | COMPLETED | NON-CRITICAL
1916 | 2024-07-11 | Jul 2024 | 13:55 | 14:10 | 0.010417 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | ICU | REMOTE NOT WORKING | FOUND THAT REMOTE WAS LOCKED , UNLOCKED AND RECTIFIEDD THE ISSUE | NO | NO | 2024-07-11 | July 2024 | 14:15 | 2024-07-11 13:55 | 2024-07-11 14:15 | 0:20 | COMPLETED | NON-CRITICAL
1917 | 2024-07-11 | Jul 2024 | 15:40 | 15:45 | 0.0034722 | BLOOD GAS ANALYSER | GEM 3500 | INSTRUMENTATION LABORATORY | SOORAJ | ICU | ERROR 1.13 | FOUND NURSING STAFF RAN SAMPLE WITH CLOT. RAN HEPARIN AND REFIXED CARTRIDGE. ISSUE RECTIFIED. | NO | NO | 2024-07-11 | July 2024 | 16:15 | 2024-07-11 15:40 | 2024-07-11 16:15 | 0:35 | COMPLETED | SEMI-CRITICAL
1918 | 2024-07-11 | Jul 2024 | 13:30 | 13:45 | 0.010417 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | RAMKIE | CATHLAB | TEST NOT PASSED | CHECKED AND RECCONNECTED ECG CABLEM PROBLEM RECTIFIED | NO | NO | 2024-07-11 | July 2024 | 14:30 | 2024-07-11 13:30 | 2024-07-11 14:30 | 1:00 | COMPLETED | CRITICAL
1919 | 2024-07-12 | Jul 2024 | 7:00 | 7:05 | 0.0034722 | PATIENT MONITOR | PHILIPS | MR 200 | HAFIZ | RADIOLOGY | UNOBLE TO MEASURE NIBP. | ISSUE WITH HOSE, REFIXED IT PROPERLY AND OBSERVED. ISSUE RECTIFIED. | NO | NO | 2024-07-12 | July 2024 | 7:20 | 2024-07-12 07:00 | 2024-07-12 07:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1920 | 2024-07-12 | Jul 2024 | 14:15 | 14:30 | 0.010417 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | EMERGENCY | NOT ABLE TO FIX THE SIDE RAIL | LUBRICATED THE SIDE RAIL AND RECTIFIED THE ISSUE | NO | NO | 2024-07-12 | July 2024 | 14:37 | 2024-07-12 14:15 | 2024-07-12 14:37 | 0:22 | COMPLETED | NON-CRITICAL
1921 | 2024-07-13 | Jul 2024 | 8:15 | 8:20 | 0.0034722 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | HAFIZ | ICU | HOLDER CLAMP ISSUE. | REPLACED THE HOLDER CLAMP AND ISSUE RECTIFIED. | HOLDER CLAMP | NO | 2024-07-13 | July 2024 | 8:30 | 2024-07-13 08:15 | 2024-07-13 08:30 | 0:15 | COMPLETED | NON-CRITICAL
1922 | 2024-07-13 | Jul 2024 | 11:00 | 11:14 | 0.0097222 | SYRINGE PUMP | SMITHS MEDICAL | GRASEBY 2100 | JERLIN | ICU | NOT WORKING | REPLACED THE MAIN BOARD AND RECTIFIED THE ISSUE | NO | NO | 2024-07-13 | July 2024 | 11:25 | 2024-07-13 11:00 | 2024-07-13 11:25 | 0:25 | COMPLETED | NON-CRITICAL
1923 | 2024-07-13 | Jul 2024 | 19:22 | 19:25 | 0.0020833 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIQ S7 | SOORAJ | RADIOLOGY | ABDOMEN PROBE NOT DETETCING. | FOUND PROBE CONNECTED IMPROPERLY. FIXED PROPERLY AND ISSUE RECTIFIED. | NO | NO | 2024-07-13 | July 2024 | 20:15 | 2024-07-13 19:22 | 2024-07-13 20:15 | 0:53 | COMPLETED | SEMI-CRITICAL
1924 | 2024-07-13 | Jul 2024 | 2:00 | 2:10 | 0.0069444 | VENTILATOR, NEONATAL | MAQUET | SERVO-I | THANISH | ICU | NOT ABLE TO PERFORM PRE USE CHECK | PERFORMED PREUSE CHECK AND RECTIFIED THE ISSUE | NO | NO | 2024-07-13 | July 2024 | 2:30 | 2024-07-13 02:00 | 2024-07-13 02:30 | 0:30 | COMPLETED | CRITICAL
1925 | 2024-07-14 | Jul 2024 | 7:15 | 7:20 | 0.0034722 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | SANTHOSH | RADIOLOGY | CAERESTREAM SOFTWARE ISSUE | ISSUE RECTIFIED | NO | NO | 2024-07-14 | July 2024 | 7:30 | 2024-07-14 07:15 | 2024-07-14 07:30 | 0:15 | COMPLETED | SEMI-CRITICAL
1926 | 2024-07-15 | Jul 2024 | 3:30 | 3:40 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | THANISH | WARD | NOT WORKING | CHANGED THE POWER CORD AND RECTIFIED THE ISSUE | POWER CORD | NO | 2024-07-15 | July 2024 | 3:55 | 2024-07-15 03:30 | 2024-07-15 03:55 | 0:25 | COMPLETED | NON-CRITICAL
1927 | 2024-07-16 | Jul 2024 | 5:20 | 5:30 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | WARD | NOT WORKING. | ISSUE WITH THE POWERSOCKET, CHANGE. NOW WORKING FINE. | NO | NO | 2024-07-16 | July 2024 | 5:40 | 2024-07-16 05:20 | 2024-07-16 05:40 | 0:20 | COMPLETED | NON-CRITICAL
1928 | 2024-07-16 | Jul 2024 | 15:00 | 15:15 | 0.010417 | BIS MONITOR | COVIDIEN | BIS VISTA | JERLIN | OT | BIS MONITOR NOT WORKING | CHECKED AND IT'S WORKING GOOD. | NO | NO | 2024-07-16 | July 2024 | 15:35 | 2024-07-16 15:00 | 2024-07-16 15:35 | 0:35 | COMPLETED | SEMI-CRITICAL
1929 | 2024-07-16 | Jul 2024 | 8:00 | 8:10 | 0.0069444 | PULMONORY FUNCTION TESTING MACHINE | NDD | EASY ONE PRO | JERLIN | OPD | NOT ABLE TO CONNECT WITH THE MODULE | RESET DONE IN THE BLUETOOTH SETTINGS AND RECTIFIED THE ISSUE | NO | NO | 2024-07-16 | July 2024 | 8:26 | 2024-07-16 08:00 | 2024-07-16 08:26 | 0:26 | COMPLETED | SEMI-CRITICAL
1930 | 2024-07-16 | Jul 2024 | 20:00 | 20:05 | 0.0034722 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | SOORAJ | RADIOLOGY | NO POWER BACKUP FOR MOBILE XRAY TAB | FOUND THE BATTERY OF TAB DEFECTIVE. NEW UPS CONNECTED AND BACKUP PROVIDED. ISSUE RECTIFIED. | UPS | NO | 2024-07-16 | July 2024 | 21:30 | 2024-07-16 20:00 | 2024-07-16 21:30 | 1:30 | COMPLETED | SEMI-CRITICAL
1931 | 2024-07-16 | Jul 2024 | 14:30 | 14:45 | 0.010417 | VENTILATOR, UNIVERSAL | MAQUET | SERVO-I | RAMKIE | ICU | TEST NOT PASSED | CHECKED AND CASSETTE RECONNECTED, DIAPHRAGM CLEANED AND PROBLEM CLEARED | NO | NO | 2024-07-16 | July 2024 | 15:30 | 2024-07-16 14:30 | 2024-07-16 15:30 | 1:00 | COMPLETED | CRITICAL
1932 | 2024-07-16 | Jul 2024 | 10:30 | 10:35 | 0.0034722 | EXAMINATION COUCH | MIDMARK JANOK | E122012 | SANTHOSH | OPD | NOT WORKING | ISSUE RECTIFIED | CARTRIDGE | NO | 2024-07-16 | July 2024 | 10:40 | 2024-07-16 10:30 | 2024-07-16 10:40 | 0:10 | COMPLETED | NON-CRITICAL
1933 | 2024-07-17 | Jul 2024 | 21:00 | 21:05 | 0.0034722 | DIALYSIS MACHINE | FRESENIUS | 4008S NG | HAFIZ | DIALYSIS | UNOBLE TO MEASURE NIBP. | REPLACED 4 NOS. NIBP CUFFS WITH HOSES. | BP CUFF | NO | 2024-07-17 | July 2024 | 21:30 | 2024-07-17 21:00 | 2024-07-17 21:30 | 0:30 | COMPLETED | CRITICAL
1934 | 2024-07-17 | Jul 2024 | 22:10 | 22:14 | 0.0027778 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | HAFIZ | EMERGENCY | ECG LEADWIRES DAMAGED | REPLACED IT WITH A NEW 10-LEADS ECG LEADWIRE SET. WORKING GOOD. | 10-LEADS ECG CABLE | NO | 2024-07-17 | July 2024 | 22:30 | 2024-07-17 22:10 | 2024-07-17 22:30 | 0:20 | COMPLETED | SEMI-CRITICAL
1935 | 2024-07-17 | Jul 2024 | 14:40 | 14:50 | 0.0069444 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | JERLIN | EMERGENCY | V2 LEAD NOT WORKING | FOUND THAT THE ELECTRODE WAS TORN . REPLACED THE CABLE AND ISSUE RECTIFIED | 10-LEADS ECG CABLE | NO | 2024-07-17 | July 2024 | 15:05 | 2024-07-17 14:40 | 2024-07-17 15:05 | 0:25 | COMPLETED | SEMI-CRITICAL
1936 | 2024-07-17 | Jul 2024 | 17:40 | 17:45 | 0.0034722 | BLOOD GAS ANALYSER | GEM 3500 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | CADRIDGE FINISHED | REPLACED CADRIDGE | CARTRIDGE | NO | 2024-07-17 | July 2024 | 17:50 | 2024-07-17 17:40 | 2024-07-17 17:50 | 0:10 | COMPLETED | SEMI-CRITICAL
1937 | 2024-07-17 | Jul 2024 | 11:15 | 11:20 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | THANISH | OT | NOT WORKING | CONNECTED THE TEMP SENSOR PROPERLY AND RECTIFIED THE ISSUE. | NO | NO | 2024-07-17 | July 2024 | 11:30 | 2024-07-17 11:15 | 2024-07-17 11:30 | 0:15 | COMPLETED | SEMI-CRITICAL
1938 | 2024-07-18 | Jul 2024 | 1:00 | 1:04 | 0.0027778 | BIOCHEMISTRY ANALYSER | ROCHE | COBAS C 501 | HAFIZ | LABORATORY | SIPPER NOZZLE UP/DOWN ERROR | REFIXED THE NOZZLE CORRECTLY AND RECTIFIED. | NO | NO | 2024-07-18 | July 2024 | 5:30 | 2024-07-18 01:00 | 2024-07-18 05:30 | 4:30 | COMPLETED | SEMI-CRITICAL
1939 | 2024-07-18 | Jul 2024 | 5:00 | 5:02 | 0.0013889 | BABY WEIGHING SCALE WITH INFANTOMETER | SECA | 334 | HAFIZ | WARD | NOT WORKING. | BATTERY REPLACED AND WORKING. | BATTERY | NO | 2024-07-18 | July 2024 | 5:30 | 2024-07-18 05:00 | 2024-07-18 05:30 | 0:30 | COMPLETED | NON-CRITICAL
1940 | 2024-07-18 | Jul 2024 | 7:30 | 7:35 | 0.0034722 | VENTILATOR, UNIVERSAL | MINDRAY MEDICAL INDIA | SV 300 | JERLIN | ICU | CHECKOUT FAILED | CLEANED THE INSPIRATORY AND EXPIRATORY VALVE AND ISSUE RECTIFIED | NO | NO | 2024-07-18 | July 2024 | 8:00 | 2024-07-18 07:30 | 2024-07-18 08:00 | 0:30 | COMPLETED | CRITICAL
1941 | 2024-07-18 | Jul 2024 | 12:00 | 12:05 | 0.0034722 | TABLE TOP STEAM STERILIZER | GETINGE | HS 22K7 | SANTHOSH | CSSD | NOT WORKING | ISSUE RECTIFIED | NO | NO | 2024-07-18 | July 2024 | 12:10 | 2024-07-18 12:00 | 2024-07-18 12:10 | 0:10 | COMPLETED | SEMI-CRITICAL
1942 | 2024-07-19 | Jul 2024 | 16:15 | 16:20 | 0.0034722 | HEATER COOLER | GETINGE | HCU40 | JERLIN | OT | LEAK IN THE HEMOTHERM BLANKET | LEAK ARRESTED AND RECTIFIED THE ISSUE | NO | NO | 2024-07-19 | July 2024 | 16:45 | 2024-07-19 16:15 | 2024-07-19 16:45 | 0:30 | COMPLETED | CRITICAL
1943 | 2024-07-19 | Jul 2024 | 15:00 | 15:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AVANCE CS2 | SANTHOSH | OT | FLASH CARD PROBLEM | REPLACED FROM ANOTHER MACHINE | NO | NO | 2024-07-19 | July 2024 | 15:10 | 2024-07-19 15:00 | 2024-07-19 15:10 | 0:10 | COMPLETED | CRITICAL
1944 | 2024-07-19 | Jul 2024 | 16:30 | 16:40 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 450 | THANISH | ICU | PEDIATRIC TEMPERATURE PROBE NOT GIVEN | REPLACEMENT PROVIDE AND CHECKED | PEDIATRIC TEMP PROBE | NO | 2024-07-19 | July 2024 | 16:50 | 2024-07-19 16:30 | 2024-07-19 16:50 | 0:20 | COMPLETED | SEMI-CRITICAL
1945 | 2024-07-19 | Jul 2024 | 15:00 | 15:15 | 0.010417 | OXYGEN THERAPY UNIT | INSPIRED | O2FLO | THANISH | ICU | CHECKOUT NOT PASSED | CLEANED THE FILTER AND NOW ITS WORKING GOOD | NO | NO | 2024-07-19 | July 2024 | 15:30 | 2024-07-19 15:00 | 2024-07-19 15:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1946 | 2024-07-20 | Jul 2024 | 15:10 | 15:15 | 0.0034722 | VENTILATOR, PEDIATRIC | MINDRAY MEDICAL INDIA | SV 300 | SANTHOSH | ICU | O2 CELL FAIL | O2 CELL DEFECTIVE AND NEED REPLACEMENT | NO | O2 CELL | 2024-07-20 | July 2024 | 15:20 | 2024-07-20 15:10 | 2024-07-20 15:20 | 0:10 | PENDING | CRITICAL
1947 | 2024-07-21 | Jul 2024 | 10:15 | 10:20 | 0.0034722 | PATIENT WARMER | SMITHS MEDICAL | L1CW | SOORAJ | WARD | WARMER NOT SWITCHING ON | FOUND HOSE CONENCTED IMPROPERLY. RECONNECTED AND ISSUE RECTIFIED. | NO | NO | 2024-07-21 | July 2024 | 10:45 | 2024-07-21 10:15 | 2024-07-21 10:45 | 0:30 | COMPLETED | SEMI-CRITICAL
1948 | 2024-07-21 | Jul 2024 | 12:50 | 12:55 | 0.0034722 | HEAD LIGHT SOURCE, LED | SUNOPTIC | SSL-5500 | SOORAJ | ICU | HEADLIGHT BATTERY NOT CHARGING | FOUND THE BATTERY TO BE DEFECTIVE. ALTERNOTE ONE MORE BATTERY AVAILABLE | NO | NO | 2024-07-21 | July 2024 | 13:50 | 2024-07-21 12:50 | 2024-07-21 13:50 | 1:00 | COMPLETED | SEMI-CRITICAL
1949 | 2024-07-21 | Jul 2024 | 12:10 | 12:15 | 0.0034722 | BLOOD GAS ANALYSER | GEM 3500 | INSTRUMENTATION LABORATORY | SOORAJ | ICU | CARTRIDGE EMPTY MESSAGE | NEW CARTRIDGE REPLACED. | CARTRIDGE | NO | 2024-07-21 | July 2024 | 12:45 | 2024-07-21 12:10 | 2024-07-21 12:45 | 0:35 | COMPLETED | SEMI-CRITICAL
1950 | 2024-07-21 | Jul 2024 | 15:30 | 15:35 | 0.0034722 | STEAM STERILIZER | GETINGE | GSS67H102E | SOORAJ | CSSD | INTERNAL LEAKAGE INSIDE | FOUND ISSUE WITH HOSE. NEEDS REPLACEMENT. | NO | HOSE | 2024-07-21 | July 2024 | 17:00 | 2024-07-21 15:30 | 2024-07-21 17:00 | 1:30 | PENDING | SEMI-CRITICAL
1951 | 2024-07-21 | Jul 2024 | 17:00 | 17:00 | 0 | TABLE TOP STEAM STERILIZER | GETINGE | HS 22K7 | SOORAJ | CSSD | PRESSURE NOT RAISING | FOUND ISSUE WITH DOOR GASKET. REPLACED GASKET AND ISSUE RECTIFIED. | GASKET | NO | 2024-07-21 | July 2024 | 20:30 | 2024-07-21 17:00 | 2024-07-21 20:30 | 3:30 | COMPLETED | SEMI-CRITICAL
1952 | 2024-07-21 | Jul 2024 | 12:15 | 12:20 | 0.0034722 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | SANTHOSH | RADIOLOGY | SOFTWARE ISSUE | ISSUE RECTIFIED | NO | NO | 2024-07-21 | July 2024 | 12:25 | 2024-07-21 12:15 | 2024-07-21 12:25 | 0:10 | COMPLETED | SEMI-CRITICAL
1953 | 2024-07-22 | Jul 2024 | 9:45 | 9:50 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | WARD | NOT WORKING. | RESET DONE, OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-07-22 | July 2024 | 15:00 | 2024-07-22 09:45 | 2024-07-22 15:00 | 5:15 | COMPLETED | NON-CRITICAL
1954 | 2024-07-22 | Jul 2024 | 18:00 | 18:02 | 0.0013889 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | NEED TO REPLACE A NEW CARTRIDGE. | REPLACED A NEW CARTRIDGE. | CARTRIDGE | NO | 2024-07-22 | July 2024 | 18:40 | 2024-07-22 18:00 | 2024-07-22 18:40 | 0:40 | COMPLETED | SEMI-CRITICAL
1955 | 2024-07-22 | Jul 2024 | 19:15 | 19:20 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | JERLIN | ICU | CARTRIDGE EXPIRED | REPLACED ANEW CARTRIDGE AND CVP DONE | CARTRIDGE | NO | 2024-07-22 | July 2024 | 19:45 | 2024-07-22 19:15 | 2024-07-22 19:45 | 0:30 | COMPLETED | SEMI-CRITICAL
1956 | 2024-07-22 | Jul 2024 | 18:25 | 18:30 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | SANTHOSH | ICU | NOT GETTING HEAT | ISSUE RECTIFIED | NO | NO | 2024-07-22 | July 2024 | 18:30 | 2024-07-22 18:25 | 2024-07-22 18:30 | 0:05 | COMPLETED | SEMI-CRITICAL
1957 | 2024-07-23 | Jul 2024 | 17:05 | 17:10 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | BLOODBANK | CATRIDGE REMOVED | CATDRIDGE REPLACED | CARTRIDGE | NO | 2024-07-23 | July 2024 | 17:20 | 2024-07-23 17:05 | 2024-07-23 17:20 | 0:15 | COMPLETED | SEMI-CRITICAL
1958 | 2024-07-23 | Jul 2024 | 11:40 | 11:45 | 0.0034722 | MICROSCOPE OPERATING, SURGICAL | CARL ZEISS | OPMI SENSARA | THANISH | OT | NOT ABLE TO USE SLAVE MONITOR | CHANGED SETTINGS AND CONNECTED TO THE SLAVE MONITOR | NO | NO | 2024-07-23 | July 2024 | 11:55 | 2024-07-23 11:40 | 2024-07-23 11:55 | 0:15 | COMPLETED | SEMI-CRITICAL
1959 | 2024-07-24 | Jul 2024 | 12:04 | 12:09 | 0.0034722 | MAMMOGRAPH UNIT | KIRAN | FELICIA | HAFIZ | RADIOLOGY | HANGING ISSUE. | RESTARTED THE SOFTWARE AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-07-24 | July 2024 | 12:34 | 2024-07-24 12:04 | 2024-07-24 12:34 | 0:30 | COMPLETED | SEMI-CRITICAL
1960 | 2024-07-24 | Jul 2024 | 18:00 | 18:05 | 0.0034722 | STEAM STERILIZER | GETINGE | GSS67H102E | HAFIZ | CSSD | PRECIPITATION FORMED INSIDE. | FOUND TDS VALUE GREATER THAN NORMAL RANGE. RESOLVED IT WITH THE SUPPORT OF MAINTENONCE TEAM. | NO | NO | 2024-07-24 | July 2024 | 18:30 | 2024-07-24 18:00 | 2024-07-24 18:30 | 0:30 | COMPLETED | SEMI-CRITICAL
1961 | 2024-07-24 | Jul 2024 | 15:30 | 15:50 | 0.013889 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | RAMKIE | ICU | TEST NOT PASSED | CHECKED AND RECCONNECTED ECG CABLEM PROBLEM RECTIFIED | NO | NO | 2024-07-24 | July 2024 | 16:20 | 2024-07-24 15:30 | 2024-07-24 16:20 | 0:50 | COMPLETED | CRITICAL
1962 | 2024-07-24 | Jul 2024 | 14:05 | 14:10 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | SANTHOSH | WARD | NOT WORKING | ISSUE RECTIFIED | NO | NO | 2024-07-24 | July 2024 | 14:20 | 2024-07-24 14:05 | 2024-07-24 14:20 | 0:15 | COMPLETED | NON-CRITICAL
1963 | 2024-07-24 | Jul 2024 | 14:05 | 14:10 | 0.0034722 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | SANTHOSH | WARD | NOT WORKING | ISSUE RECTIFIED | NO | NO | 2024-07-24 | July 2024 | 14:20 | 2024-07-24 14:05 | 2024-07-24 14:20 | 0:15 | COMPLETED | NON-CRITICAL
1964 | 2024-07-24 | Jul 2024 | 8:00 | 8:10 | 0.0069444 | PATIENT MONITOR | PHILIPS | MX 550 | THANISH | OT | SLAVE MONITOR NOT WORKING | CHANGED THE SETTINGS AND RECTIFIED THE ISSUE. | NO | NO | 2024-07-24 | July 2024 | 8:20 | 2024-07-24 08:00 | 2024-07-24 08:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1965 | 2024-07-24 | Jul 2024 | 11:00 | 11:10 | 0.0069444 | MONITOR HIGH DEFINITION | KARL STORZ | TM220 | THANISH | OT | DISPLAY COLOUR CHANGED | CAMERA CABLE CONNECTED PROPERLY AND RECTIFIED THE ISSUE. | NO | NO | 2024-07-24 | July 2024 | 11:20 | 2024-07-24 11:00 | 2024-07-24 11:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1966 | 2024-07-24 | Jul 2024 | 10:00 | 10:15 | 0.010417 | ACT MACHINE | HELENO LABORATORIES | ACTALYKE | THANISH | CATHLAB | POL ERROR | CHECKED WITH ANOTHER SAMPLE AND IT'S WORKING FINE. | NO | NO | 2024-07-24 | July 2024 | 10:25 | 2024-07-24 10:00 | 2024-07-24 10:25 | 0:25 | COMPLETED | SEMI-CRITICAL
1967 | 2024-07-24 | Jul 2024 | 7:15 | 7:20 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | THANISH | OT | LEAK IN BP CUFF | FIXED IT AND NOW IT'S WORKING GOOD | NO | NO | 2024-07-24 | July 2024 | 7:25 | 2024-07-24 07:15 | 2024-07-24 07:25 | 0:10 | COMPLETED | SEMI-CRITICAL
1968 | 2024-07-25 | Jul 2024 | 15:30 | 15:35 | 0.0034722 | LABOUR BED | ARJO HUNTLEIGH | 170000E1L | HAFIZ | WARD | COUCH'S REMOTE MISSING | REPLACED A NEW REMOTE UNDER CONTRACT. | REMOTE | NO | 2024-07-25 | July 2024 | 16:00 | 2024-07-25 15:30 | 2024-07-25 16:00 | 0:30 | COMPLETED | NON-CRITICAL
1969 | 2024-07-26 | Jul 2024 | 19:00 | 19:00 | 0 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | CVP DUE | RUNNED ALL THE CVP AND READY TO USE. | CARTRIDGE | NO | 2024-07-26 | July 2024 | 19:15 | 2024-07-26 19:00 | 2024-07-26 19:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1970 | 2024-07-27 | Jul 2024 | 10:10 | 10:12 | 0.0013889 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | HAFIZ | ENDOSCOPY | OP CHECK FAILED. | PROPERLY FIXED THERAPY PORT AND RERUNNED THE TEST AND PASSED. | NO | NO | 2024-07-27 | July 2024 | 10:20 | 2024-07-27 10:10 | 2024-07-27 10:20 | 0:10 | COMPLETED | CRITICAL
1971 | 2024-07-27 | Jul 2024 | 20:30 | 20:40 | 0.0069444 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | JERLIN | OT | MOTOR FAULT | REPLACED THE MOTOR AND RECTIFIED THE ISSUE | MOTOR | NO | 2024-07-27 | July 2024 | 21:05 | 2024-07-27 20:30 | 2024-07-27 21:05 | 0:35 | COMPLETED | NON-CRITICAL
1972 | 2024-07-27 | Jul 2024 | 16:30 | 16:50 | 0.013889 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID E95 | RAMKIE | OPD | HANGING PROBLEM | COMPANY PERSON UPDATED SAFTWARE AND PROBLEM RECTIFIED | NO | NO | 2024-07-27 | July 2024 | 17:30 | 2024-07-27 16:30 | 2024-07-27 17:30 | 1:00 | COMPLETED | SEMI-CRITICAL
1973 | 2024-07-28 | Jul 2024 | 3:30 | 3:35 | 0.0034722 | VENTILATOR, UNIVERSAL | MAQUET | SERVO-I | JERLIN | ICU | INTERNOL LEAKAGE | FIXED THE EXPIRATORY CASSETTE PROPERLY AND RECTIFIED THE ISSUE | NO | NO | 2024-07-28 | July 2024 | 3:55 | 2024-07-28 03:30 | 2024-07-28 03:55 | 0:25 | COMPLETED | CRITICAL
1974 | 2024-07-28 | Jul 2024 | 2:55 | 3:00 | 0.0034722 | CEREBRAL OXIMETRY SYSTEM | COVIDIEN | SOMANETICS | JERLIN | OT | READING NOT DISPLAYED PROPERLY | FOUND THAT THE CABLE WAS LOOSELY CONNECTED , FIXED AND ISSUE RECTIFIED. | NO | NO | 2024-07-28 | July 2024 | 3:15 | 2024-07-28 02:55 | 2024-07-28 03:15 | 0:20 | COMPLETED | SEMI-CRITICAL
1975 | 2024-07-28 | Jul 2024 | 16:00 | 16:30 | 0.020833 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | RAMKIE | WARD | NOT WORKING | RESET DONE AND PROBLEM RECTIFIED | NO | NO | 2024-07-28 | July 2024 | 17:30 | 2024-07-28 16:00 | 2024-07-28 17:30 | 1:30 | COMPLETED | NON-CRITICAL
1976 | 2024-07-29 | Jul 2024 | 23:25 | 23:30 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | WARD | NOT WORKING | POWER CORD WAS REMOVED FROM THE CONTROL BOX , FIXED IT AND RECTIFIED THE ISSUE. | NO | NO | 2024-07-29 | July 2024 | 23:45 | 2024-07-29 23:25 | 2024-07-29 23:45 | 0:20 | COMPLETED | NON-CRITICAL
1977 | 2024-07-30 | Jul 2024 | 7:00 | 7:05 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SOORAJ | OT | AC POWER FAIL ISSUE | CHECEKD AND FOUND POWER CHORD LOOSE CONTACT. RECONNECTED AND ISSUE RECTIFIED. | NO | NO | 2024-07-30 | July 2024 | 7:45 | 2024-07-30 07:00 | 2024-07-30 07:45 | 0:45 | COMPLETED | CRITICAL
1978 | 2024-07-30 | Jul 2024 | 16:20 | 16:40 | 0.013889 | ECHO MACHINE | GE HEALTHCARE | VIVID S70 | RAMKIE | ICU | SOFTWARE CORRUPTED | COMPANY CHECKED AND UPDATED SOFTWARE, PROBLEM RECTIFIED | NO | NO | 2024-07-30 | July 2024 | 17:20 | 2024-07-30 16:20 | 2024-07-30 17:20 | 1:00 | COMPLETED | SEMI-CRITICAL
1979 | 2024-07-31 | Jul 2024 | 11:30 | 11:35 | 0.0034722 | WASHER,ENDOSCOPY WASHING | OLYMPUS | OER-AW | HAFIZ | ENDOSCOPY | E-01 ERROR | ISSUE WITH WATER SUPPLY, CLEANED THE FILTER AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-07-31 | July 2024 | 12:00 | 2024-07-31 11:30 | 2024-07-31 12:00 | 0:30 | COMPLETED | SEMI-CRITICAL
1980 | 2024-07-31 | Jul 2024 | 8:50 | 8:55 | 0.0034722 | TOURNIQUET,DUAL CUFF | VBM MEDZINTECHNIK | 9000 | SOORAJ | OT | LEAK ALARM | HOSE RECONNECTED AND ISSUE RECTIFIED | NO | NO | 2024-07-31 | July 2024 | 9:30 | 2024-07-31 08:50 | 2024-07-31 09:30 | 0:40 | COMPLETED | SEMI-CRITICAL
1981 | 2024-07-31 | Jul 2024 | 9:30 | 9:30 | 0 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | REDUCED SODA LIME QTY, REFIXED AL CONNECTIONS, ISSUE RECTIFIED. | NO | NO | 2024-07-31 | July 2024 | 9:50 | 2024-07-31 09:30 | 2024-07-31 09:50 | 0:20 | COMPLETED | CRITICAL
1982 | 2024-07-31 | Jul 2024 | 12:20 | 12:25 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SOORAJ | OT | NO BATTERY BACKUP ISSUE | POWER MANOGEMENT BOARD DEFECTIVE. CHANGED PMB BOARD AND ISSUE RECTIFIED | POWER MANOGEMENT BOARD | NO | 2024-07-31 | July 2024 | 13:40 | 2024-07-31 12:20 | 2024-07-31 13:40 | 1:20 | COMPLETED | CRITICAL
1983 | 2024-07-31 | Jul 2024 | 10:20 | 10:40 | 0.013889 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | RAMKIE | ICU | VOLUME DELIVERY RECTIFIED | INTRUCTED IT IS WARNING, NORMAL IN PRVC MODE AND NO PROBLEM WITH VENTILATOR | NO | NO | 2024-07-31 | July 2024 | 11:00 | 2024-07-31 10:20 | 2024-07-31 11:00 | 0:40 | COMPLETED | CRITICAL
1984 | 2024-08-01 | Aug 2024 | 9:00 | 9:04 | 0.0027778 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | SODA LIME QTY ADJUSTED AND CLEANED MOISTURE. ISSUE RECTIFIED | NO | NO | 2024-08-01 | August 2024 | 9:30 | 2024-08-01 09:00 | 2024-08-01 09:30 | 0:30 | COMPLETED | CRITICAL
1985 | 2024-08-01 | Aug 2024 | 9:35 | 9:40 | 0.0034722 | CT MACHINE | PHILIPS | INGENUITY | SOORAJ | RADIOLOGY | RING ARTEFACTS SEEN IN SCANS | PERFORMED AIR CALIBRATION AND ISSUE RECTIFIED | NO | NO | 2024-08-01 | August 2024 | 10:15 | 2024-08-01 09:35 | 2024-08-01 10:15 | 0:40 | COMPLETED | CRITICAL
1986 | 2024-08-01 | Aug 2024 | 11:00 | 11:05 | 0.0034722 | CAMERA SYSTEM | KARL STORZ | TC200 | SOORAJ | OT | LAP TOWER MENU DISABLED | FOUND THE LINK CABLE LOOSE CONTACT. ADJUSTED THE CABLE AND ISSUE RECTIFIED. | NO | NO | 2024-08-01 | August 2024 | 12:00 | 2024-08-01 11:00 | 2024-08-01 12:00 | 1:00 | COMPLETED | SEMI-CRITICAL
1987 | 2024-08-01 | Aug 2024 | 13:20 | 13:25 | 0.0034722 | BATTERY CHARGING UNIT | STRYKER | 7110-120-000 | SOORAJ | OT | DRILL NOT WORKING | FOUND BATTERY TO BE DEFECTIVE | NO | BATTERY | 2024-08-01 | August 2024 | 13:40 | 2024-08-01 13:20 | 2024-08-01 13:40 | 0:20 | PENDING | SEMI-CRITICAL
1988 | 2024-08-01 | Aug 2024 | 17:00 | 17:02 | 0.0013889 | PATIENT MONITOR | PHILIPS | MX 450 | HAFIZ | ENDOSCOPY | CAN'T ABLE TO MEASURE ETCO2 | FOUND WATER TRAP BROKEN, REPLACED IT AND OBSERVED. NOW WORKING GOOD. | WATER TRAP | NO | 2024-08-01 | August 2024 | 17:15 | 2024-08-01 17:00 | 2024-08-01 17:15 | 0:15 | COMPLETED | SEMI-CRITICAL
1989 | 2024-08-01 | Aug 2024 | 14:30 | 14:35 | 0.0034722 | BLOOD GAS ANALYSER | GEM 3500 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | PCO2 ISSUE | ISSUE RECTIFIED | NO | NO | 2024-08-01 | August 2024 | 14:40 | 2024-08-01 14:30 | 2024-08-01 14:40 | 0:10 | COMPLETED | SEMI-CRITICAL
1990 | 2024-08-01 | Aug 2024 | 13:40 | 13:45 | 0.0034722 | STEAM STERILIZER | GETINGE | GSS67H102E | SANTHOSH | CSSD | STEAM TIME OUT | WATERSETTLED ISSUE RECTIFIED | NO | NO | 2024-08-01 | August 2024 | 13:50 | 2024-08-01 13:40 | 2024-08-01 13:50 | 0:10 | COMPLETED | SEMI-CRITICAL
1991 | 2024-08-01 | Aug 2024 | 10:30 | 10:50 | 0.013889 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | RAMKIE | OPD | NOT WORKING | CHECKED AND NOT ISSUES FOUND, KEPT UNDER OBSERVATION | NO | NO | 2024-08-01 | August 2024 | 11:05 | 2024-08-01 10:30 | 2024-08-01 11:05 | 0:35 | COMPLETED | SEMI-CRITICAL
1992 | 2024-08-01 | Aug 2024 | 9:10 | 9:15 | 0.0034722 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 2000 | THANISH | OPD | DISPLAY NOT WORKING | CLEANED THE DISPLAY CABLE AND RECONNECTED IT.WORKING FINE. | NO | NO | 2024-08-01 | August 2024 | 9:40 | 2024-08-01 09:10 | 2024-08-01 09:40 | 0:30 | COMPLETED | SEMI-CRITICAL
1993 | 2024-08-02 | Aug 2024 | 7:10 | 7:15 | 0.0034722 | CT MACHINE | PHILIPS | INGENUITY | SOORAJ | RADIOLOGY | GANTRY SHUTDOWN | RESTARTED THE SYSTEM THRICE AND ISSUE RECTIFIED. | NO | NO | 2024-08-02 | August 2024 | 7:55 | 2024-08-02 07:10 | 2024-08-02 07:55 | 0:45 | COMPLETED | CRITICAL
1994 | 2024-08-02 | Aug 2024 | 8:00 | 8:05 | 0.0034722 | HEATER COOLER | GETINGE | HCU40 | SOORAJ | OT | HEMOTHERM BLANKET LEAK. | LEAK ARRESTED AND ISSUE RECTIFIED. | NO | NO | 2024-08-02 | August 2024 | 9:00 | 2024-08-02 08:00 | 2024-08-02 09:00 | 1:00 | COMPLETED | CRITICAL
1995 | 2024-08-02 | Aug 2024 | 9:40 | 9:45 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SOORAJ | ICU | INCORRECT DATE AND TIME SETTING | REBOOTED AND CONFIGURED THE SYSTEM. ISSUE RECTIFIED. | NO | NO | 2024-08-02 | August 2024 | 11:30 | 2024-08-02 09:40 | 2024-08-02 11:30 | 1:50 | COMPLETED | SEMI-CRITICAL
1996 | 2024-08-02 | Aug 2024 | 11:25 | 11:35 | 0.0069444 | CAMERA SYSTEM | KARL STORZ | TC200 | SOORAJ | OT | LAP TOWER MENU DISABLED | FOUND THE LINK CABLE LOOSE CONTACT. ADJUSTED THE CABLE AND ISSUE RECTIFIED. | NO | NO | 2024-08-02 | August 2024 | 12:00 | 2024-08-02 11:25 | 2024-08-02 12:00 | 0:35 | COMPLETED | SEMI-CRITICAL
1997 | 2024-08-02 | Aug 2024 | 12:30 | 12:35 | 0.0034722 | CAMERA SYSTEM | KARL STORZ | TC200 | SOORAJ | OT | YELLOW COLOUR ON DISPLAY | FOUND DIRT ON CAMERA LENS. CLEANED THE LENS AND ISSUE RECTIFIED. | NO | NO | 2024-08-02 | August 2024 | 13:00 | 2024-08-02 12:30 | 2024-08-02 13:00 | 0:30 | COMPLETED | SEMI-CRITICAL
1998 | 2024-08-02 | Aug 2024 | 17:00 | 17:05 | 0.0034722 | C-MAC MONITOR | KARL STORZ | 84032ZX | SOORAJ | OT | VIDEO LINK CABLE NOT WORKING | VIDEO CABLE SERVICE DONE AND ISSUE RECTIFIED. | NO | NO | 2024-08-02 | August 2024 | 17:20 | 2024-08-02 17:00 | 2024-08-02 17:20 | 0:20 | COMPLETED | SEMI-CRITICAL
1999 | 2024-08-02 | Aug 2024 | 17:30 | 17:35 | 0.0034722 | TOURNIQUET,DUAL CUFF | VBM MEDZINTECHNIK | 9000 | SOORAJ | OT | LEAK ALARM | FOUND ISSUE WITH 24" BP CUFF. MACHINE CHECKED AND WORKING FINE | NO | NO | 2024-08-02 | August 2024 | 18:00 | 2024-08-02 17:30 | 2024-08-02 18:00 | 0:30 | COMPLETED | SEMI-CRITICAL
2000 | 2024-08-02 | Aug 2024 | 10:00 | 10:03 | 0.0020833 | RADIOGRAPHIC UNIT | CARESTREAM | DRF ASCEND ANOLOG HF | HAFIZ | RADIOLOGY | DETECTOR BATTERY ISSUE | TWO DETECTOR BATTERIES ISSUED AND OBSERVED. MODEL: DRX-1 SYSTEM 1 AND S.NO: 492200031, 492200172 | DETECTOR BATTERY | NO | 2024-08-02 | August 2024 | 10:15 | 2024-08-02 10:00 | 2024-08-02 10:15 | 0:15 | COMPLETED | SEMI-CRITICAL
2001 | 2024-08-02 | Aug 2024 | 17:00 | 17:04 | 0.0027778 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | HAFIZ | DIALYSIS | BRAKE ISSUE - BROKEN | BRAKE REFIXED INTO THE COT AND OBSERVED. WORKING GOOD. | NO | NO | 2024-08-02 | August 2024 | 17:20 | 2024-08-02 17:00 | 2024-08-02 17:20 | 0:20 | COMPLETED | NON-CRITICAL
2002 | 2024-08-02 | Aug 2024 | 17:00 | 17:05 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | DATE AND TIME ISSUE | ISSUE RECTIFIED | NO | NO | 2024-08-02 | August 2024 | 17:10 | 2024-08-02 17:00 | 2024-08-02 17:10 | 0:10 | COMPLETED | SEMI-CRITICAL
2003 | 2024-08-02 | Aug 2024 | 11:30 | 11:40 | 0.0069444 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | RAMKIE | OT | NOT WORKING | CHECKED AND WIRE SOLDERING | NO | NO | 2024-08-02 | August 2024 | 12:00 | 2024-08-02 11:30 | 2024-08-02 12:00 | 0:30 | COMPLETED | SEMI-CRITICAL
2004 | 2024-08-03 | Aug 2024 | 7:50 | 7:55 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | REDUCED SODA LIME QTY AND ISSUE RECTIFIED | NO | NO | 2024-08-03 | August 2024 | 8:30 | 2024-08-03 07:50 | 2024-08-03 08:30 | 0:40 | COMPLETED | CRITICAL
2005 | 2024-08-03 | Aug 2024 | 15:40 | 15:45 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | CARESTATION 650 | SOORAJ | OT | AC POWER FAIL ISSUE | CLEANED THE POWER MANOGEMENT BOARD AND RELOADED SOFTWARE. ISSUE RECTIFIED. | NO | NO | 2024-08-03 | August 2024 | 16:30 | 2024-08-03 15:40 | 2024-08-03 16:30 | 0:50 | COMPLETED | CRITICAL
2006 | 2024-08-05 | Aug 2024 | 8:00 | 8:06 | 0.0041667 | INFUSION PUMP | SMITHS MEDICAL | GRASEBY 1200 | HAFIZ | ICU | NOT WORKING | REPLACED THE HOLDER CLAMP, AND PRESSURE SENSOR. NOW WORKING GOOD. | HOLDER CLAMP, PRESSURE SENSOR | NO | 2024-08-05 | August 2024 | 8:20 | 2024-08-05 08:00 | 2024-08-05 08:20 | 0:20 | COMPLETED | NON-CRITICAL
2007 | 2024-08-05 | Aug 2024 | 10:00 | 10:05 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC E-R7 | HAFIZ | ICU | ECHO PROBE NOT WORKING | REPLACED 3-SC ECHO PROBE (S.NO: 1332067WX6), AND OBSERVED. NOW WORKING GOOD. | 3SC ECHO PROBE | NO | 2024-08-05 | August 2024 | 10:30 | 2024-08-05 10:00 | 2024-08-05 10:30 | 0:30 | COMPLETED | SEMI-CRITICAL
2008 | 2024-08-05 | Aug 2024 | 12:20 | 12:21 | 0.00069444 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | HAFIZ | OT | NOT WORKING | PROPERLY FIXED THE HOSE AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-08-05 | August 2024 | 12:35 | 2024-08-05 12:20 | 2024-08-05 12:35 | 0:15 | COMPLETED | SEMI-CRITICAL
2009 | 2024-08-05 | Aug 2024 | 17:00 | 17:02 | 0.0013889 | ANKLE MOTION DEVICE | ANKLE MOTION PRIVATE LIMITED | AM-03 | HAFIZ | PHYSIOTHERAPY | NOT WORKING | ISSUE WITH ADAPTORS, SO REPLACED 2 ADAPTOR. NOW ISSUE RECTIFIED. | ANKLE MOTION ADAPTOR | NO | 2024-08-05 | August 2024 | 17:30 | 2024-08-05 17:00 | 2024-08-05 17:30 | 0:30 | COMPLETED | SEMI-CRITICAL
2010 | 2024-08-05 | Aug 2024 | 12:30 | 12:35 | 0.0034722 | USG MACHINE | WIPRO GE HEALTH CARE | LOGIC V2 | JERLIN | EMERGENCY | ARTIFACTS IN ECHO IMAGING | CHANGED THE ECHO PROBE AND RECTIFIED THEISSUE | ADULT ECHO PROBE | NO | 2024-08-05 | August 2024 | 15:35 | 2024-08-05 12:30 | 2024-08-05 15:35 | 3:05 | COMPLETED | SEMI-CRITICAL
2011 | 2024-08-06 | Aug 2024 | 6:15 | 6:20 | 0.0034722 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | SOORAJ | ICU | ERROR TE:27 | RESTARTED THE SYSTEM AND RAN PRE USE CHECK. ISSUE RECTIFIED. | NO | NO | 2024-08-06 | August 2024 | 6:45 | 2024-08-06 06:15 | 2024-08-06 06:45 | 0:30 | COMPLETED | CRITICAL
2012 | 2024-08-06 | Aug 2024 | 21:28 | 21:30 | 0.0013889 | BIPAP MACHINE | PHILIPS | A40 | SOORAJ | ICU | TIDAL VOLUME NOT SHOWING | SETTINGS CONFIGURED AND ISSUE RECTIFIED. | NO | NO | 2024-08-06 | August 2024 | 21:45 | 2024-08-06 21:28 | 2024-08-06 21:45 | 0:17 | COMPLETED | SEMI-CRITICAL
2013 | 2024-08-06 | Aug 2024 | 23:30 | 23:40 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | SOORAJ | WARD | COT FUNCTION NOT WORKING | POWER SOCKET DEFECTIVE. CCONNECTED TO ANOTHER POWER SOCKET AND ISSUE RECTIFIED. | NO | NO | 2024-08-06 | August 2024 | 23:50 | 2024-08-06 23:30 | 2024-08-06 23:50 | 0:20 | COMPLETED | NON-CRITICAL
2014 | 2024-08-06 | Aug 2024 | 18:00 | 18:03 | 0.0020833 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | HAFIZ | WARD | NOT WORKING | RESET DONE, AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-08-06 | August 2024 | 18:10 | 2024-08-06 18:00 | 2024-08-06 18:10 | 0:10 | COMPLETED | NON-CRITICAL
2015 | 2024-08-06 | Aug 2024 | 9:20 | 9:27 | 0.0048611 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX | SANTHOSH | OT | ERROR-198 | RESET DONE ,NOW WORKING | NO | NO | 2024-08-06 | August 2024 | 9:30 | 2024-08-06 09:20 | 2024-08-06 09:30 | 0:10 | COMPLETED | SEMI-CRITICAL
2016 | 2024-08-06 | Aug 2024 | 12:20 | 12:25 | 0.0034722 | FLUID WARMER SYSTEM | MATRIX | MN05 | SANTHOSH | OT | BLANKET LEAK | CLEARED DATA NOW SAVING IMAGE | NO | NO | 2024-08-06 | August 2024 | 12:30 | 2024-08-06 12:20 | 2024-08-06 12:30 | 0:10 | COMPLETED | SEMI-CRITICAL
2017 | 2024-08-06 | Aug 2024 | 9:45 | 10:00 | 0.010417 | PATIENT MONITOR | PHILIPS | CM12 | RAMKIE | OT | ECG NOT WORKING | FOUND THE ECG CABLE NOT WORKING | NO | NO | 2024-08-06 | August 2024 | 10:20 | 2024-08-06 09:45 | 2024-08-06 10:20 | 0:35 | COMPLETED | SEMI-CRITICAL
2018 | 2024-08-07 | Aug 2024 | 0:15 | 0:20 | 0.0034722 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | SOORAJ | WARD | SPO2 FUNCTION NOT WORKING. | SPO2 SENSOR CONNECTOR PIN BENT. STRAIGHTENED PIN AND ISSUE RECTIFIED. | NO | NO | 2024-08-07 | August 2024 | 0:45 | 2024-08-07 00:15 | 2024-08-07 00:45 | 0:30 | COMPLETED | SEMI-CRITICAL
2019 | 2024-08-07 | Aug 2024 | 1:45 | 1:50 | 0.0034722 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | SOORAJ | WARD | BP CUFF NOT WORKING | FOUND LEAK WITH CUFF. REPLACED CUFF AND ISSUE RECTIFIED | NO | NO | 2024-08-07 | August 2024 | 2:15 | 2024-08-07 01:45 | 2024-08-07 02:15 | 0:30 | COMPLETED | SEMI-CRITICAL
2020 | 2024-08-07 | Aug 2024 | 21:35 | 21:40 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | SOORAJ | WARD | THERAPY DISABLED ALARM | RAN PRE OP CHECK AND ISSUE RECTIFIED. | NO | NO | 2024-08-07 | August 2024 | 21:55 | 2024-08-07 21:35 | 2024-08-07 21:55 | 0:20 | COMPLETED | CRITICAL
2021 | 2024-08-07 | Aug 2024 | 10:30 | 10:40 | 0.0069444 | MOVEMENT THERAPHY UNIT, UPPER AND LOWER LIMBS | RECK | MOTOMED LETTO2 | JERLIN | PHYSIOTHERAPY | BRAKE ISSUE | FIXED AND RECTIFIED THE ISSUE | NO | NO | 2024-08-07 | August 2024 | 10:55 | 2024-08-07 10:30 | 2024-08-07 10:55 | 0:25 | COMPLETED | SEMI-CRITICAL
2022 | 2024-08-08 | Aug 2024 | 5:28 | 5:35 | 0.0048611 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | SOORAJ | WARD | COT FUNCTIONS NOT WORKING | COT RESET DONE AND ISSUE RECTIFIED. | NO | NO | 2024-08-08 | August 2024 | 6:30 | 2024-08-08 05:28 | 2024-08-08 06:30 | 1:02 | COMPLETED | NON-CRITICAL
2023 | 2024-08-08 | Aug 2024 | 20:45 | 20:50 | 0.0034722 | IMMUNOLOGY ANALYSER | COBAS E 601 | ROCHE | SOORAJ | LABORATORY | INTERMITTENT STANDBY MODE | REMOVED AND CLEANED THE REAGENT TRAY. OPEN CAP CHECK PROCESS RAN AND PASSED. ISSUE RECTIFIED. | NO | NO | 2024-08-08 | August 2024 | 21:30 | 2024-08-08 20:45 | 2024-08-08 21:30 | 0:45 | COMPLETED | SEMI-CRITICAL
2024 | 2024-08-09 | Aug 2024 | 22:00 | 22:05 | 0.0034722 | DIALYSIS MACHINE | FRESENIUS | 4008S NG | SOORAJ | DIALYSIS | FLOW ALARM | CLEANED MOTOR AND RAN CHECKS, ISSUE RECTIFIED | NO | NO | 2024-08-09 | August 2024 | 22:30 | 2024-08-09 22:00 | 2024-08-09 22:30 | 0:30 | COMPLETED | CRITICAL
2025 | 2024-08-09 | Aug 2024 | 22:30 | 22:35 | 0.0034722 | DIALYSIS MACHINE | FRESENIUS | 4008S NG | SOORAJ | DIALYSIS | HEPARIN ALARM | BEARING NEEDS REPLACEMENT BUT MACHINE IS WORKING FINE. | NO | NO | 2024-08-09 | August 2024 | 23:00 | 2024-08-09 22:30 | 2024-08-09 23:00 | 0:30 | COMPLETED | CRITICAL
2026 | 2024-08-09 | Aug 2024 | 23:00 | 23:00 | 0 | DIALYSIS MACHINE | FRESENIUS | 4008S NG | SOORAJ | DIALYSIS | BLOOD LEAK ALARM | ISSUE RECTIFIED BY ADJUSTING VOLTAGE | NO | NO | 2024-08-09 | August 2024 | 23:20 | 2024-08-09 23:00 | 2024-08-09 23:20 | 0:20 | COMPLETED | CRITICAL
2027 | 2024-08-09 | Aug 2024 | 23:20 | 23:20 | 0 | DIALYSIS MACHINE | FRESENIUS | 4008 S | SOORAJ | DIALYSIS | BLOOD LEAK ALARM | BLEACH CALIBRATION DONE AND VOLTAGE ADJUSTED, ISSUE RECTIFIED | NO | NO | 2024-08-10 | August 2024 | 0:00 | 2024-08-09 23:20 | 2024-08-10 | 0:40 | COMPLETED | CRITICAL
2028 | 2024-08-09 | Aug 2024 | 0:00 | 0:00 | 0 | DIALYSIS MACHINE | FRESENIUS | 4008 S NG | SOORAJ | DIALYSIS | BP NOT WORKING | BP CUFF COONECTOR BROKEN. TEMPORARILY RECTIFIED THE ISSUE | NO | NO | 2024-08-09 | August 2024 | 0:40 | 2024-08-09 | 2024-08-09 00:40 | 0:40 | COMPLETED | CRITICAL
2029 | 2024-08-09 | Aug 2024 | 0:40 | 0:40 | 0 | DIALYSIS MACHINE | FRESENIUS | 4008 S NG | SOORAJ | DIALYSIS | BP NOT WORKING | BP CUFF COONECTOR BROKEN. TEMPORARILY RECTIFIED THE ISSUE | NO | NO | 2024-08-09 | August 2024 | 1:20 | 2024-08-09 00:40 | 2024-08-09 01:20 | 0:40 | COMPLETED | CRITICAL
2030 | 2024-08-09 | Aug 2024 | 18:00 | 18:03 | 0.0020833 | DIALYSIS MACHINE | FRESENIUS | 4008 S | HAFIZ | DIALYSIS | BP CUFFS NOT WORKING | ISSUED 4 NOS BP CUFFS TO DIALYSIS. NOW WORKING GOOD. | BP CUFF | NO | 2024-08-09 | August 2024 | 18:10 | 2024-08-09 18:00 | 2024-08-09 18:10 | 0:10 | COMPLETED | CRITICAL
2031 | 2024-08-09 | Aug 2024 | 9:30 | 9:40 | 0.0069444 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | JERLIN | WARD | REMOTE NOT WORKING | REPLACED THE REMOTE AND RECTIFIED THE ISSUE | REMOTE | NO | 2024-08-09 | August 2024 | 9:54 | 2024-08-09 09:30 | 2024-08-09 09:54 | 0:24 | COMPLETED | NON-CRITICAL
2032 | 2024-08-09 | Aug 2024 | 10:35 | 10:40 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SANTHOSH | OT | CIRCUIT LEAK TEST FAILED | TIE WRAPPED | NO | NO | 2024-08-09 | August 2024 | 10:55 | 2024-08-09 10:35 | 2024-08-09 10:55 | 0:20 | COMPLETED | CRITICAL
2033 | 2024-08-10 | Aug 2024 | 10:00 | 10:04 | 0.0027778 | AUTOMATED ELISA READER | ROCHE | COBAS E411 | HAFIZ | LABORATORY | ERROR, NOT WORKING | FOUND NEEDLE TIP BEHIND THE TRAY, CLEARED IT AND OBSERVED. WORKING GOOD. | NO | NO | 2024-08-10 | August 2024 | 10:20 | 2024-08-10 10:00 | 2024-08-10 10:20 | 0:20 | COMPLETED | SEMI-CRITICAL
2034 | 2024-08-10 | Aug 2024 | 11:00 | 11:04 | 0.0027778 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 4000 | HAFIZ | ICU | NEED TO REPLACE CARTRIDGE | CARTRIDGE LOADED AND CVP TESTS DONE. NOW WORKING GOOD. | CARTRIDGE | NO | 2024-08-10 | August 2024 | 12:00 | 2024-08-10 11:00 | 2024-08-10 12:00 | 1:00 | COMPLETED | SEMI-CRITICAL
2035 | 2024-08-10 | Aug 2024 | 10:45 | 11:00 | 0.010417 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | RAMKIE | OT | NOT WORKING | CHECKED AND RESET DONE | NO | NO | 2024-08-10 | August 2024 | 11:20 | 2024-08-10 10:45 | 2024-08-10 11:20 | 0:35 | COMPLETED | NON-CRITICAL
2036 | 2024-08-11 | Aug 2024 | 18:30 | 18:35 | 0.0034722 | NITRIC OXIDE DELIVERY SYSTEM | WIPRO GE HEALTH CARE | SLE | SANTHOSH | ICU | CALIBRATION DUE | CALIBRATION DONE | NO | NO | 2024-08-11 | August 2024 | 18:55 | 2024-08-11 18:30 | 2024-08-11 18:55 | 0:25 | COMPLETED | CRITICAL
2037 | 2024-08-12 | Aug 2024 | 7:00 | 7:02 | 0.0013889 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | CLEANED THE ABSORBER CANISTER AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-08-12 | August 2024 | 7:20 | 2024-08-12 07:00 | 2024-08-12 07:20 | 0:20 | COMPLETED | CRITICAL
2038 | 2024-08-12 | Aug 2024 | 7:20 | 7:22 | 0.0013889 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | HAFIZ | OT | CIRCUIT LEAK TEST FAILED | CLEANED THE ABSORBER CANISTER AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-08-12 | August 2024 | 7:40 | 2024-08-12 07:20 | 2024-08-12 07:40 | 0:20 | COMPLETED | CRITICAL
2039 | 2024-08-12 | Aug 2024 | 8:35 | 8:37 | 0.0013889 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | HAFIZ | WARD | NOT WORKING | ISSUE WITH BULB ELECTRODES. CLEANED AND REFIXED, NOW WORKING GOOD. | NO | NO | 2024-08-12 | August 2024 | 8:50 | 2024-08-12 08:35 | 2024-08-12 08:50 | 0:15 | COMPLETED | SEMI-CRITICAL
2040 | 2024-08-12 | Aug 2024 | 15:00 | 15:05 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 8000 | JERLIN | WARD | NOT WORKING | REPLACED THE HEAD END ACTUATOR AND RECTIFIED THE ISSUE | ACTUATOR | NO | 2024-08-12 | August 2024 | 15:20 | 2024-08-12 15:00 | 2024-08-12 15:20 | 0:20 | COMPLETED | NON-CRITICAL
2041 | 2024-08-12 | Aug 2024 | 10:45 | 10:50 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | WARD | CATRIDGE DEFECT | ISSUE RECTIFIED | NO | NO | 2024-08-12 | August 2024 | 11:00 | 2024-08-12 10:45 | 2024-08-12 11:00 | 0:15 | COMPLETED | SEMI-CRITICAL
2042 | 2024-08-13 | Aug 2024 | 2:15 | 2:20 | 0.0034722 | PATIENT MONITOR | PHILIPS | MX 450 | SANTHOSH | OPD | PRESSURE TRANSDUCER FAILURE | REPLACED NEW TRANSDUCER | TRANSDUCER | NO | 2024-08-13 | August 2024 | 2:30 | 2024-08-13 02:15 | 2024-08-13 02:30 | 0:15 | COMPLETED | SEMI-CRITICAL
2043 | 2024-08-14 | Aug 2024 | 14:30 | 14:32 | 0.0013889 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | HAFIZ | OT | ANGULATION NOT WORKING | REMOVED, RECONNECTED THE PROBE, OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-08-14 | August 2024 | 14:45 | 2024-08-14 14:30 | 2024-08-14 14:45 | 0:15 | COMPLETED | SEMI-CRITICAL
2044 | 2024-08-14 | Aug 2024 | 7:40 | 7:42 | 0.0013889 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | HAFIZ | OT | NOT WORKING | SOLDERED THE WIRES AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-08-14 | August 2024 | 8:00 | 2024-08-14 07:40 | 2024-08-14 08:00 | 0:20 | COMPLETED | SEMI-CRITICAL
2045 | 2024-08-14 | Aug 2024 | 8:00 | 8:03 | 0.0020833 | PATIENT MONITOR | PHILIPS | CM12 | HAFIZ | WARD | SPO2 NOT WORKING | RFIXED IT PROPERLY AND OPBSERVED. NOW WORKING GOOD. | NO | NO | 2024-08-14 | August 2024 | 8:15 | 2024-08-14 08:00 | 2024-08-14 08:15 | 0:15 | COMPLETED | SEMI-CRITICAL
2046 | 2024-08-14 | Aug 2024 | 8:30 | 8:32 | 0.0013889 | NEBULIZER | PARI | PARI BOY SX | HAFIZ | WARD | NEBULIZER NOT WORKING | FIXED THE CONNECTOR AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-08-14 | August 2024 | 8:45 | 2024-08-14 08:30 | 2024-08-14 08:45 | 0:15 | COMPLETED | SEMI-CRITICAL
2047 | 2024-08-14 | Aug 2024 | 0:45 | 0:50 | 0.0034722 | DIGITAL PLATFORM SCALE | SECA | 676 | SANTHOSH | DIALYSIS | ERROR 1.16 | ADJUST SETTIGS ISSUE RECTIFIED | NO | NO | 2024-08-14 | August 2024 | 1:00 | 2024-08-14 00:45 | 2024-08-14 01:00 | 0:15 | COMPLETED | NON-CRITICAL
2048 | 2024-08-14 | Aug 2024 | 11:10 | 11:20 | 0.0069444 | DEFIBRILLATOR MONITOR WITH AED | PHILIPS | EFFICIA DFM 100 | RAMKIE | OT | TEST NOT PASSED | CHECKED AND RECONNEDTED THE PADDLES | NO | NO | 2024-08-14 | August 2024 | 11:30 | 2024-08-14 11:10 | 2024-08-14 11:30 | 0:20 | COMPLETED | CRITICAL
2049 | 2024-08-15 | Aug 2024 | 12:15 | 12:20 | 0.0034722 | NEBULIZER | JK MEDICALS | ULTIMA | JERLIN | WARD | NOT WORKING PROPERLY | REPLACED THE MOTOR AND RECTIFIED THE ISSUE | MOTOR | NO | 2024-08-15 | August 2024 | 12:30 | 2024-08-15 12:15 | 2024-08-15 12:30 | 0:15 | COMPLETED | SEMI-CRITICAL
2050 | 2024-08-15 | Aug 2024 | 8:00 | 8:05 | 0.0034722 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | CVP FAILED | CVP PASSED ISSUE RECTIFIED | NO | NO | 2024-08-15 | August 2024 | 8:10 | 2024-08-15 08:00 | 2024-08-15 08:10 | 0:10 | COMPLETED | SEMI-CRITICAL
2051 | 2024-08-15 | Aug 2024 | 9:15 | 9:20 | 0.0034722 | HEATER COOLER | GETINGE | HCU40 | THANISH | OT | HEMOTHERM BLANKET LEAKAGE | PATCHWORK DONE FOR HEMOTHERM BLANKET | NO | NO | 2024-08-15 | August 2024 | 9:45 | 2024-08-15 09:15 | 2024-08-15 09:45 | 0:30 | COMPLETED | CRITICAL
2052 | 2024-08-16 | Aug 2024 | 11:30 | 11:40 | 0.0069444 | STEAM STERILIZER | GETINGE | GSS67H102E | RAMKIE | CSSD | ERROR | CHECKED AND FOUND WATER PRESSURE NOT SUFFICIENT | NO | NO | 2024-08-16 | August 2024 | 11:50 | 2024-08-16 11:30 | 2024-08-16 11:50 | 0:20 | COMPLETED | SEMI-CRITICAL
2053 | 2024-08-16 | Aug 2024 | 16:10 | 16:15 | 0.0034722 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | THANISH | WARD | SHOCK ABORTED ISSUE | RECTIFIED THE PADDLES.WORKING FINE. | NO | NO | 2024-08-16 | August 2024 | 16:50 | 2024-08-16 16:10 | 2024-08-16 16:50 | 0:40 | COMPLETED | CRITICAL
2054 | 2024-08-17 | Aug 2024 | 17:10 | 17:15 | 0.0034722 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 9000 | SOORAJ | ICU | COT LOCKED | RESET DONE AND ISSUE RECTIFIED | NO | NO | 2024-08-17 | August 2024 | 17:30 | 2024-08-17 17:10 | 2024-08-17 17:30 | 0:20 | COMPLETED | NON-CRITICAL
2055 | 2024-08-17 | Aug 2024 | 20:00 | 20:01 | 0.00069444 | BLOOD GAS ANALYSER | INSTRUMENTATION LABORATORY | GEM 3500 | HAFIZ | ICU | NEED TO REPLACE CARTRIDGE | CARTRIDGE LOADED AND IS READY TO USE. | CARTRIDGE | NO | 2024-08-17 | August 2024 | 20:40 | 2024-08-17 20:00 | 2024-08-17 20:40 | 0:40 | COMPLETED | SEMI-CRITICAL
2056 | 2024-08-17 | Aug 2024 | 13:30 | 13:40 | 0.0069444 | VENTILATOR, UNIVERSAL | MAQUET | SERVO-I | RAMKIE | ICU | TEST NOT PASSED | CHECKED AND CASSETTE RECONNECTED, DIAPHRAGM CLEANED AND PROBLEM CLEARED | NO | NO | 2024-08-17 | August 2024 | 13:50 | 2024-08-17 13:30 | 2024-08-17 13:50 | 0:20 | COMPLETED | CRITICAL
2057 | 2024-08-17 | Aug 2024 | 10:10 | 10:15 | 0.0034722 | EEG MACHINE | NOTUS NEUROLOGY | NICOLET EEG V32 | THANISH | OPD | EEG LAPTOP BOOTING ISSUE | REPLACED THE DATA HARD-DISK AND RE-INSTALLED THE SOFTWARE.WORKING FINE. | NO | NO | 2024-08-17 | August 2024 | 18:30 | 2024-08-17 10:10 | 2024-08-17 18:30 | 8:20 | COMPLETED | SEMI-CRITICAL
2058 | 2024-08-18 | Aug 2024 | 17:20 | 17:30 | 0.0069444 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | SOORAJ | ICU | O2 CONCENTATION HIGH ALARM | O2 CELL ADJUSTMENT DONE AND ISSUE RECTIFIED. | NO | NO | 2024-08-18 | August 2024 | 18:00 | 2024-08-18 17:20 | 2024-08-18 18:00 | 0:40 | COMPLETED | CRITICAL
2059 | 2024-08-18 | Aug 2024 | 6:49 | 6:51 | 0.0013889 | BIPAP MACHINE | PHILIPS | A40 | HAFIZ | WARD | SETTINGS TO BE CHNGED | SETTINGS CONFIGURED. | NO | NO | 2024-08-18 | August 2024 | 7:00 | 2024-08-18 06:49 | 2024-08-18 07:00 | 0:11 | COMPLETED | SEMI-CRITICAL
2060 | 2024-08-19 | Aug 2024 | 11:40 | 11:45 | 0.0034722 | ELECTROTHERAPHY WITH LASER | CHATTANOOGA | INTELECT ADVANCED | SANTHOSH | PHYSIOTHERAPY | POWER CORD BURNT | REPLACED NEW POWER CORD | POWERCORD | NO | 2024-08-19 | August 2024 | 23:50 | 2024-08-19 11:40 | 2024-08-19 23:50 | 12:10 | COMPLETED | SEMI-CRITICAL
2061 | 2024-08-19 | Aug 2024 | 12:30 | 12:35 | 0.0034722 | LABOUR BED | ARJO HUNTLEIGH | 170000E1L | SANTHOSH | WARD | NOT WORKING | ISSUE RECTIFIED | NO | NO | 2024-08-19 | August 2024 | 12:40 | 2024-08-19 12:30 | 2024-08-19 12:40 | 0:10 | COMPLETED | NON-CRITICAL
2062 | 2024-08-19 | Aug 2024 | 17:35 | 17:40 | 0.0034722 | PATIENT MONITOR | PHILIPS | CM12 | SANTHOSH | WARD | NOT WORKING | ISSUE RECTIFIED | NO | NO | 2024-08-19 | August 2024 | 17:50 | 2024-08-19 17:35 | 2024-08-19 17:50 | 0:15 | COMPLETED | SEMI-CRITICAL
2063 | 2024-08-19 | Aug 2024 | 11:50 | 11:55 | 0.0034722 | PORTABLE X-RAY MACHINE | SKANRAY | SKANMOBILE | SANTHOSH | RADIOLOGY | UPS BATTERY BACKUP ISSUE | CLEARED DATA NOW WORKING GOOD | NO | NO | 2024-08-19 | August 2024 | 12:05 | 2024-08-19 11:50 | 2024-08-19 12:05 | 0:15 | COMPLETED | SEMI-CRITICAL
2064 | 2024-08-19 | Aug 2024 | 14:20 | 14:40 | 0.013889 | DEFIBRILLATOR MONITOR | PHILIPS | EFFICIA DFM 100 | RAMKIE | ICU | TEST NOT PASSED | CHECKED AND RECCONNECTED ECG CABLEM PROBLEM RECTIFIED | NO | NO | 2024-08-19 | August 2024 | 14:50 | 2024-08-19 14:20 | 2024-08-19 14:50 | 0:30 | COMPLETED | CRITICAL
2065 | 2024-08-19 | Aug 2024 | 14:10 | 14:15 | 0.0034722 | RADIANT HEAT WARMER | WIPRO GE HEALTH CARE | LULLABY | THANISH | ICU | TEMPERATURE PROBE FAILED | REPLACED WITH NEW TEMPERATURE PROBE.WORKING FINE. | SKIN TEMPERATURE PROBE | NO | 2024-08-19 | August 2024 | 14:30 | 2024-08-19 14:10 | 2024-08-19 14:30 | 0:20 | COMPLETED | SEMI-CRITICAL
2066 | 2024-08-20 | Aug 2024 | 16:00 | 16:05 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | SOORAJ | ICU | NOT SWITCHING ON | ISSUED WITH THREE BOARDS. REPLACED AND ISSUE RECTIFIED. | I/O BOARD MST BOARD CPU BOARD | NO | 2024-08-20 | August 2024 | 17:00 | 2024-08-20 16:00 | 2024-08-20 17:00 | 1:00 | COMPLETED | SEMI-CRITICAL
2067 | 2024-08-20 | Aug 2024 | 12:30 | 12:35 | 0.0034722 | MICROSCOPE OPERATING, SURGICAL | CARL ZEISS | OPMI SENSARA | JERLIN | OT | LAMP HOURS CROSSED AND INTENSITY IS LOW | REPLACED ANEW LAMP AND RECTIFIED THE ISSUE | XENON LAMP | NO | 2024-08-20 | August 2024 | 12:50 | 2024-08-20 12:30 | 2024-08-20 12:50 | 0:20 | COMPLETED | SEMI-CRITICAL
2068 | 2024-08-20 | Aug 2024 | 11:30 | 11:35 | 0.0034722 | TRANSPORT VENTILATOR | RESMED | ASTRAL 100 | THANISH | EMERGENCY | BIPAP MACHINE NOT WORKING | FACTORY RESET DONE.WORKING FINE | NO | NO | 2024-08-20 | August 2024 | 11:55 | 2024-08-20 11:30 | 2024-08-20 11:55 | 0:25 | COMPLETED | SEMI-CRITICAL
2069 | 2024-08-21 | Aug 2024 | 11:30 | 11:40 | 0.0069444 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | AISYS CS2 | JERLIN | OT | CHECKOUT FAILED | CHANGED THE GAS MIXER ASSEMBLY AND RECTIFIED THE ISSUE | GAS MIXER ASSEMBLY | NO | 2024-08-21 | August 2024 | 12:50 | 2024-08-21 11:30 | 2024-08-21 12:50 | 1:20 | COMPLETED | CRITICAL
2070 | 2024-08-21 | Aug 2024 | 15:00 | 15:15 | 0.010417 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | JERLIN | OPD | NOT ABLE TO ACQUIRE ECG WAVES | FIXED THE ECG CABLE PROPERLY AD RECTIFIED THE ISSUE | NO | NO | 2024-08-21 | August 2024 | 15:25 | 2024-08-21 15:00 | 2024-08-21 15:25 | 0:25 | COMPLETED | SEMI-CRITICAL
2071 | 2024-08-21 | Aug 2024 | 9:10 | 9:20 | 0.0069444 | BMD MACHINE | HOLOGIC | ASY-05119 | JERLIN | RADIOLOGY | CONSOLE SYSTEM NOT WORKING | RESTATED THE CPU AND RECTIFIED THE ISSUE | NO | NO | 2024-08-21 | August 2024 | 9:30 | 2024-08-21 09:10 | 2024-08-21 09:30 | 0:20 | COMPLETED | SEMI-CRITICAL
2072 | 2024-08-21 | Aug 2024 | 14:00 | 14:05 | 0.0034722 | BMI MACHINE | AKAS MEDICAL | QVS-100 | SANTHOSH | OPD | POWER CORD ISSUE | ISSUE RECTIFIED | NO | NO | 2024-08-21 | August 2024 | 14:20 | 2024-08-21 14:00 | 2024-08-21 14:20 | 0:20 | COMPLETED | NON-CRITICAL
2073 | 2024-08-21 | Aug 2024 | 10:10 | 10:15 | 0.0034722 | PULMONORY FUNCTION TESTING MACHINE | NDD | EASY ONE PRO | THANISH | PHYSIOTHERAPY | PFT NOT WORKING | DELETED THE OLD DATA AND RECONFIGURED IT.WORKING FINE. | NO | NO | 2024-08-21 | August 2024 | 10:45 | 2024-08-21 10:10 | 2024-08-21 10:45 | 0:35 | COMPLETED | SEMI-CRITICAL
2074 | 2024-08-22 | Aug 2024 | 10:00 | 10:10 | 0.0069444 | RADIANT HEAT WARMER | WIPRO GE HEALTH CARE | LULLABY | JERLIN | ICU | TEMP PROBE NOT WORKING | RPLACED THE TEMPERATURE PROBE AND RECTIFIED THE ISSUE | SKIN TEMPERATURE PROBE | NO | 2024-08-22 | August 2024 | 10:15 | 2024-08-22 10:00 | 2024-08-22 10:15 | 0:15 | COMPLETED | SEMI-CRITICAL
2075 | 2024-08-22 | Aug 2024 | 11:15 | 11:20 | 0.0034722 | WARMING UNIT,PATIENT CARE | SMITHS MEDICAL | EQ 5000 | JERLIN | OT | NOT WORKING | CONNECTED THE TEMP SENSOR PROPERLY AND RECTIFIED THE ISSUE. | NO | NO | 2024-08-22 | August 2024 | 11:30 | 2024-08-22 11:15 | 2024-08-22 11:30 | 0:15 | COMPLETED | SEMI-CRITICAL
2076 | 2024-08-22 | Aug 2024 | 15:20 | 15:30 | 0.0069444 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID E95 | RAMKIE | OPD | HANGING PROBLEM | COMPANY PERSON UPDATED SAFTWARE AND PROBLEM RECTIFIED | NO | NO | 2024-08-22 | August 2024 | 15:55 | 2024-08-22 15:20 | 2024-08-22 15:55 | 0:35 | COMPLETED | SEMI-CRITICAL
2077 | 2024-08-22 | Aug 2024 | 11:10 | 11:15 | 0.0034722 | ECHO MACHINE | WIPRO GE HEALTH CARE | VIVID IQ | THANISH | RADIOLOGY | BOOTING ISSUE | REPLACED THE CPU,IO BOARD&MST BOARD AND REINSTALLED THE SOFTWARE.WORKING FINE. | CPU,IO BOARD&MST BOARD | NO | 2024-08-22 | August 2024 | 17:00 | 2024-08-22 11:10 | 2024-08-22 17:00 | 5:50 | COMPLETED | SEMI-CRITICAL
2078 | 2024-08-23 | Aug 2024 | 16:30 | 16:34 | 0.0027778 | RESUSCITATION UNIT | WIPRO GE HEALTH CARE | LULLABY RESUS PLUS | HAFIZ | ICU | FLOW CONTROL KNOB MIS-ALIGNED | REALIGNED AND OBSERVED. NOW WORKING GOOD. | NO | NO | 2024-08-23 | August 2024 | 17:00 | 2024-08-23 16:30 | 2024-08-23 17:00 | 0:30 | COMPLETED | SEMI-CRITICAL
2079 | 2024-08-23 | Aug 2024 | 11:00 | 11:00 | 0 | PATIENT MONITOR | PHILIPS | MX 450 | SANTHOSH | ICU | CO MODULE NOT WORKING | REPLACED ANOTHER MODULE | CO MODULE | NO | 2024-08-23 | August 2024 | 11:10 | 2024-08-23 11:00 | 2024-08-23 11:10 | 0:10 | COMPLETED | SEMI-CRITICAL
2080 | 2024-08-23 | Aug 2024 | 16:00 | 16:10 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | CVP 2 FAILED | ISSUE RECTIFIED | NO | NO | 2024-08-23 | August 2024 | 16:10 | 2024-08-23 16:00 | 2024-08-23 16:10 | 0:10 | COMPLETED | SEMI-CRITICAL
2081 | 2024-08-23 | Aug 2024 | 10:05 | 10:10 | 0.0034722 | PRINTER,COMPUTED RADIOGRAPHY | CARESTREAM | DRY VIEW 6950 | THANISH | RADIOLOGY | PRINTER NOT WORKING | FILM JAMMED IN LOWER TRAY AND REMOVED IT.WORKING FINE. | NO | NO | 2024-08-23 | August 2024 | 10:30 | 2024-08-23 10:05 | 2024-08-23 10:30 | 0:25 | COMPLETED | SEMI-CRITICAL
2082 | 2024-08-24 | Aug 2024 | 14:00 | 14:10 | 0.0069444 | ELECTRO SURGICAL UNIT | COVIDIEN | FORCE FX | JERLIN | OT | CONTINUOUS ALARM AFTER SWITCHING ON | RESET DONE AND RECTIFED THE ISSUE | NO | NO | 2024-08-24 | August 2024 | 14:20 | 2024-08-24 14:00 | 2024-08-24 14:20 | 0:20 | COMPLETED | SEMI-CRITICAL
2083 | 2024-08-24 | Aug 2024 | 13:30 | 13:40 | 0.0069444 | HUMIDIFIER AIR | FISHER AND PAYKEL | MR810 | JERLIN | ICU | NOT WORKING | TEMPERATURE CUTOFF RESET DONE AND RECTIFIED THE ISSUE | NO | NO | 2024-08-24 | August 2024 | 13:50 | 2024-08-24 13:30 | 2024-08-24 13:50 | 0:20 | COMPLETED | SEMI-CRITICAL
2084 | 2024-08-24 | Aug 2024 | 1:00 | 1:06 | 0.0041667 | OT TABLE | MAQUET | MEERA | JERLIN | OT | SLIDE FUNCTION NOT WORKING | LUBRICATED THE SLIDE RAILING AND RECTIFIED THE ISSUE | NO | NO | 2024-08-24 | August 2024 | 1:15 | 2024-08-24 01:00 | 2024-08-24 01:15 | 0:15 | COMPLETED | NON-CRITICAL
2085 | 2024-08-24 | Aug 2024 | 14:00 | 14:10 | 0.0069444 | ECG MACHINE | WIPRO GE HEALTH CARE | MAC 600 | JERLIN | WARD | NOT ABLE TO ACQUIRE ECG WAVES | FIXED THE ECG CABLE PROPERLY AD RECTIFIED THE ISSUE | NO | NO | 2024-08-24 | August 2024 | 14:20 | 2024-08-24 14:00 | 2024-08-24 14:20 | 0:20 | COMPLETED | SEMI-CRITICAL
2086 | 2024-08-24 | Aug 2024 | 19:00 | 19:10 | 0.0069444 | PATIENT MONITOR | MINDRAY MEDICAL INDIA | UMEC 12 | SANTHOSH | WARD | BP ISSUE | ISSUE RECTIFIED | NO | NO | 2024-08-24 | August 2024 | 19:10 | 2024-08-24 19:00 | 2024-08-24 19:10 | 0:10 | COMPLETED | SEMI-CRITICAL
2087 | 2024-08-24 | Aug 2024 | 16:20 | 16:40 | 0.013889 | MULTIFUNCTIONAL BED | ARJO HUNTLEIGH | ENTERPRISE 5000 | RAMKIE | WARD | NOT WORKING | RESET DONE AND PROBLEM RECTIFIED | NO | NO | 2024-08-24 | August 2024 | 16:55 | 2024-08-24 16:20 | 2024-08-24 16:55 | 0:35 | COMPLETED | NON-CRITICAL
2088 | 2024-08-24 | Aug 2024 | 19:30 | 19:35 | 0.0034722 | CT MACHINE | PHILIPS | INGENUITY | THANISH | RADIOLOGY | TABLE MOVEMENT AND SYSTEM HANGING ISSUE | REPLACED LEAD SCREW,BALL BEARING,ENCODER AND HARD DISK | LEAD SCREW,BALL BEARING,ENCODER AND HARD DISK | NO | 2024-08-25 | August 2024 | 7:40 | 2024-08-24 19:30 | 2024-08-25 07:40 | 12:10 | COMPLETED | CRITICAL
2089 | 2024-08-26 | Aug 2024 | 8:30 | 8:35 | 0.0034722 | ANESTHESIA MACHINE | WIPRO GE HEALTH CARE | 9100 CNXT | SOORAJ | OT | CIRCUIT LEAK TEST FAILED | LEAK ARRESTED BY REDUCING SODA LIME QTY | NO | NO | 2024-08-26 | August 2024 | 9:10 | 2024-08-26 08:30 | 2024-08-26 09:10 | 0:40 | COMPLETED | CRITICAL
2090 | 2024-08-26 | Aug 2024 | 12:20 | 12:40 | 0.013889 | ECHO MACHINE | GE HEALTHCARE | VIVID S70 | RAMKIE | ICU | SOFTWARE CORRUPTED | COMPANY CHECKED AND UPDATED SOFTWARE, PROBLEM RECTIFIED | NO | NO | 2024-08-26 | August 2024 | 12:55 | 2024-08-26 12:20 | 2024-08-26 12:55 | 0:35 | COMPLETED | SEMI-CRITICAL
2091 | 2024-08-27 | Aug 2024 | 17:40 | 17:50 | 0.0069444 | PATIENT MONITOR | PHILIPS | CM12 | SANTHOSH | WARD | CONTINEOUS ALAMING | ALARM MUTED | NO | NO | 2024-08-27 | August 2024 | 18:20 | 2024-08-27 17:40 | 2024-08-27 18:20 | 0:40 | COMPLETED | SEMI-CRITICAL
2092 | 2024-08-27 | Aug 2024 | 23:20 | 23:30 | 0.0069444 | BLOOD GAS ANALYSER | GEM 4000 | INSTRUMENTATION LABORATORY | SANTHOSH | ICU | CATRIDGE DEFECT | REPLACED NEW CATRIDGE | CARTRIDGE | NO | 2024-08-27 | August 2024 | 23:30 | 2024-08-27 23:20 | 2024-08-27 23:30 | 0:10 | COMPLETED | SEMI-CRITICAL
2093 | 2024-08-28 | Aug 2024 | 12:30 | 12:50 | 0.013889 | VENTILATOR, PEDIATRIC | MAQUET | SERVO AIR | RAMKIE | ICU | VOLUME DELIVERY RECTIFIED | INTRUCTED IT IS WARNING, NORMAL IN PRVC MODE AND NO PROBLEM WITH VENTILATOR | NO | NO | 2024-08-28 | August 2024 | 13:20 | 2024-08-28 12:30 | 2024-08-28 13:20 | 0:50 | COMPLETED | CRITICAL
2094 | 2024-08-29 | Aug 2024 | 12:50 | 13:20 | 0.020833 | WASHER DISINFECTOR | GETINGE | 46-5-403 | RAMKIE | CSSD | ERROR 01 | CHECKED AND CHANGED FILTER AND RECTIFIED | NO | NO | 2024-08-29 | August 2024 | 13:40 | 2024-08-29 12:50 | 2024-08-29 13:40 | 0:50 | COMPLETED | SEMI-CRITICAL
2095 | 2024-08-30 | Aug 2024 | 13:00 | 13:10 | 0.0069444 | TREADMILL TEST | WIPRO GE HEALTH CARE | T2100-ST2 | SANTHOSH | OPD | PRINTER ISSUE | ISSUE RECTIFIED | NO | NO | 2024-08-30 | August 2024 | 13:20 | 2024-08-30 13:00 | 2024-08-30 13:20 | 0:20 | COMPLETED | NON-CRITICAL
Formulas:
STARTING DATE & TIME = [CALL DATE]+[CALL RECEIVED TIME]
ENDING DATE & TIME = [END DATE]+[END TIME]
END DATE = [CALL DATE]